## Hanna Code
Lot | Hanna SFG Code | Recipe | Preparation Week |  | Hanna SFG Code | HI701-0 | HI7014-0 | HI93701-0 | HI96779B-0 | MI526-0
0965 | HI701-0 | CP-R25 | W 10/1 |  | Parameter Formula | Cl2 | Cl2 | Cl2 | ClO2 | Chlorine
Description | Exp | Line |  |  | QC Department | Registration Book | QC Type
Free Chlorine Reagent | 2027-03-01 | L57 Powder |  |  | Laboratory | TABLET 3 | Production
Ref Weight. | Min (mg) | Max (mg)
120 | 110 | 130
REAGENT SET 1 | [5 empty] | REAGENT SET 2
Reagent A Lot | Reagent B Lot | Reagent C Lot | Reagent D Lot | Reagent E Lot |  | Reagent A Lot | Reagent B Lot | Reagent C Lot | Reagent D Lot | Reagent E Lot
Expiration A | Expiration B | Expiration C | Expiration D | Expiration E |  | Expiration A | Expiration B | Expiration C | Expiration D | Expiration E
Reagent Range Min | Reagent Range Max
0.00 mg/L | 2.50 mg/L
Prep. Operator | First day Prod. | Last day Prod. | Machine |  |  | Old Lot A | Old Lot B | Old Lot D
ERIKA | 2025-03-10 | 2025-03-14 | 4HEAD |  |  | 0927
[6 empty] | Expiration A | Expiration B | Expiration D
2027-02-01
Meter
Meter 1 Family | Meter 1 Code | Meter 2 Family | Meter 2 Code
HI 801 | CTO-775 | HI97711 | CTO-633
ph Meter | Description | Turbid. Meter | Description | Spectr. Meter | Description
|  | HI98703 | CTO-860
Standard | mg/L | [5 empty] | Tolerance |  |  |  |  | pH
STD Number | STD Value | STD Min | STD Max |  | STD MR |  | Fixed | AndOr | Percentage | QCRestriction |  | pH Value | pH Min | pH Max
2 | 0.50 | 0.47 | 0.53 |  | MR67C |  | 0.03 | & | 0.9% | 100% |  | 6.0 | 5.7 | 6.3
1 | 0.00 | 0.00 | 0.03
3 | 2.00 | 1.95 | 2.05
Readings Table
Standard | STD Value | # | TEST | QC DATE | QC TIME | PROD. DATE | PROD. TIME | PROD. OPERATOR | HEAD | METER 1 | METER 2 | METER 3 | METER 4 | SPECTR. [ABS] | pH 1 | pH 2 | pH 3 | TURB. | WEIGHT [mg] | REAGENT SET | QC OPERATOR | CORRECTION | phNumber | STD | STD_ID | NOTE
2 | 0.50 | 1 | Old A | 2025-03-13 | 08:48 | 2025-03-13 | 06:58 | ANA | 4 | 0.47 | 0.48 | / | / | 0.250 |  |  |  | 3.47 |  |  | Kinga |  | 3 | 2 | 95753
1 | 0.00 | 2 | Old A | 2025-03-13 | 08:49 | 2025-03-13 | 06:58 | ANA | 4 | 0.00 | 0.00 | / | / | 0.002 |  |  |  | 0.60 |  |  | Kinga |  | 3 | 1 | 95753
3 | 2.00 | 3 | Old A | 2025-03-13 | 08:49 | 2025-03-13 | 06:58 | ANA | 4 | 1.98 | 1.96 | / | / | 1.003 |  |  |  | 1.66 |  |  | Kinga |  | 3 | 3 | 95753
2 | 0.50 | 4 | P/Prod | 2025-03-13 | 08:53 | 2025-03-13 | 06:58 | ANA | 4 | 0.47 | 0.49 | / | / | 0.248 | 6.169 |  |  | 1.04 | 110 |  | Kinga |  | 3 | 2 | 95753
2 | 0.50 | 5 | P/Prod | 2025-03-13 | 08:55 | 2025-03-13 | 06:58 | ANA | 4 | 0.48 | 0.47 | / | / | 0.252 |  |  |  | 1.07 | 122 |  | Kinga |  | 3 | 2 | 95753
3 | 2.00 | 6 | P/Prod | 2025-03-13 | 08:56 | 2025-03-13 | 06:58 | ANA | 4 | 1.95 | 1.96 | / | / | 0.990 | 6.181 |  |  | 0.73 | 102 |  | Kinga |  | 3 | 3 | 95753
3 | 2.00 | 7 | P/Prod | 2025-03-13 | 08:57 | 2025-03-13 | 06:58 | ANA | 4 | 1.98 | 1.97 | / | / | 1.004 |  |  |  | 0.71 | 126 |  | Kinga |  | 3 | 3 | 95753
1 | 0.00 | 8 | P/Prod | 2025-03-13 | 08:59 | 2025-03-13 | 06:58 | ANA | 4 | 0.00 | 0.00 | / | / | 0.005 | 6.424 |  |  | 0.19 |  |  | Kinga |  | 3 | 1 | 95753
2 | 0.50 | 9 | P/Prod | 2025-03-13 | 14:44 | 2025-03-13 | 14:22 | DORINA | 4 | 0.48 | 0.48 | / | / | 0.255 |  |  |  | 0.37 |  |  | Kinga |  | 3 | 2 | 95753
2 | 0.50 | 10 | P/Prod | 2025-03-13 | 14:45 | 2025-03-13 | 14:22 | DORINA | 4 | 0.48 | 0.47 | / | / | 0.255 | 6.293 |  |  | 0.22 |  |  | Kinga |  | 3 | 2 | 95753
3 | 2.00 | 11 | P/Prod | 2025-03-13 | 14:46 | 2025-03-13 | 14:22 | DORINA | 4 | 1.96 | 1.97 | / | / | 0.972 | 6.341 |  |  | 0.24 |  |  | Kinga |  | 3 | 3 | 95753
3 | 2.00 | 12 | P/Prod | 2025-03-13 | 14:47 | 2025-03-13 | 14:22 | DORINA | 4 | 1.95 | 1.95 | / | / | 0.988 |  |  |  | 0.52 |  |  | Kinga |  | 3 | 3 | 95753
1 | 0.00 | 13 | P/Prod | 2025-03-13 | 14:49 | 2025-03-13 | 14:22 | DORINA | 4 | 0.01 | 0.00 | / | / | 0.012 | 6.317 |  |  | 0.67 |  |  | Kinga |  | 3 | 1 | 95753
1 | 0.00 | 14 | P/Prod | 2025-03-13 | 14:49 | 2025-03-13 | 14:22 | DORINA | 4 | 0.00 | 0.00 | / | / | 0.007 |  |  |  | 0.56 |  |  | Kinga |  | 3 | 1 | 95753
2 | 0.50 | 15 | Old A | 2025-03-14 | 08:43 | 2025-03-14 | 06:45 | ANA | 4 | 0.48 | 0.47 | / | / | 0.252 |  |  |  | 7.41 |  |  | Kinga |  | 3 | 2 | 95753
3 | 2.00 | 16 | P/Prod | 2025-03-14 | 08:44 | 2025-03-14 | 06:45 | ANA | 4 | 1.97 | 1.99 | / | / | 1.002 |  |  |  | 0.72 |  |  | Kinga |  | 3 | 3 | 95753
1 | 0.00 | 17 | Old A | 2025-03-14 | 08:45 | 2025-03-14 | 06:45 | ANA | 4 | 0.00 | 0.00 | / | / | 0.003 |  |  |  | 10.6 |  |  | Kinga |  | 3 | 1 | 95753
3 | 2.00 | 18 | P/Prod | 2025-03-14 | 08:49 | 2025-03-14 | 06:45 | ANA | 4 | 1.95 | 1.99 | / | / | 0.986 |  |  |  | 0.58 |  |  | Kinga |  | 3 | 3 | 95753
3 | 2.00 | 19 | P/Prod | 2025-03-14 | 08:50 | 2025-03-14 | 06:45 | ANA | 4 | 1.95 | 1.96 | / | / | 0.984 | 6.537 |  |  | 0.92 |  |  | Kinga |  | 3 | 3 | 95753
2 | 0.50 | 20 | P/Prod | 2025-03-14 | 08:52 | 2025-03-14 | 06:45 | ANA | 4 | 0.49 | 0.47 | / | / | 0.260 |  |  |  | 5.70 | 114 |  | Kinga |  | 3 | 2 | 95753
2 | 0.50 | 21 | P/Prod | 2025-03-14 | 08:53 | 2025-03-14 | 06:45 | ANA | 4 | 0.49 | 0.47 | / | / | 0.257 | 6.222 |  |  | 5.70 | 128 |  | Kinga |  | 3 | 2 | 95753
1 | 0.00 | 22 | P/Prod | 2025-03-14 | 08:54 | 2025-03-14 | 06:45 | ANA | 4 | 0.00 | 0.00 | / | / | 0.004 |  |  |  | 1.03 | 111 |  | Kinga |  | 3 | 1 | 95753
1 | 0.00 | 23 | P/Prod | 2025-03-14 | 08:55 | 2025-03-14 | 06:45 | ANA | 4 | 0.00 | 0.00 | / | / | 0.006 | 5.921 |  |  | 4.73 | 111 |  | Kinga |  | 3 | 1 | 95753
Standard # 2
Total Tests | Total Readings | Total Average | Selected Tests | Selected Readings | Selected Average
6 | 12 | 0.48 mg/L | 6 | 12 | 0.48 mg/L
Standard # 1
Total Tests | Total Readings | Total Average | Selected Tests | Selected Readings | Selected Average
5 | 10 | 0 mg/L | 5 | 10 | 0 mg/L
Standard # 3
Total Tests | Total Readings | Total Average | Selected Tests | Selected Readings | Selected Average
7 | 14 | 1.96 mg/L | 7 | 14 | 1.96 mg/L
Specifications
Standard Value [ppm] | Target Value ÷ U [ppm] | Mean Value [ppm] | Passed
0.00 | 0.00 ÷ 0.03 | 0.00 | YES
0.50 | 0.47 ÷ 0.53 | 0.48 | YES
2.00 | 1.95 ÷ 2.05 | 1.96 | YES
Validation Date | by