## Hanna Code
Lot | Hanna SFG Code | Recipe | Preparation Week |  | Hanna SFG Code | HI705C-0 | HI706B-0 | HI717B-0 | HI93705C-0 | HI93706B-0 | HI93717B-0
1204 | HI93706B-0 | CP-R30 | W 43/1 |  | Parameter Formula | SiO2 | P-PO4 | PO4 | SiO2 | P-PO4 | PO4
Description | Exp | Line |  |  | QC Department | Registration Book | QC Type
Phosphorus Reagent B | 2030-10-01 | L57 Powder |  |  | Laboratory | TABLET 3 | Production
Ref Weight. | Min (mg) | Max (mg)
170 | 160 | 180
REAGENT SET 1 | [5 empty] | REAGENT SET 2
Reagent A Lot | Reagent B Lot | Reagent C Lot | Reagent D Lot | Reagent E Lot |  | Reagent A Lot | Reagent B Lot | Reagent C Lot | Reagent D Lot | Reagent E Lot
2878 | 1204
Expiration A | Expiration B | Expiration C | Expiration D | Expiration E |  | Expiration A | Expiration B | Expiration C | Expiration D | Expiration E
2030-10-01 | 2030-10-01
Reagent Range Min | Reagent Range Max
0.0 mg/L | 15.0 mg/L
Prep. Operator | First day Prod. | Last day Prod. | Machine |  |  | Old Lot A | Old Lot B | Old Lot D
ERIKA | 2025-10-27 | 2025-10-28 | 2 HEAD |  |  |  | 1152
[6 empty] | Expiration A | Expiration B | Expiration D
2030-09-01
Meter
Meter 1 Family | Meter 1 Code
HI97706 | CTO-625
ph Meter | Description | Turbid. Meter | Description | Spectr. Meter | Description
|  | H98703 | CTO-778
Standard | mg/L | [5 empty] | Tolerance |  |  |  |  | pH
STD Number | STD Value | STD Min | STD Max |  | STD MR |  | Fixed | AndOr | Percentage | QCRestriction |  | pH Value | pH Min | pH Max
2 | 3.1 | 2.7 | 3.5 |  | MR104 |  | 0.3 | & | 4% | 100% |  | 0.6 | 0.4 | 0.8
3 | 12.0 | 11.2 | 12.8 | [8 empty] | 1.4 | 1.2 | 1.6
1 | 0.0 | 0.0 | 0.3
Readings Table
Standard | STD Value | # | TEST | QC DATE | QC TIME | PROD. DATE | PROD. TIME | PROD. OPERATOR | HEAD | M. 1 HI97706 | M. 2 Label17 | M. 3 Label17 | M. 4 Label17 | SPECTR. [ABS] | pH 1 | pH 2 | pH 3 | TURB. | WEIGHT [mg] | REAGENT SET | QC OPERATOR | CORRECTION | phNumber | STD | STD_ID | NOTE
2 | 3.1 | 1 | Old B | 2025-10-28 | 08:30 | 2025-10-27 | 16:06 | ELENA | 2 | 3.3 | / | / | / |  |  |  |  | 25.2 |  | 1 | Andrea V |  | 3 | 2
2 | 3.1 | 2 | Old B | 2025-10-28 | 08:32 | 2025-10-27 | 16:06 | ELENA | 2 | 3.2 | / | / | / |  |  |  |  | 18.7 |  | 1 | Andrea V |  | 3 | 2
3 | 12.0 | 3 | Old B | 2025-10-28 | 08:32 | 2025-10-27 | 16:06 | ELENA | 2 | 12.3 | / | / | / | [6 empty] | 1 | Andrea V |  | 3 | 3 | 82713
3 | 12.0 | 4 | Old B | 2025-10-28 | 08:32 | 2025-10-27 | 16:06 | ELENA | 2 | 12.4 | / | / | / | [6 empty] | 1 | Andrea V |  | 3 | 3 | 82713
1 | 0.0 | 5 | Old B | 2025-10-28 | 08:35 | 2025-10-27 | 16:06 | ELENA | 2 | 0.0 | / | / | / |  |  |  |  | 16.5 |  | 1 | Andrea V |  | 3 | 1 | 82713
1 | 0.0 | 6 | Old B | 2025-10-28 | 08:35 | 2025-10-27 | 16:06 | ELENA | 2 | 0.0 | / | / | / |  |  |  |  | 14.9 |  | 1 | Andrea V |  | 3 | 1 | 82713
1 | 0.0 | 7 | P/FINAL | 2025-10-28 | 08:36 | 2025-10-27 | 16:06 | ELENA | 2 | 0.0 | / | / | / |  |  |  |  | 13.2 | 166 | 1 | Andrea V |  | 3 | 1 | 82713
1 | 0.0 | 8 | P/FINAL | 2025-10-28 | 08:38 | 2025-10-27 | 16:06 | ELENA | 2 | 0.0 | / | / | / |  |  | 0.684 |  | 16.0 | 168 | 1 | Andrea V |  | 3 | 1 | 82713
2 | 3.1 | 9 | P/FINAL | 2025-10-28 | 08:41 | 2025-10-27 | 16:06 | ELENA | 2 | 3.2 | / | / | / |  |  |  |  | 18.5 | 174 | 1 | Andrea V |  | 3 | 2 | 82713
2 | 3.1 | 10 | P/FINAL | 2025-10-28 | 08:42 | 2025-10-27 | 16:06 | ELENA | 2 | 3.1 | / | / | / |  |  | 0.775 |  | 10.0 | 168 | 1 | Andrea V |  | 3 | 2 | 82713
2 | 3.1 | 11 | P/FINAL | 2025-10-28 | 08:43 | 2025-10-27 | 16:06 | ELENA | 2 | 3.1 | / | / | / |  |  |  |  | 14.1 | 178 | 1 | Andrea V |  | 3 | 2 | 82713
3 | 12.0 | 12 | P/FINAL | 2025-10-28 | 08:44 | 2025-10-27 | 16:06 | ELENA | 2 | 12.2 | / | / | / | [5 empty] | 177 | 1 | Andrea V |  | 3 | 3 | 82713
3 | 12.0 | 13 | P/FINAL | 2025-10-28 | 08:45 | 2025-10-27 | 16:06 | ELENA | 2 | 12.3 | / | / | / | [6 empty] | 1 | Andrea V |  | 3 | 3 | 82713
3 | 12.0 | 14 | P/FINAL | 2025-10-28 | 08:45 | 2025-10-27 | 16:06 | ELENA | 2 | 12.3 | / | / | / |  |  | 0.674 |  |  |  | 1 | Andrea V |  | 3 | 3 | 82713
Standard # 2
Total Tests | Total Readings | Total Average | Selected Tests | Selected Readings | Selected Average
3 | 3 | 3.1 mg/L | 3 | 3 | 3.1 mg/L
Standard # 3
Total Tests | Total Readings | Total Average | Selected Tests | Selected Readings | Selected Average
3 | 3 | 12.3 mg/L | 3 | 3 | 12.3 mg/L
Standard # 1
Total Tests | Total Readings | Total Average | Selected Tests | Selected Readings | Selected Average
2 | 2 | 0 mg/L | 2 | 2 | 0 mg/L
Specifications
Standard Value [ppm] | Target Value ÷ U [ppm] | Mean Value [ppm] | Passed
0.0 | 0.0 ÷ 0.3 | 0.0 | YES
3.1 | 2.7 ÷ 3.5 | 3.1 | YES
12.0 | 11.2 ÷ 12.8 | 12.3 | YES
Validation Date | by