## Hanna Code
Lot | Hanna SFG Code | Recipe | Preparation Week |  | Hanna SFG Code | HI782-0
1045 | HI782-0 | CP-R80 | W20_2025 |  | Parameter Formula | NO3
Description | Exp | Line |  |  | QC Department | Registration Book | QC Type
Marine Nitrate HR Reagent | 2030-05-01 | L57 Powder |  |  | Laboratory | Tablet 3 | Production
Ref Weight. | Min (mg) | Max (mg)
206 | 196 | 216
REAGENT SET 1 | [5 empty] | REAGENT SET 2
Reagent A Lot | Reagent B Lot | Reagent C Lot | Reagent D Lot | Reagent E Lot |  | Reagent A Lot | Reagent B Lot | Reagent C Lot | Reagent D Lot | Reagent E Lot
1045
Expiration A | Expiration B | Expiration C | Expiration D | Expiration E |  | Expiration A | Expiration B | Expiration C | Expiration D | Expiration E
2030-05-01
Reagent Range Min | Reagent Range Max
0.0 mg/L | 75.0 mg/L
Prep. Operator | First day Prod. | Last day Prod. | Machine |  |  | Old Lot A | Old Lot B | Old Lot D
Mark/Erika | 2025-05-22 | 2025-05-26 | 2HEAD |  |  | 1025
[6 empty] | Expiration A | Expiration B | Expiration D
2030-04-01
Meter
Meter 1 Family | Meter 1 Code | Meter 2 Family | Meter 2 Code
HI97115 | CTO-708 | HI97115 | CTO-740
ph Meter | Description | Turbid. Meter | Description | Spectr. Meter | Description
|  | HI98703 | CTO-778
Standard | mg/L | [5 empty] | Tolerance |  |  |  |  | pH
STD Number | STD Value | STD Min | STD Max |  | STD MR |  | Fixed | AndOr | Percentage | QCRestriction |  | pH Value | pH Min | pH Max
2 | 15.0 | 12.3 | 17.7 |  | MR101 |  | 2.0 | & | 4.9% | 100%
3 | 35.0 | 31.3 | 38.7
4 | 60.0 | 55.1 | 64.9 | [8 empty] | 1.7 | 1.5 | 1.9
1 | 0.0 | 0.0 | 2.0
Readings Table
Standard | STD Value | # | TEST | QC DATE | QC TIME | PROD. DATE | PROD. TIME | PROD. OPERATOR | HEAD | METER 1 | METER 2 | METER 3 | METER 4 | SPECTR. [ABS] | pH 1 | pH 2 | pH 3 | TURB. | WEIGHT [mg] | REAGENT SET | QC OPERATOR | CORRECTION | phNumber | STD | STD_ID | NOTE
2 | 15.0 | 1 | Old A | 2025-05-15 | 17:37 | 2025-05-15 | 16:08 | Mark/Erika |  | 14.8 | 14.9 | / | / |  |  |  | 1.983 | 2.15 | 214 | 1 | Monika |  | 3 | 2
2 | 15.0 | 2 | Old A | 2025-05-15 | 17:38 | 2025-05-15 | 16:08 | Mark/Erika |  | 15.1 | 15.1 | / | / |  |  |  |  | 1.41 | 206 | 1 | Monika |  | 3 | 2
3 | 35.0 | 3 | Old A | 2025-05-15 | 17:39 | 2025-05-15 | 16:08 | Mark/Erika |  | 33.2 | 33.4 | / | / |  |  |  | 1.950 | 1.67 | 206 | 1 | Monika |  | 3 | 3 | 84815
3 | 35.0 | 4 | Old A | 2025-05-15 | 17:40 | 2025-05-15 | 16:08 | Mark/Erika |  | 36.6 | 36.7 | / | / |  |  |  |  | 2.11 | 212 | 1 | Monika |  | 3 | 3 | 84815
4 | 60.0 | 5 | Old A | 2025-05-15 | 17:41 | 2025-05-15 | 16:08 | Mark/Erika |  | 56.8 | 55.7 | / | / |  |  |  | 1.899 | 1.32 | 208 | 1 | Monika |  | 3 | 4 | 23707
4 | 60.0 | 6 | Old A | 2025-05-15 | 17:41 | 2025-05-15 | 16:08 | Mark/Erika |  | 60.0 | 59.7 | / | / |  |  |  |  | 1.59 | 215 | 1 | Monika |  | 3 | 4 | 23707
2 | 15.0 | 7 | Valid | 2025-05-15 | 17:46 | 2025-05-15 | 16:08 | Mark/Erika |  | 17.3 | 17.3 | / | / |  |  |  | 1.950 | 0.71 | 206 | 1 | Monika |  | 3 | 2 | 23707
2 | 15.0 | 8 | Valid | 2025-05-15 | 17:47 | 2025-05-15 | 16:08 | Mark/Erika |  | 17.9 | 17.8 | / | / |  |  |  | 1.935 | 0.75 | 206 | 1 | Monika |  | 3 | 2 | 23707
3 | 35.0 | 9 | Valid | 2025-05-15 | 17:49 | 2025-05-15 | 16:08 | Mark/Erika |  | 42.1 | 42.1 | / | / |  |  |  | 1.940 | 1.76 | 206 | 1 | Monika |  | 3 | 3 | 84815
3 | 35.0 | 10 | Valid | 2025-05-15 | 17:49 | 2025-05-15 | 16:08 | Mark/Erika |  | 37.2 | 37.3 | / | / |  |  |  | 1.931 | 1.27 | 206 | 1 | Monika |  | 3 | 3 | 84815
4 | 60.0 | 11 | Valid | 2025-05-15 | 17:52 | 2025-05-15 | 16:08 | Mark/Erika |  | 61.1 | 62.1 | / | / |  |  |  | 1.946 | 1.36 | 206 | 1 | Monika |  | 3 | 4 | 23707
4 | 60.0 | 12 | Valid | 2025-05-15 | 17:53 | 2025-05-15 | 16:08 | Mark/Erika |  | 61.5 | 61.1 | / | / |  |  |  | 1.945 | 1.57 | 206 | 1 | Monika |  | 3 | 4 | 23707
1 | 0.0 | 13 | ASTM D665 | 2025-05-15 | 18:07 | 2025-05-15 | 16:08 | Mark/Erika |  | 0.0 | 0.0 | / | / |  |  |  | 1.950 | 1.41 | 206 | 1 | Monika |  | 3 | 1 |  | LOT:1230604035
3 | 35.0 | 14 | Valid | 2025-05-15 | 18:09 | 2025-05-15 | 16:08 | Mark/Erika |  | 37.6 | 37.7 | / | / |  |  |  | 1.950 | 1.27 | 206 | 1 | Monika |  | 3 | 3 | 84815
2 | 15.0 | 15 | P/FINAL | 2025-05-21 | 09:24 | 2025-05-21 | 08:52 | GIORGIANA | 2 | 14.6 | 14.5 | / | / |  |  |  | 1.876 | 0.96 | 203 | 1 | Brigitta P |  | 3 | 2 | 113042
2 | 15.0 | 16 | P/FINAL | 2025-05-21 | 09:25 | 2025-05-21 | 08:52 | GIORGIANA | 2 | 12.4 | 12.4 | / | / |  |  |  |  | 1.14 | 214 | 1 | Brigitta P |  | 3 | 2 | 113042
3 | 35.0 | 17 | P/FINAL | 2025-05-21 | 09:25 | 2025-05-21 | 08:52 | GIORGIANA | 2 | 36.2 | 36.2 | / | / |  |  |  | 1.860 | 1.14 | 210 | 1 | Brigitta P |  | 3 | 3 | 113042
3 | 35.0 | 18 | P/FINAL | 2025-05-21 | 09:26 | 2025-05-21 | 08:52 | GIORGIANA | 2 | 38.2 | 38.3 | / | / |  |  |  |  | 1.79 | 215 | 1 | Brigitta P |  | 3 | 3 | 113042
4 | 60.0 | 19 | P/FINAL | 2025-05-21 | 09:28 | 2025-05-21 | 08:52 | GIORGIANA | 2 | 58.5 | 58.0 | / | / |  |  |  | 1.880 | 1 | 203 | 1 | Brigitta P |  | 3 | 4 | 113042
4 | 60.0 | 20 | P/FINAL | 2025-05-21 | 09:29 | 2025-05-21 | 08:52 | GIORGIANA | 2 | 59.4 | 60.0 | / | / |  |  |  |  | 0.96 | 205 | 1 | Brigitta P |  | 3 | 4 | 113042
2 | 15.0 | 21 | P/FINAL | 2025-05-21 | 10:50 | 2025-05-21 | 10:34 | GIORGIANA | 2 | 15.7 | 15.9 | / | / |  |  |  | 1.872 | 0.74 | 204 | 1 | Brigitta P |  | 3 | 2 | 113042
2 | 15.0 | 22 | P/FINAL | 2025-05-21 | 10:50 | 2025-05-21 | 10:34 | GIORGIANA | 2 | 16.3 | 16.3 | / | / |  |  |  |  | 0.74 | 208 | 1 | Brigitta P |  | 3 | 2 | 113042
3 | 35.0 | 23 | P/FINAL | 2025-05-21 | 10:54 | 2025-05-21 | 10:34 | GIORGIANA | 2 | 35.2 | 35.3 | / | / |  |  |  | 1.882 | 0.83 | 200 | 1 | Brigitta P |  | 3 | 3 | 113042
3 | 35.0 | 24 | P/FINAL | 2025-05-21 | 10:54 | 2025-05-21 | 10:34 | GIORGIANA | 2 | 37.0 | 37.1 | / | / |  |  |  |  | 0.83 | 211 | 1 | Brigitta P |  | 3 | 3 | 113042
4 | 60.0 | 25 | P/FINAL | 2025-05-21 | 10:56 | 2025-05-21 | 10:34 | GIORGIANA | 2 | 59.7 | 59.8 | / | / |  |  |  | 1.885 | 1.32 | 193 | 1 | Brigitta P |  | 3 | 4 | 113042
4 | 60.0 | 26 | P/FINAL | 2025-05-21 | 10:56 | 2025-05-21 | 10:34 | GIORGIANA | 2 | 61.4 | 60.9 | / | / |  |  |  |  | 2.34 | 203 | 1 | Brigitta P |  | 3 | 4 | 113042
2 | 15.0 | 27 | P/FINAL | 2025-05-21 | 12:47 | 2025-05-21 | 12:21 | GIORGIANA | 2 | 15.4 | 15.3 | / | / |  |  |  | 1.886 | 1.81 | 215 | 1 | Kinga |  | 3 | 2 | 113042
2 | 15.0 | 28 | P/FINAL | 2025-05-21 | 12:48 | 2025-05-21 | 12:21 | GIORGIANA | 2 | 15.1 | 15.1 | / | / |  |  |  |  | 1.18 | 209 | 1 | Kinga |  | 3 | 2 | 113042
3 | 35.0 | 29 | P/FINAL | 2025-05-21 | 12:49 | 2025-05-21 | 12:21 | GIORGIANA | 2 | 36.5 | 36.5 | / | / |  |  |  | 1.889 | 4.05 | 210 | 1 | Kinga |  | 3 | 3 | 113042
3 | 35.0 | 30 | P/FINAL | 2025-05-21 | 12:50 | 2025-05-21 | 12:21 | GIORGIANA | 2 | 35.3 | 34.9 | / | / |  |  |  |  | 6.39 | 206 | 1 | Kinga |  | 3 | 3 | 113042
4 | 60.0 | 31 | P/FINAL | 2025-05-21 | 12:51 | 2025-05-21 | 12:21 | GIORGIANA | 2 | 58.9 | 58.3 | / | / |  |  |  | 1.878 | 1.13 | 213 | 1 | Kinga |  | 3 | 4 | 113042
4 | 60.0 | 32 | P/FINAL | 2025-05-21 | 12:52 | 2025-05-21 | 12:21 | GIORGIANA | 2 | 58.4 | 58.6 | / | / |  |  |  |  | 2.50 | 216 | 1 | Kinga |  | 3 | 4 | 113042
2 | 15.0 | 33 | Old A | 2025-05-21 | 13:01 | 2025-05-21 | 12:21 | GIORGIANA | 2 | 14.1 | 14.1 | / | / |  |  |  | 1.905 | 3.20 |  | 1 | Kinga |  | 3 | 2 | 124249
3 | 35.0 | 34 | Old A | 2025-05-21 | 13:36 | 2025-05-21 | 12:21 | GIORGIANA | 2 | 35.2 | 35.3 | / | / |  |  |  |  | 1.37 |  | 1 | Kinga |  | 3 | 3 | 124249
4 | 60.0 | 35 | Old A | 2025-05-21 | 13:36 | 2025-05-21 | 12:21 | GIORGIANA | 2 | 59.3 | 59.0 | / | / |  |  |  |  | 1.88 |  | 1 | Kinga |  | 3 | 4 | 124249
2 | 15.0 | 36 | P/FINAL | 2025-05-21 | 14:53 | 2025-05-21 | 14:30 | ROLAND | 2 | 13.8 | 13.8 | / | / |  |  |  |  | 0.52 |  | 1 | Kinga |  | 3 | 2 | 113042
2 | 15.0 | 37 | P/FINAL | 2025-05-21 | 14:54 | 2025-05-21 | 14:30 | ROLAND | 2 | 14.9 | 15.1 | / | / |  |  |  |  | 0.57 |  | 1 | Kinga |  | 3 | 2 | 113042
3 | 35.0 | 38 | P/FINAL | 2025-05-21 | 14:57 | 2025-05-21 | 14:30 | ROLAND | 2 | 33.4 | 33.7 | / | / |  |  |  |  | 0.55 |  | 1 | Kinga |  | 3 | 3 | 113042
3 | 35.0 | 39 | P/FINAL | 2025-05-21 | 14:57 | 2025-05-21 | 14:30 | ROLAND | 2 | 35.4 | 35.5 | / | / |  |  |  |  | 0.57 |  | 1 | Kinga |  | 3 | 3 | 113042
4 | 60.0 | 40 | P/FINAL | 2025-05-21 | 15:09 | 2025-05-21 | 14:30 | ROLAND | 2 | 55.1 | 55.1 | / | / |  |  |  |  | 0.62 |  | 1 | Kinga |  | 3 | 4 | 113042
4 | 60.0 | 41 | P/FINAL | 2025-05-21 | 15:10 | 2025-05-21 | 14:30 | ROLAND | 2 | 58.8 | 59.2 | / | / |  |  |  |  | 0.83 |  | 1 | Kinga |  | 3 | 4 | 113042
2 | 15.0 | 42 | P/FINAL | 2025-05-21 | 17:29 | 2025-05-21 | 16:59 | ROLAND | 2 | 13.9 | 13.9 | / | / |  |  |  | 1.905 | 0.67 |  | 1 | Kinga |  | 3 | 2 | 113042
2 | 15.0 | 43 | P/FINAL | 2025-05-21 | 17:29 | 2025-05-21 | 16:59 | ROLAND | 2 | 14.7 | 14.7 | / | / |  |  |  |  | 0.67 |  | 1 | Kinga |  | 3 | 2 | 113042
3 | 35.0 | 44 | P/FINAL | 2025-05-21 | 17:32 | 2025-05-21 | 16:59 | ROLAND | 2 | 31.4 | 31.6 | / | / |  |  |  | 1.905 | 0.54 |  | 1 | Kinga |  | 3 | 3 | 113042
3 | 35.0 | 45 | P/FINAL | 2025-05-21 | 17:33 | 2025-05-21 | 16:59 | ROLAND | 2 | 36.8 | 36.9 | / | / |  |  |  |  | 0.75 |  | 1 | Kinga |  | 3 | 3 | 113042
4 | 60.0 | 46 | P/FINAL | 2025-05-21 | 17:44 | 2025-05-21 | 16:59 | ROLAND | 2 | 60.9 | 60.9 | / | / |  |  |  | 1.843 | 1.13 |  | 1 | Kinga |  | 3 | 4 | 113042
4 | 60.0 | 47 | P/FINAL | 2025-05-21 | 17:45 | 2025-05-21 | 16:59 | ROLAND | 2 | 59.0 | 59.0 | / | / |  |  |  |  | 1.23 |  | 1 | Kinga |  | 3 | 4 | 113042
2 | 15.0 | 48 | P/FINAL | 19:2025/05/21 | 19:30 | 2025-05-21 | 19:02 | ROLAND | 2 | 14.3 | 14.3 | / | / |  |  |  | 1.843 | 0.60 |  | 1 | Kinga |  | 3 | 2 | 113042
2 | 15.0 | 49 | P/FINAL | 19:2025/05/21 | 19:30 | 2025-05-21 | 19:02 | ROLAND | 2 | 17.3 | 17.3 | / | / |  |  |  |  | 0.77 |  | 1 | Kinga |  | 3 | 2 | 113042
3 | 35.0 | 50 | P/FINAL | 19:2025/05/21 | 19:34 | 2025-05-21 | 19:02 | ROLAND | 2 | 34.2 | 34.2 | / | / |  |  |  | 1.909 | 0.65 |  | 1 | Kinga |  | 3 | 3 | 113042
3 | 35.0 | 51 | P/FINAL | 19:2025/05/21 | 19:34 | 2025-05-21 | 19:02 | ROLAND | 2 | 36.8 | 36.3 | / | / |  |  |  |  | 0.69 |  | 1 | Kinga |  | 3 | 3 | 113042
4 | 60.0 | 52 | P/FINAL | 19:2025/05/21 | 19:45 | 2025-05-21 | 19:02 | ROLAND | 2 | 63.6 | 63.6 | / | / |  |  |  | 1.910 | 0.95 |  | 1 | Kinga |  | 3 | 4 | 113042
4 | 60.0 | 53 | P/FINAL | 19:2025/05/21 | 19:46 | 2025-05-21 | 19:02 | ROLAND | 2 | 59.1 | 58.7 | / | / |  |  |  | 1.928 | 1.29 |  | 1 | Kinga |  | 3 | 4 | 113042
2 | 15.0 | 54 | P/FINAL | 19:2025/05/21 | 21:17 | 2025-05-21 | 20:56 | ROLAND | 2 | 14.2 | 14.3 | / | / |  |  |  |  | 0.55 |  | 1 | Kinga |  | 3 | 2 | 113042
2 | 15.0 | 55 | P/FINAL | 19:2025/05/21 | 21:18 | 2025-05-21 | 20:56 | ROLAND | 2 | 14.6 | 14.8 | / | / |  |  |  | 1.924 | 0.52 |  | 1 | Kinga |  | 3 | 2 | 113042
3 | 35.0 | 56 | P/FINAL | 19:2025/05/21 | 21:21 | 2025-05-21 | 20:56 | ROLAND | 2 | 35.8 | 36.3 | / | / |  |  |  |  | 0.54 |  | 1 | Kinga |  | 3 | 3 | 113042
3 | 35.0 | 57 | P/FINAL | 19:2025/05/21 | 21:22 | 2025-05-21 | 20:56 | ROLAND | 2 | 33.3 | 33.3 | / | / |  |  |  | 1.925 | 0.47 |  | 1 | Kinga |  | 3 | 3 | 113042
4 | 60.0 | 58 | P/FINAL | 19:2025/05/21 | 21:33 | 2025-05-21 | 20:56 | ROLAND | 2 | 56.4 | 56.5 | / | / |  |  |  |  | 0.65 |  | 1 | Kinga |  | 3 | 4 | 113042
4 | 60.0 | 59 | P/FINAL | 19:2025/05/21 | 21:33 | 2025-05-21 | 20:56 | ROLAND | 2 | 64.9 | 64.4 | / | / |  |  |  | 1.880 | 1.09 |  | 1 | Kinga |  | 3 | 4 | 113042
2 | 15.0 | 60 | P/FINAL | 2025-05-22 | 07:30 | 2025-05-22 | 06:24 | GIORGIANA | 2 | 15.8 | 15.8 | / | / |  |  |  |  | 0.62 | 205 | 1 | Brigitta P |  | 3 | 2 | 113042
2 | 15.0 | 61 | P/FINAL | 2025-05-22 | 07:30 | 2025-05-22 | 06:24 | GIORGIANA | 2 | 15.1 | 15.1 | / | / |  |  |  |  | 0.45 | 193 | 1 | Brigitta P |  | 3 | 2 | 113042
3 | 35.0 | 62 | P/FINAL | 2025-05-22 | 07:31 | 2025-05-22 | 06:24 | GIORGIANA | 2 | 38.7 | 38.7 | / | / |  |  |  |  | 3.13 | 191 | 1 | Brigitta P |  | 3 | 3 | 113042
3 | 35.0 | 63 | P/FINAL | 2025-05-22 | 07:32 | 2025-05-22 | 06:24 | GIORGIANA | 2 | 38.9 | 38.7 | / | / |  |  |  |  | 2.15 | 205 | 1 | Brigitta P |  | 3 | 3 | 113042
4 | 60.0 | 64 | P/FINAL | 2025-05-22 | 07:33 | 2025-05-22 | 06:24 | GIORGIANA | 2 | 64.8 | 64.8 | / | / |  |  |  |  | 0.91 | 189 | 1 | Brigitta P |  | 3 | 4 | 113042
4 | 60.0 | 65 | P/FINAL | 2025-05-22 | 07:33 | 2025-05-22 | 06:24 | GIORGIANA | 2 | 64.5 | 64.5 | / | / |  |  |  |  | 0.91 | 195 | 1 | Brigitta P |  | 3 | 4 | 113042
2 | 15.0 | 66 | P/FINAL | 2025-05-22 | 08:43 | 2025-05-22 | 08:12 | GIORGIANA | 2 | 14.8 | 14.9 | / | / |  |  |  | 1.944 | 0.52 | 195 | 1 | Brigitta P |  | 3 | 2 | 113042
2 | 15.0 | 67 | P/FINAL | 2025-05-22 | 08:44 | 2025-05-22 | 08:12 | GIORGIANA | 2 | 16.0 | 16.1 | / | / |  |  |  |  | 0.50 | 209 | 1 | Brigitta P |  | 3 | 2 | 113042
3 | 35.0 | 68 | P/FINAL | 2025-05-22 | 08:46 | 2025-05-22 | 08:12 | GIORGIANA | 2 | 36.2 | 36.5 | / | / |  |  |  | 1.907 | 0.78 | 195 | 1 | Brigitta P |  | 3 | 3 | 113042
3 | 35.0 | 69 | P/FINAL | 2025-05-22 | 08:47 | 2025-05-22 | 08:12 | GIORGIANA | 2 | 37.9 | 38.1 | / | / |  |  |  |  | 0.80 | 195 | 1 | Brigitta P |  | 3 | 3 | 113042
4 | 60.0 | 70 | P/FINAL | 2025-05-22 | 08:50 | 2025-05-22 | 08:12 | GIORGIANA | 2 | 59.5 | 59.6 | / | / |  |  |  | 1.900 | 0.85 | 202 | 1 | Brigitta P |  | 3 | 4 | 113042
4 | 60.0 | 71 | P/FINAL | 2025-05-22 | 08:52 | 2025-05-22 | 08:12 | GIORGIANA | 2 | 61.5 | 61.5 | / | / |  |  |  |  | 0.82 | 202 | 1 | Brigitta P |  | 3 | 4 | 113042
2 | 15.0 | 72 | P/FINAL | 2025-05-22 | 11:04 | 2025-05-22 | 10:32 | GIORGIANA | 2 | 15.0 | 14.9 | / | / |  |  |  |  | 0.69 | 204 | 1 | Brigitta P |  | 3 | 2 | 113042
2 | 15.0 | 73 | P/FINAL | 2025-05-22 | 11:04 | 2025-05-22 | 10:32 | GIORGIANA | 2 | 16.5 | 15.9 | / | / |  |  |  | 1.873 | 0.80 | 212 | 1 | Brigitta P |  | 3 | 2 | 113042
3 | 35.0 | 74 | P/FINAL | 2025-05-22 | 11:07 | 2025-05-22 | 10:32 | GIORGIANA | 2 | 36.1 | 36.1 | / | / |  |  |  | 1.885 | 0.80 | 208 | 1 | Brigitta P |  | 3 | 3 | 113042
3 | 35.0 | 75 | P/FINAL | 2025-05-22 | 11:08 | 2025-05-22 | 10:32 | GIORGIANA | 2 | 36.5 | 36.7 | / | / |  |  |  |  | 0.79 | 209 | 1 | Brigitta P |  | 3 | 3 | 113042
4 | 60.0 | 76 | P/FINAL | 2025-05-22 | 11:05 | 2025-05-22 | 10:32 | GIORGIANA | 2 | 59.5 | 59.5 | / | / |  |  |  | 1.900 | 0.89 | 215 | 1 | Brigitta P |  | 3 | 3 | 113042
4 | 60.0 | 77 | P/FINAL | 2025-05-22 | 11:05 | 2025-05-22 | 10:32 | GIORGIANA | 2 | 62.1 | 62.3 | / | / |  |  |  |  | 0.1.05 | 212 | 1 | Brigitta P |  | 3 | 3 | 113042
2 | 15.0 | 78 | P/FINAL | 2025-05-22 | 12:34 | 2025-05-22 | 12:17 | GIORGIANA | 2 | 14.6 | 14.8 | / | / |  |  |  |  | 0.76 | 198 | 1 | Brigitta P |  | 3 | 2 | 124249
2 | 15.0 | 79 | P/FINAL | 2025-05-22 | 12:34 | 2025-05-22 | 12:17 | GIORGIANA | 2 | 15.7 | 15.8 | / | / |  |  |  |  | 0.62 | 210 | 1 | Brigitta P |  | 3 | 2 | 124249
3 | 35.0 | 80 | P/FINAL | 2025-05-22 | 12:37 | 2025-05-22 | 12:17 | GIORGIANA | 2 | 33.5 | 33.5 | / | / |  |  |  |  | 0.65 | 196 | 1 | Brigitta P |  | 3 | 3 | 124249
3 | 35.0 | 81 | P/FINAL | 2025-05-22 | 12:38 | 2025-05-22 | 12:17 | GIORGIANA | 2 | 35.5 | 36.1 | / | / |  |  |  |  | 0.64 | 202 | 1 | Brigitta P |  | 3 | 3 | 124249
4 | 60.0 | 82 | P/FINAL | 2025-05-22 | 12:40 | 2025-05-22 | 12:17 | GIORGIANA | 2 | 59.4 | 59.4 | / | / |  |  |  |  | 0.87 | 198 | 1 | Brigitta P |  | 3 | 4 | 124249
4 | 60.0 | 83 | P/FINAL | 2025-05-22 | 12:41 | 2025-05-22 | 12:17 | GIORGIANA | 2 | 58.0 | 58.1 | / | / |  |  |  |  | 0.66 | 198 | 1 | Brigitta P |  | 3 | 4 | 124249
2 | 15.0 | 84 | P/FINAL | 2025-05-22 | 14:40 | 2025-05-22 | 14:19 | ROLAND | 2 | 14.3 | 14.2 | / | / |  |  |  | 1.857 | 0.50 |  | 1 | Brigitta P |  | 3 | 2 | 124249
2 | 15.0 | 85 | P/FINAL | 2025-05-22 | 14:41 | 2025-05-22 | 14:19 | ROLAND | 2 | 14.3 | 14.4 | / | / |  |  |  |  | 0.67 |  | 1 | Brigitta P |  | 3 | 2 | 124249
3 | 35.0 | 86 | P/FINAL | 2025-05-22 | 14:43 | 2025-05-22 | 14:19 | ROLAND | 2 | 32.3 | 32.4 | / | / |  |  |  | 1.899 | 0.53 |  | 1 | Brigitta P |  | 3 | 3 | 124249
3 | 35.0 | 87 | P/FINAL | 2025-05-22 | 14:44 | 2025-05-22 | 14:19 | ROLAND | 2 | 34.4 | 34.4 | / | / |  |  |  |  | 0.52 |  | 1 | Brigitta P |  | 3 | 3 | 124249
4 | 60.0 | 88 | P/FINAL | 2025-05-22 | 14:55 | 2025-05-22 | 14:19 | ROLAND | 2 | 55.3 | 55.6 | / | / |  |  |  | 1.893 | 0.67 |  | 1 | Brigitta P |  | 3 | 4 | 124249
4 | 60.0 | 89 | P/FINAL | 2025-05-22 | 14:56 | 2025-05-22 | 14:19 | ROLAND | 2 | 61.5 | 62.0 | / | / |  |  |  |  | 0.77 |  | 1 | Brigitta P |  | 3 | 4 | 124249
2 | 15.0 | 90 | P/FINAL | 2025-05-22 | 16:08 | 2025-05-22 | 15:45 | ROLAND |  | 14.5 | 14.6 | / | / |  |  |  | 1.883 | 0.43 |  | 1 | Brigitta P |  | 3 | 2 | 124249
2 | 15.0 | 91 | P/FINAL | 2025-05-22 | 16:09 | 2025-05-22 | 15:45 | ROLAND |  | 14.5 | 14.6 | / | / |  |  |  |  | 0.46 |  | 1 | Brigitta P |  | 3 | 2 | 124249
3 | 35.0 | 92 | P/FINAL | 2025-05-22 | 16:11 | 2025-05-22 | 15:45 | ROLAND |  | 36.9 | 37.4 | / | / |  |  |  | 1.954 | 0.68 |  | 1 | Brigitta P |  | 3 | 3 | 124249
3 | 35.0 | 93 | P/FINAL | 2025-05-22 | 16:12 | 2025-05-22 | 15:45 | ROLAND |  | 34.1 | 34.5 | / | / |  |  |  |  | 0.56 |  | 1 | Brigitta P |  | 3 | 3 | 124249
4 | 60.0 | 94 | P/FINAL | 2025-05-22 | 16:14 | 2025-05-22 | 15:45 | ROLAND |  | 60.9 | 60.5 | / | / |  |  |  | 1.928 | 0.90 |  | 1 | Brigitta P |  | 3 | 4 | 124249
4 | 60.0 | 95 | P/FINAL | 2025-05-22 | 16:15 | 2025-05-22 | 15:45 | ROLAND |  | 55.7 | 55.3 | / | / |  |  |  |  | 0.54 |  | 1 | Brigitta P |  | 3 | 4 | 124249
2 | 15.0 | 96 | P/FINAL | 2025-05-22 | 19:53 | 2025-05-22 | 19:17 | ROLAND | 2 | 13.1 | 13.2 | / | / |  |  |  | 1.928 | 0.58 |  | 1 | Brigitta P |  | 3 | 2 | 124249
2 | 15.0 | 97 | P/FINAL | 2025-05-22 | 19:53 | 2025-05-22 | 19:17 | ROLAND | 2 | 15.7 | 15.7 | / | / |  |  |  |  | 0.45 |  | 1 | Brigitta P |  | 3 | 2 | 124249
3 | 35.0 | 98 | P/FINAL | 2025-05-22 | 19:55 | 2025-05-22 | 19:17 | ROLAND | 2 | 35.1 | 35.2 | / | / |  |  |  | 1.885 | 0.50 |  | 1 | Brigitta P |  | 3 | 3 | 124249
3 | 35.0 | 99 | P/FINAL | 2025-05-22 | 19:56 | 2025-05-22 | 19:17 | ROLAND | 2 | 32.6 | 32.6 | / | / |  |  |  |  | 0.55 |  | 1 | Brigitta P |  | 3 | 3 | 124249
4 | 60.0 | 100 | P/FINAL | 2025-05-22 | 20:07 | 2025-05-22 | 19:17 | ROLAND | 2 | 58.6 | 58.5 | / | / |  |  |  | 1.938 | 0.59 |  | 1 | Brigitta P |  | 3 | 4 | 124249
4 | 60.0 | 101 | P/FINAL | 2025-05-22 | 20:08 | 2025-05-22 | 19:17 | ROLAND | 2 | 59.5 | 59.6 | / | / |  |  |  |  | 0.58 |  | 1 | Brigitta P |  | 3 | 4 | 124249
2 | 15.0 | 102 | P/FINAL | 2025-05-22 | 21:00 | 2025-05-22 | 20:32 | ROLAND | 2 | 14.8 | 14.9 | / | / |  |  |  | 1.968 | 0.61 |  | 1 | Brigitta P |  | 3 | 2 | 124249
2 | 15.0 | 103 | P/FINAL | 2025-05-22 | 21:01 | 2025-05-22 | 20:32 | ROLAND | 2 | 16.0 | 16.2 | / | / |  |  |  |  | 0.45 |  | 1 | Brigitta P |  | 3 | 2 | 124249
3 | 35.0 | 104 | P/FINAL | 2025-05-22 | 21:03 | 2025-05-22 | 20:32 | ROLAND | 2 | 37.9 | 38.2 | / | / |  |  |  | 1.885 | 0.60 |  | 1 | Brigitta P |  | 3 | 3 | 124249
3 | 35.0 | 105 | P/FINAL | 2025-05-22 | 21:04 | 2025-05-22 | 20:32 | ROLAND | 2 | 34.0 | 34.0 | / | / |  |  |  |  | 0.55 |  | 1 | Brigitta P |  | 3 | 3 | 124249
4 | 60.0 | 106 | P/FINAL | 2025-05-22 | 21:15 | 2025-05-22 | 20:32 | ROLAND | 2 | 56.0 | 55.8 | / | / |  |  |  | 1.928 | 1.03 |  | 1 | Brigitta P |  | 3 | 4 | 124249
4 | 60.0 | 107 | P/FINAL | 2025-05-22 | 21:15 | 2025-05-22 | 20:32 | ROLAND | 2 | 56.8 | 56.7 | / | / | 0 |  |  |  | 0.63 |  | 1 | Brigitta P |  | 3 | 4 | 124249
2 | 15.0 | 108 | P/FINAL | 2025-05-23 | 07:40 | 2025-05-23 |  | ELENA | 2 | 15.5 | 15.5 | / | / |  |  |  |  | 0.53 |  | 1 | Brigitta P |  | 3 | 2
2 | 15.0 | 109 | P/FINAL | 2025-05-23 | 07:41 | 2025-05-23 |  | ELENA | 2 | 16.4 | 16.4 | / | / |  |  |  |  | 0.52 |  | 1 | Brigitta P |  | 3 | 2
3 | 35.0 | 110 | P/FINAL | 2025-05-23 | 07:42 | 2025-05-23 | 06:48 | ELENA | 2 | 36.3 | 36.3 | / | / |  |  |  |  | 0.63 |  | 1 | Brigitta P |  | 3 | 3
3 | 35.0 | 111 | P/FINAL | 2025-05-23 | 07:43 | 2025-05-23 | 06:48 | ELENA | 2 | 38.1 | 38.3 | / | / |  |  |  |  | 0.68 |  | 1 | Brigitta P |  | 3 | 3
4 | 60.0 | 112 | P/FINAL | 2025-05-23 | 07:45 | 2025-05-23 | 06:48 | ELENA | 2 | 58.2 | 58.3 | / | / |  |  |  |  | 0.75 |  | 1 | Brigitta P |  | 3 | 4
4 | 60.0 | 113 | P/FINAL | 2025-05-23 | 07:45 | 2025-05-23 | 06:48 | ELENA | 2 | 60.0 | 60.0 | / | / |  |  |  |  | 0.74 |  | 1 | Brigitta P |  | 3 | 4
2 | 15.0 | 114 | P/FINAL | 2025-05-23 | 09:19 | 2025-05-23 | 08:22 | ELENA | 2 | 15.6 | 15.6 | / | / |  |  |  |  | 0.46 |  | 1 | Brigitta P |  | 3 | 2 | 124249
2 | 15.0 | 115 | P/FINAL | 2025-05-23 | 09:20 | 2025-05-23 | 08:22 | ELENA | 2 | 16.4 | 16.4 | / | / |  |  |  |  | 0.58 |  | 1 | Brigitta P |  | 3 | 2 | 124249
4 | 60.0 | 116 | P/FINAL | 2025-05-23 | 09:21 | 2025-05-23 | 08:22 | ELENA | 2 | 57.0 | 56.8 | / | / |  |  |  |  | 1.17 |  | 1 | Brigitta P |  | 3 | 4 | 124249
4 | 60.0 | 117 | P/FINAL | 2025-05-23 | 09:22 | 2025-05-23 | 08:22 | ELENA | 2 | 64.1 | 64.0 | / | / |  |  |  |  | 0.71 |  | 1 | Brigitta P |  | 3 | 4 | 124249
3 | 35.0 | 118 | P/FINAL | 2025-05-23 | 09:24 | 2025-05-23 | 08:22 | ELENA | 2 | 35.4 | 35.6 | / | / |  |  |  |  | 0.60 |  | 1 | Brigitta P |  | 3 | 3 | 124249
3 | 35.0 | 119 | P/FINAL | 2025-05-23 | 09:26 | 2025-05-23 | 08:22 | ELENA | 2 | 37.2 | 37.4 | / | / |  |  |  |  | 0.53 |  | 1 | Brigitta P |  | 3 | 3 | 124249
2 | 15.0 | 120 | P/FINAL | 2025-05-23 | 10:35 | 2025-05-23 | 09:56 | ELENA | 2 | 15.2 | 15.3 | / | / |  |  |  | 1.956 | 0.67 | 192 | 1 | Kinga |  | 3 | 2 | 124249
2 | 15.0 | 121 | P/FINAL | 2025-05-23 | 10:36 | 2025-05-23 | 09:56 | ELENA | 2 | 15.1 | 15.3 | / | / |  |  |  |  | 1.89 | 194 | 1 | Kinga |  | 3 | 2 | 124249
3 | 35.0 | 122 | P/FINAL | 2025-05-23 | 10:39 | 2025-05-23 | 09:56 | ELENA | 2 | 34.6 | 34.7 | / | / |  |  |  |  | 1.45 | 209 | 1 | Kinga |  | 3 | 3 | 124249
3 | 35.0 | 123 | P/FINAL | 2025-05-23 | 10:40 | 2025-05-23 | 09:56 | ELENA | 2 | 34.2 | 34.3 | / | / |  |  |  | 1.934 | 0.74 | 199 | 1 | Kinga |  | 3 | 3 | 124249
4 | 60.0 | 124 | P/FINAL | 2025-05-23 | 10:44 | 2025-05-23 | 09:56 | ELENA | 2 | 58.0 | 58.7 | / | / |  |  |  |  | 0.73 | 202 | 1 | Kinga |  | 3 | 4 | 124249
4 | 60.0 | 125 | P/FINAL | 2025-05-23 | 10:44 | 2025-05-23 | 09:56 | ELENA | 2 | 55.2 | 55.5 | / | / |  |  |  | 1.918 | 0.84 | 203 | 1 | Kinga |  | 3 | 4 | 124249
2 | 15.0 | 126 | P/FINAL | 2025-05-23 | 13:45 | 2025-05-23 | 12:42 | ELENA | 2 | 15.2 | 15.2 | / | / |  |  |  |  | 3.43 | 198 | 1 | Kinga |  | 3 | 2 | 124249
2 | 15.0 | 127 | P/FINAL | 2025-05-23 | 13:45 | 2025-05-23 | 12:42 | ELENA | 2 | 15.1 | 15.2 | / | / |  |  |  | 1.882 | 3.38 | 198 | 1 | Kinga |  | 3 | 2 | 124249
3 | 35.0 | 128 | P/FINAL | 2025-05-23 | 13:49 | 2025-05-23 | 12:42 | ELENA | 2 | 34.1 | 34.1 | / | / |  |  |  |  | 0.87 | 203 | 1 | Kinga |  | 3 | 3 | 124249
3 | 35.0 | 129 | P/FINAL | 2025-05-23 | 13:49 | 2025-05-23 | 12:42 | ELENA | 2 | 35.7 | 35.2 | / | / |  |  |  | 1.868 | 0.82 | 221 | 1 | Kinga |  | 3 | 3 | 124249
4 | 60.0 | 130 | P/FINAL | 2025-05-23 | 13:51 | 2025-05-23 | 12:42 | ELENA | 2 | 58.2 | 58.4 | / | / |  |  |  |  | 1.86 | 221 | 1 | Kinga |  | 3 | 4 | 124249
4 | 60.0 | 131 | P/FINAL | 2025-05-23 | 13:52 | 2025-05-23 | 12:42 | ELENA | 2 | 60.1 | 60.3 | / | / |  |  |  | 1.885 | 4.12 | 186 | 1 | Kinga |  | 3 | 4 | 124249
2 | 15.0 | 132 | P/FINAL | 2025-05-23 | 15:03 | 2025-05-23 | 14:42 | ROLAND | 2 | 15.5 | 15.3 | / | / |  |  |  |  | 0.78 | 209 | 1 | Kinga |  | 3 | 2 | 124249
2 | 15.0 | 133 | P/FINAL | 2025-05-23 | 15:04 | 2025-05-23 | 14:42 | ROLAND | 2 | 16.2 | 16.3 | / | / |  |  |  | 1.887 | 3.02 | 205 | 1 | Kinga |  | 3 | 2 | 124249
3 | 35.0 | 134 | P/FINAL | 2025-05-23 | 15:05 | 2025-05-23 | 14:42 | ROLAND | 2 | 36.1 | 36.1 | / | / |  |  |  |  | 0.97 | 190 | 1 | Kinga |  | 3 | 3 | 124249
3 | 35.0 | 135 | P/FINAL | 2025-05-23 | 15:05 | 2025-05-23 | 14:42 | ROLAND | 2 | 36.0 | 36.1 | / | / |  |  |  | 1.869 | 1.00 | 191 | 1 | Kinga |  | 3 | 3 | 124249
4 | 60.0 | 136 | P/FINAL | 2025-05-23 | 15:07 | 2025-05-23 | 14:42 | ROLAND | 2 | 56.1 | 56.7 | / | / |  |  |  |  | 1.18 | 205 | 1 | Kinga |  | 3 | 4 | 124249
4 | 60.0 | 137 | P/FINAL | 2025-05-23 | 15:08 | 2025-05-23 | 14:42 | ROLAND | 2 | 57.4 | 57.5 | / | / |  |  |  |  | 1.52 | 207 | 1 | Kinga |  | 3 | 4 | 124249
2 | 15.0 | 138 | P/FINAL | 2025-05-23 | 15:45 | 2025-05-23 | 15:27 | ROLAND | 2 | 14.8 | 14.9 | / | / |  |  |  |  | 1.20 | 201 | 1 | Kinga |  | 3 | 2 | 124249
2 | 15.0 | 139 | P/FINAL | 2025-05-23 | 15:45 | 2025-05-23 | 15:27 | ROLAND | 2 | 15.0 | 15.1 | / | / |  |  |  |  | 1.21 | 210 | 1 | Kinga |  | 3 | 2 | 124249
3 | 35.0 | 140 | P/FINAL | 2025-05-23 | 15:47 | 2025-05-23 | 15:27 | ROLAND | 2 | 33.2 | 33.6 | / | / |  |  |  |  | 0.71 | 184 | 1 | Kinga |  | 3 | 3 | 124249
3 | 35.0 | 141 | P/FINAL | 2025-05-23 | 15:48 | 2025-05-23 | 15:27 | ROLAND | 2 | 33.3 | 33.4 | / | / |  |  |  |  | 0.66 | 186 | 1 | Kinga |  | 3 | 3 | 124249
4 | 60.0 | 142 | P/FINAL | 2025-05-23 | 15:49 | 2025-05-23 | 15:27 | ROLAND | 2 | 61.7 | 61.8 | / | / |  |  |  |  | 1.10 | 213 | 1 | Kinga |  | 3 | 4 | 124249
4 | 60.0 | 143 | P/FINAL | 2025-05-23 | 15:49 | 2025-05-23 | 15:27 | ROLAND | 2 | 58.8 | 58.9 | / | / |  |  |  |  | 0.61 | 185 | 1 | Kinga |  | 3 | 4 | 124249
2 | 15.0 | 144 | P/FINAL | 2025-05-26 | 08:18 | 2025-05-26 | 07:55 | ROLAND | 2 | 15.1 | 15.1 | / | / |  |  |  | 1.827 | 1.07 | 184 | 1 | Kinga |  | 3 | 2 | 124249
2 | 15.0 | 145 | P/FINAL | 2025-05-26 | 08:19 | 2025-05-26 | 07:55 | ROLAND | 2 | 15.4 | 15.5 | / | / |  |  |  |  | 0.76 | 205 | 1 | Kinga |  | 3 | 2 | 124249
3 | 35.0 | 146 | P/FINAL | 2025-05-26 | 08:19 | 2025-05-26 | 07:55 | ROLAND | 2 | 35.0 | 35.1 | / | / |  |  |  | 1.816 | 1.01 | 188 | 1 | Kinga |  | 3 | 3 | 124249
3 | 35.0 | 147 | P/FINAL | 2025-05-26 | 08:20 | 2025-05-26 | 07:55 | ROLAND | 2 | 36.8 | 36.8 | / | / |  |  |  |  | 1.05 | 195 | 1 | Kinga |  | 3 | 3 | 124249
4 | 60.0 | 148 | P/FINAL | 2025-05-26 | 08:24 | 2025-05-26 | 07:55 | ROLAND | 2 | 59.8 | 60.0 | / | / |  |  |  | 1.816 | 1.57 | 202 | 1 | Kinga |  | 3 | 4 | 124249
4 | 60.0 | 149 | P/FINAL | 2025-05-26 | 08:25 | 2025-05-26 | 07:55 | ROLAND | 2 | 60.6 | 61.3 | / | / |  |  |  |  | 1.03 | 190 | 1 | Kinga |  | 3 | 4 | 124249
2 | 15.0 | 150 | P/FINAL | 2025-05-26 | 09:20 | 2025-05-26 | 09:01 | ROLAND | 2 | 15.6 | 15.7 | / | / |  |  |  |  | 18.2 | 189 | 1 | Kinga |  | 3 | 2 | 124249
2 | 15.0 | 151 | P/FINAL | 2025-05-26 | 09:21 | 2025-05-26 | 09:01 | ROLAND | 2 | 15.7 | 15.8 | / | / |  |  |  | 1.861 | 3.26 | 197 | 1 | Kinga |  | 3 | 2 | 124249
3 | 35.0 | 152 | P/FINAL | 2025-05-26 | 09:22 | 2025-05-26 | 09:01 | ROLAND | 2 | 35.7 | 35.7 | / | / |  |  |  |  | 3.01 | 205 | 1 | Kinga |  | 3 | 3 | 124249
3 | 35.0 | 153 | P/FINAL | 2025-05-26 | 09:23 | 2025-05-26 | 09:01 | ROLAND | 2 | 36.7 | 36.9 | / | / |  |  |  | 1.861 | 2.05 | 200 | 1 | Kinga |  | 3 | 3 | 124249
4 | 60.0 | 154 | P/FINAL | 2025-05-26 | 09:25 | 2025-05-26 | 09:01 | ROLAND | 2 | 58.0 | 58.9 | / | / |  |  |  |  | 0.72 | 215 | 1 | Kinga |  | 3 | 4 | 124249
4 | 60.0 | 155 | P/FINAL | 2025-05-26 | 09:25 | 2025-05-26 | 09:01 | ROLAND | 2 | 57.9 | 57.9 | / | / |  |  |  | 1.871 | 2.07 | 216 | 1 | Kinga |  | 3 | 4 | 124249
2 | 15.0 | 156 | Old A | 2025-05-26 | 10:44 | 2025-05-26 | 11:11 | ROLAND | 2 | 15.5 | 15.2 | / | / |  |  |  |  | 0.71 |  | 1 | Kinga |  | 3 | 2 | 93950
3 | 35.0 | 157 | Old A | 2025-05-26 | 10:45 | 2025-05-26 | 11:11 | ROLAND | 2 | 36.5 | 36.4 | / | / |  |  |  |  | 0.71 |  | 1 | Kinga |  | 3 | 3 | 93950
4 | 60.0 | 158 | Old A | 2025-05-26 | 10:45 | 2025-05-26 | 11:11 | ROLAND | 2 | 62.1 | 61.1 | / | / |  |  |  |  | 0.97 |  | 1 | Kinga |  | 3 | 4 | 93950
2 | 15.0 | 159 | P/FINAL | 2025-05-26 | 11:31 | 2025-05-26 | 11:11 | ROLAND | 2 | 15.7 | 15.7 | / | / |  |  |  | 1.873 | 0.99 | 196 | 1 | Kinga |  | 3 | 2 | 93950
2 | 15.0 | 160 | P/FINAL | 2025-05-26 | 11:32 | 2025-05-26 | 11:11 | ROLAND | 2 | 15.7 | 15.4 | / | / |  |  |  | 1.869 | 1.60 | 180 | 1 | Kinga |  | 3 | 2 | 93950
3 | 35.0 | 161 | P/FINAL | 2025-05-26 | 11:32 | 2025-05-26 | 11:11 | ROLAND | 2 | 35.8 | 35.9 | / | / |  |  |  | 1.890 | 2.95 | 211 | 1 | Kinga |  | 3 | 3 | 93950
3 | 35.0 | 162 | P/FINAL | 2025-05-26 | 11:33 | 2025-05-26 | 11:11 | ROLAND | 2 | 33.7 | 33.6 | / | / |  |  |  | 1.895 | 2.32 | 181 | 1 | Kinga |  | 3 | 3 | 93950
4 | 60.0 | 163 | P/FINAL | 2025-05-26 | 11:45 | 2025-05-26 | 11:11 | ROLAND | 2 | 56.6 | 56.9 | / | / |  |  |  | 1.918 | 1.55 | 211 | 1 | Kinga |  | 3 | 4 | 93950
4 | 60.0 | 164 | P/FINAL | 2025-05-26 | 11:46 | 2025-05-26 | 11:11 | ROLAND | 2 | 56.9 | 57.5 | / | / |  |  |  | 1.887 | 0.77 | 199 | 1 | Kinga |  | 3 | 4 | 93950
2 | 15.0 | 165 | P/FINAL | 2025-05-26 | 12:56 | 2025-05-26 | 12:42 | ROLAND | 2 | 15.6 | 15.6 | / | / |  |  |  | 1.889 | 7.33 | 197 | 1 | Kinga |  | 3 | 2 | 93950
2 | 15.0 | 166 | P/FINAL | 2025-05-26 | 12:56 | 2025-05-26 | 12:42 | ROLAND | 2 | 15.6 | 15.6 | / | / |  |  |  | 1.901 | 1.86 | 195 | 1 | Kinga |  | 3 | 2 | 93950
3 | 35.0 | 167 | P/FINAL | 2025-05-26 | 13:00 | 2025-05-26 | 12:42 | ROLAND | 2 | 34.2 | 34.4 | / | / |  |  |  | 1.909 | 0.99 | 180 | 1 | Kinga |  | 3 | 3 | 93950
3 | 35.0 | 168 | P/FINAL | 2025-05-26 | 13:00 | 2025-05-26 | 12:42 | ROLAND | 2 | 36.2 | 36.2 | / | / |  |  |  | 1.894 | 1.05 | 212 | 1 | Kinga |  | 3 | 3 | 93950
4 | 60.0 | 169 | P/FINAL | 2025-05-26 | 13:02 | 2025-05-26 | 12:42 | ROLAND | 2 | 60.6 | 59.1 | / | / |  |  |  | 1.929 | 0.76 | 202 | 1 | Kinga |  | 3 | 4 | 93950
4 | 60.0 | 170 | P/FINAL | 2025-05-26 | 13:03 | 2025-05-26 | 12:42 | ROLAND | 2 | 55.9 | 56.0 | / | / |  |  |  |  | 0.70 | 192 | 1 | Kinga |  | 3 | 4 | 93950
2 | 15.0 | 171 | P/FINAL | 2025-05-26 | 15:06 | 2025-05-26 | 14:21 | Elena | 2 | 15.1 | 15.1 | / | / |  |  |  | 1.902 | 0.48 | 208 | 1 | Monika |  | 3 | 2 | 93950
2 | 15.0 | 172 | P/FINAL | 2025-05-26 | 15:06 | 2025-05-26 | 14:21 | Elena | 2 | 14.5 | 14.7 | / | / |  |  |  |  | 0.61 | 210 | 1 | Monika |  | 3 | 2 | 93950
3 | 35.0 | 173 | P/FINAL | 2025-05-26 | 15:09 | 2025-05-26 | 14:21 | Elena | 2 | 35.2 | 35.2 | / | / |  |  |  | 1.902 | 0.72 | 211 | 1 | Monika |  | 3 | 3 | 93950
3 | 35.0 | 174 | P/FINAL | 2025-05-26 | 15:11 | 2025-05-26 | 14:21 | Elena | 2 | 37.0 | 37.0 | / | / |  |  |  |  | 0.64 | 209 | 1 | Monika |  | 3 | 3 | 93950
4 | 60.0 | 175 | P/FINAL | 2025-05-26 | 15:12 | 2025-05-26 | 14:21 | Elena | 2 | 59.5 | 59.9 | / | / |  |  |  | 1.900 | 0.81 | 212 | 1 | Monika |  | 3 | 4 | 93950
4 | 60.0 | 176 | P/FINAL | 2025-05-26 | 15:13 | 2025-05-26 | 14:21 | Elena | 2 | 56.6 | 56.5 | / | / |  |  |  |  | 0.59 | 198 | 1 | Monika |  | 3 | 4 | 93950
2 | 15.0 | 177 | P/FINAL | 2025-05-26 | 16:08 | 2025-05-26 | 15:30 | Elena | 2 | 15.1 | 15.2 | / | / |  |  |  | 1.900 | 0.66 | 203 | 1 | Monika |  | 3 | 2 | 93950
2 | 15.0 | 178 | P/FINAL | 2025-05-26 | 16:08 | 2025-05-26 | 15:30 | Elena | 2 | 15.8 | 15.7 | / | / |  |  |  |  | 0.89 | 212 | 1 | Monika |  | 3 | 2 | 93950
3 | 35.0 | 179 | P/FINAL | 2025-05-26 | 16:11 | 2025-05-26 | 15:30 | Elena | 2 | 35.4 | 35.5 | / | / |  |  |  | 1.899 | 0.57 | 200 | 1 | Monika |  | 3 | 3 | 93950
3 | 35.0 | 180 | P/FINAL | 2025-05-26 | 16:12 | 2025-05-26 | 15:30 | Elena | 2 | 37.3 | 37.4 | / | / |  |  |  |  | 0.74 | 209 | 1 | Monika |  | 3 | 3 | 93950
4 | 60.0 | 181 | P/FINAL | 2025-05-26 | 16:17 | 2025-05-26 | 15:30 | Elena | 2 | 57.2 | 56.9 | / | / |  |  |  | 1.907 | 0.66 | 206 | 1 | Monika |  | 3 | 4 | 93950
4 | 60.0 | 182 | P/FINAL | 2025-05-26 | 16:18 | 2025-05-26 | 15:30 | Elena | 2 | 58.9 | 58.8 | / | / |  |  |  |  | 0.59 | 200 | 1 | Monika |  | 3 | 4 | 93950
2 | 15.0 | 183 | P/FINAL | 2025-05-26 | 17:54 | 2025-05-26 | 17:21 | Elena | 2 | 15.6 | 15.6 | / | / |  |  |  | 1.907 | 0.54 | 212 | 1 | Monika |  | 3 | 2 | 93950
2 | 15.0 | 184 | P/FINAL | 2025-05-26 | 17:55 | 2025-05-26 | 17:21 | Elena | 2 | 16.2 | 16.2 | / | / |  |  |  |  | 0.59 | 203 | 1 | Monika |  | 3 | 2 | 93950
3 | 35.0 | 185 | P/FINAL | 2025-05-26 | 17:57 | 2025-05-26 | 17:21 | Elena | 2 | 35.8 | 35.8 | / | / |  |  |  | 1.904 | 0.53 | 205 | 1 | Monika |  | 3 | 3 | 93950
3 | 35.0 | 186 | P/FINAL | 2025-05-26 | 17:58 | 2025-05-26 | 17:21 | Elena | 2 | 37.7 | 37.7 | / | / |  |  |  |  | 0.69 | 212 | 1 | Monika |  | 3 | 3 | 93950
4 | 60.0 | 187 | P/FINAL | 2025-05-26 | 18:01 | 2025-05-26 | 17:21 | Elena | 2 | 57.1 | 57.2 | / | / |  |  |  | 1.908 | 0.67 | 194 | 1 | Monika |  | 3 | 4 | 93950
4 | 60.0 | 188 | P/FINAL | 2025-05-26 | 18:02 | 2025-05-26 | 17:21 | Elena | 2 | 57.9 | 57.9 | / | / |  |  |  |  | 0.60 | 202 | 1 | Monika |  | 3 | 4 | 93950
2 | 15.0 | 189 | P/FINAL | 2025-05-26 | 19:32 | 2025-05-26 | 18:54 | Elena | 2 | 15.7 | 15.8 | / | / |  |  |  | 1.919 | 0.78 | 206 | 1 | Monika |  | 3 | 2 | 93950 | Vacuum 10/10 OK
2 | 15.0 | 190 | P/FINAL | 2025-05-26 | 19:33 | 2025-05-26 | 18:54 | Elena | 2 | 16.3 | 16.3 | / | / |  |  |  |  | 0.52 | 205 | 1 | Monika |  | 3 | 2 | 93950
3 | 35.0 | 191 | P/FINAL | 2025-05-26 | 19:34 | 2025-05-26 | 18:54 | Elena | 2 | 35.2 | 35.2 | / | / |  |  |  | 1.922 | 0.77 | 197 | 1 | Monika |  | 3 | 3 | 93950
4 | 60.0 | 192 | P/FINAL | 2025-05-26 | 19:38 | 2025-05-26 | 18:54 | Elena | 2 | 55.1 | 55.1 | / | / |  |  |  |  | 0.97 | 204 | 1 | Monika |  | 3 | 4 | 93950
4 | 60.0 | 193 | P/FINAL | 2025-05-26 | 19:40 | 2025-05-26 | 18:54 | Elena | 2 | 58.5 | 58.8 | / | / |  |  |  | 1.918 | 0.60 | 209 | 1 | Monika |  | 3 | 4 | 93950
3 | 35.0 | 194 | P/FINAL | 2025-05-26 | 19:41 | 2025-05-26 | 18:54 | Elena | 2 | 37.1 | 37.2 | / | / |  |  |  |  | 0.58 | 213 | 1 | Monika |  | 3 | 3 | 93950
2 | 15.0 | 195 | P/FINAL | 2025-05-26 | 20:56 | 2025-05-26 | 20:10 | Elena | 2 | 16.1 | 16.1 | / | / |  |  |  | 1.918 | 0.66 | 212 | 1 | Monika |  | 3 | 2 | 93950 | Vacuum 10/10 OK
2 | 15.0 | 196 | P/FINAL | 2025-05-26 | 20:57 | 2025-05-26 | 20:10 | Elena | 2 | 16.6 | 16.6 | / | / |  |  |  |  | 0.58 | 204 | 1 | Monika |  | 3 | 2 | 93950
3 | 35.0 | 197 | P/FINAL | 2025-05-26 | 21:01 | 2025-05-26 | 20:10 | Elena | 2 | 35.5 | 35.5 | / | / |  |  |  | 1.912 | 0.68 | 205 | 1 | Monika |  | 3 | 3 | 93950
3 | 35.0 | 198 | P/FINAL | 2025-05-26 | 21:01 | 2025-05-26 | 20:10 | Elena | 2 | 36.3 | 36.4 | / | / |  |  |  |  | 0.55 | 205 | 1 | Monika |  | 3 | 3 | 93950
4 | 60.0 | 199 | P/FINAL | 2025-05-26 | 21:04 | 2025-05-26 | 20:10 | Elena | 2 | 58.2 | 59.2 | / | / |  |  |  | 1.919 | 0.79 | 215 | 1 | Monika |  | 3 | 4 | 93950
4 | 60.0 | 200 | P/FINAL | 2025-05-26 | 21:05 | 2025-05-26 | 20:10 | Elena | 2 | 63.7 | 62.5 | / | / |  |  |  |  | 0.73 | 217 | 1 | Monika |  | 3 | 4 | 93950
1 | 0.0 | 201 | P/FINAL | 2025-05-26 | 21:15 | 2025-05-26 | 20:10 | Elena | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 0.53 | 205 | 1 | Monika |  | 3 | 4 | 93950
Standard # 2
Total Tests | Total Readings | Total Average | Selected Tests | Selected Readings | Selected Average
60 | 120 | 15.3 mg/L | 60 | 120 | 15.3 mg/L
Standard # 3
Total Tests | Total Readings | Total Average | Selected Tests | Selected Readings | Selected Average
60 | 120 | 35.7 mg/L | 60 | 120 | 35.7 mg/L
Standard # 4
Total Tests | Total Readings | Total Average | Selected Tests | Selected Readings | Selected Average
60 | 120 | 59 mg/L | 60 | 120 | 59 mg/L
Standard # 1
Total Tests | Total Readings | Total Average | Selected Tests | Selected Readings | Selected Average
1 | 2 | 0 mg/L | 1 | 2 | 0 mg/L
Specifications
Standard Value [ppm] | Target Value [ppm] | Mean Value [ppm] | Passed
0.0 | 0.0 ÷ 2.0 | 0.0 | YES
15.0 | 12.3 ÷ 17.7 | 15.3 | YES
35.0 | 31.3 ÷ 38.7 | 35.7 | YES
60.0 | 55.1 ÷ 64.9 | 59.0 | YES
Validation Date | by
2025-05-28 | Eugenia