## Hanna Code
Lot | Hanna SFG Code | Recipe | Preparation Week |  | Hanna SFG Code | HI782-0
1103 | HI782-0 | CP-R80 | W 28/1 |  | Parameter Formula | NO3
Description | Exp | Line |  |  | QC Department | Registration Book | QC Type
Marine Nitrate HR Reagent | 2030-07-01 | L57 Powder |  |  | Laboratory | TABLET 3 | Production
Ref Weight. | Min (mg) | Max (mg)
206 | 196 | 216
REAGENT SET 1 | [5 empty] | REAGENT SET 2
Reagent A Lot | Reagent B Lot | Reagent C Lot | Reagent D Lot | Reagent E Lot |  | Reagent A Lot | Reagent B Lot | Reagent C Lot | Reagent D Lot | Reagent E Lot
Expiration A | Expiration B | Expiration C | Expiration D | Expiration E |  | Expiration A | Expiration B | Expiration C | Expiration D | Expiration E
Reagent Range Min | Reagent Range Max
0.0 mg/L | 75.0 mg/L
Prep. Operator | First day Prod. | Last day Prod. | Machine |  |  | Old Lot A | Old Lot B | Old Lot D
ERIKA | 2025-07-08 | 2025-07-15 | 2 HEAD |  |  | 1084
[6 empty] | Expiration A | Expiration B | Expiration D
2030-06-01
Meter
Meter 1 Family | Meter 1 Code | Meter 2 Family | Meter 2 Code
HI97115 | CTO-740 | HI97115 | CTO-730
ph Meter | Description | Turbid. Meter | Description | Spectr. Meter | Description
|  | HI98703 | CTO-861
Standard | mg/L | [5 empty] | Tolerance |  |  |  |  | pH
STD Number | STD Value | STD Min | STD Max |  | STD MR |  | Fixed | AndOr | Percentage | QCRestriction |  | pH Value | pH Min | pH Max
3 | 35.0 | 31.3 | 38.7 |  | MR101 |  | 2.0 | & | 4.9% | 100%
4 | 60.0 | 55.1 | 64.9
2 | 15.0 | 12.3 | 17.7 | [8 empty] | 1.7 | 1.5 | 1.9
1 | 0.0 | 0.0 | 2.0
Readings Table
Standard | STD Value | # | TEST | QC DATE | QC TIME | PROD. DATE | PROD. TIME | PROD. OPERATOR | HEAD | METER 1 | METER 2 | METER 3 | METER 4 | SPECTR. [ABS] | pH 1 | pH 2 | pH 3 | TURB. | WEIGHT [mg] | REAGENT SET | QC OPERATOR | CORRECTION | phNumber | STD | STD_ID | NOTE
3 | 35.0 | 1 | Old A | 2025-07-08 | 09:31 | 2025-07-08 | 08:47 | ERIKA |  | 34.5 | 34.5 | / | / |  |  |  |  | 3.61 |  |  | Kinga |  | 3 | 3 | 81230
3 | 35.0 | 2 | Old A | 2025-07-08 | 09:32 | 2025-07-08 | 08:47 | ERIKA |  | 35.2 | 34.7 | / | / |  |  |  |  | 3.74 |  |  | Kinga |  | 3 | 3 | 81230
4 | 60.0 | 3 | Old A | 2025-07-08 | 09:35 | 2025-07-08 | 08:47 | ERIKA |  | 55.4 | 55.6 | / | / |  |  |  |  | 0.78 |  |  | Kinga |  | 3 | 4 | 81230
4 | 60.0 | 4 | Old A | 2025-07-08 | 09:36 | 2025-07-08 | 08:47 | ERIKA |  | 56.0 | 56.2 | / | / |  |  |  |  | 0.78 |  |  | Kinga |  | 3 | 4 | 81230
2 | 15.0 | 5 | Old A | 2025-07-08 | 09:38 | 2025-07-08 | 08:47 | ERIKA |  | 13.2 | 13.5 | / | / |  |  |  |  | 0.70 |  |  | Kinga |  | 3 | 2 | 81230
2 | 15.0 | 6 | Old A | 2025-07-08 | 09:39 | 2025-07-08 | 08:47 | ERIKA |  | 14.4 | 14.8 | / | / |  |  |  |  | 3.48 |  |  | Kinga |  | 3 | 2 | 81230
3 | 35.0 | 7 | Valid | 2025-07-08 | 09:43 | 2025-07-08 | 08:47 | ERIKA |  | 36.2 | 36.7 | / | / |  |  |  | 1.923 | 1.75 |  |  | Kinga |  | 3 | 3 | 81230
3 | 35.0 | 8 | Valid | 2025-07-08 | 09:44 | 2025-07-08 | 08:47 | ERIKA |  | 33.2 | 33.6 | / | / |  |  |  |  | 1.25 |  |  | Kinga |  | 3 | 3 | 81230
3 | 35.0 | 9 | Valid | 2025-07-08 | 09:46 | 2025-07-08 | 08:47 | ERIKA |  | 34.0 | 33.9 | / | / |  |  |  |  | 4.90 |  |  | Kinga |  | 3 | 3 | 81230
4 | 60.0 | 10 | Valid | 2025-07-08 | 09:50 | 2025-07-08 | 08:47 | ERIKA |  | 60.9 | 60.7 | / | / |  |  |  | 1.926 | 1.30 |  |  | Kinga |  | 3 | 4 | 81230
4 | 60.0 | 11 | Valid | 2025-07-08 | 09:50 | 2025-07-08 | 08:47 | ERIKA |  | 59.0 | 58.1 | / | / |  |  |  |  | 1.50 |  |  | Kinga |  | 3 | 4 | 81230
4 | 60.0 | 12 | Valid | 2025-07-08 | 09:53 | 2025-07-08 | 08:47 | ERIKA |  | 64.2 | 63.9 | / | / |  |  |  |  | 1.59 |  |  | Kinga |  | 3 | 4 | 81230
2 | 15.0 | 13 | Valid | 2025-07-08 | 09:56 | 2025-07-08 | 08:47 | ERIKA |  | 13.5 | 13.7 | / | / |  |  |  |  | 1.00 |  |  | Kinga |  | 3 | 2 | 81230
2 | 15.0 | 14 | Valid | 2025-07-08 | 09:57 | 2025-07-08 | 08:47 | ERIKA |  | 14.2 | 14.7 | / | / |  |  |  | 1.920 | 6.93 |  |  | Kinga |  | 3 | 2 | 81230
2 | 15.0 | 15 | Old A | 2025-07-10 | 11:12 | 2025-07-10 |  | ROLAND | 2 | 16.1 | 16.2 | / | / |  |  |  |  | 1.18 |  |  | Kinga |  | 3 | 2 | 81230
2 | 15.0 | 16 | Old A | 2025-07-10 | 11:12 | 2025-07-10 |  | ROLAND | 2 | 15.7 | 15.8 | / | / |  |  |  |  | 1.18 |  |  | Kinga |  | 3 | 2 | 81230
3 | 35.0 | 17 | Old A | 2025-07-10 | 11:13 | 2025-07-10 |  | ROLAND | 2 | 35.8 | 35.8 | / | / |  |  |  |  | 1.10 |  |  | Kinga |  | 3 | 3 | 81230
3 | 35.0 | 18 | Old A | 2025-07-10 | 11:13 | 2025-07-10 |  | ROLAND | 2 | 36.1 | 36.7 | / | / |  |  |  | 1.778 | 1.27 |  |  | Kinga |  | 3 | 3 | 81230
4 | 60.0 | 19 | Old A | 2025-07-10 | 11:14 | 2025-07-10 |  | ROLAND | 2 | 57.5 | 57.4 | / | / |  |  |  |  | 1.02 |  |  | Kinga |  | 3 | 4 | 81230
4 | 60.0 | 20 | Old A | 2025-07-10 | 11:14 | 2025-07-10 |  | ROLAND | 2 | 55.5 | 55.4 | / | / |  |  |  |  | 1.02 |  |  | Kinga |  | 3 | 4 | 81230
4 | 60.0 | 21 | Old A | 2025-07-10 | 11:14 | 2025-07-10 |  | ROLAND | 2 | 55.7 | 55.6 | / | / |  |  |  |  | 0.92 |  |  | Kinga |  | 3 | 4 | 81230
2 | 15.0 | 22 | P/FINAL | 2025-07-10 | 14:43 | 2025-07-10 | 14:23 | GIORGIANA | 2 | 13.9 | 14.0 | / | / |  |  |  |  | 0.83 | 191 |  | Kinga |  | 3 | 2 | 81230
2 | 15.0 | 23 | P/FINAL | 2025-07-10 | 14:44 | 2025-07-10 | 14:23 | GIORGIANA | 2 | 15.2 | 15.2 | / | / |  |  |  | 1.784 | 5.29 | 201 |  | Kinga |  | 3 | 2 | 81230
2 | 15.0 | 24 | P/FINAL | 2025-07-10 | 14:45 | 2025-07-10 | 14:23 | GIORGIANA | 2 | 13.7 | 13.7 | / | / |  |  |  |  | 0.44 | 183 |  | Kinga |  | 3 | 2 | 81230
3 | 35.0 | 25 | P/FINAL | 2025-07-10 | 14:47 | 2025-07-10 | 14:23 | GIORGIANA | 2 | 35.2 | 34.8 | / | / |  |  |  |  | 0.79 | 193 |  | Kinga |  | 3 | 3 | 81230
3 | 35.0 | 26 | P/FINAL | 2025-07-10 | 14:48 | 2025-07-10 | 14:23 | GIORGIANA | 2 | 32.2 | 32.3 | / | / |  |  |  | 1.790 | 0.62 | 185 |  | Kinga |  | 3 | 3 | 81230
3 | 35.0 | 27 | P/FINAL | 2025-07-10 | 14:49 | 2025-07-10 | 14:23 | GIORGIANA | 2 | 33.3 | 34.1 | / | / |  |  |  |  | 1.02 | 189 |  | Kinga |  | 3 | 3 | 81230
4 | 60.0 | 28 | P/FINAL | 2025-07-10 | 14:59 | 2025-07-10 | 14:23 | GIORGIANA | 2 | 60.3 | 59.5 | / | / |  |  |  |  | 0.83 | 202 |  | Kinga |  | 3 | 4 | 81230
4 | 60.0 | 29 | P/FINAL | 2025-07-10 | 15:00 | 2025-07-10 | 14:23 | GIORGIANA | 2 | 58.8 | 59.8 | / | / |  |  |  |  | 0.89 | 197 |  | Kinga |  | 3 | 4 | 81230
4 | 60.0 | 30 | P/FINAL | 2025-07-10 | 15:00 | 2025-07-10 | 14:23 | GIORGIANA | 2 | 61.2 | 61.1 | / | / |  |  |  | 1.823 | 1.52 | 208 |  | Kinga |  | 3 | 4 | 81230
2 | 15.0 | 31 | P/FINAL | 2025-07-10 | 16:11 | 2025-07-10 | 15:57 | GIORGIANA | 2 | 15.8 | 15.7 | / | / |  |  |  |  | 1.10 | 208 |  | Kinga |  | 3 | 2
2 | 15.0 | 32 | P/FINAL | 2025-07-10 | 16:11 | 2025-07-10 | 15:57 | GIORGIANA | 2 | 15.4 | 15.3 | / | / |  |  |  |  | 2.82 | 213 |  | Kinga |  | 3 | 2
2 | 15.0 | 33 | P/FINAL | 2025-07-10 | 16:12 | 2025-07-10 | 15:57 | GIORGIANA | 2 | 15.9 | 15.8 | / | / |  |  |  | 1.753 | 1.22 | 219 |  | Kinga |  | 3 | 2
3 | 35.0 | 34 | P/FINAL | 2025-07-10 | 16:14 | 2025-07-10 | 15:57 | GIORGIANA | 2 | 34.8 | 34.0 | / | / |  |  |  |  | 0.83 | 211 |  | Kinga |  | 3 | 3
3 | 35.0 | 35 | P/FINAL | 2025-07-10 | 16:14 | 2025-07-10 | 15:57 | GIORGIANA | 2 | 35.3 | 35.6 | / | / |  |  |  |  | 0.87 | 219 |  | Kinga |  | 3 | 3
3 | 35.0 | 36 | P/FINAL | 2025-07-10 | 16:15 | 2025-07-10 | 15:57 | GIORGIANA | 2 | 35.5 | 35.6 | / | / |  |  |  | 1.808 | 0.76 | 206 |  | Kinga |  | 3 | 3
4 | 60.0 | 37 | P/FINAL | 2025-07-10 | 16:17 | 2025-07-10 | 15:57 | GIORGIANA | 2 | 57.1 | 57.2 | / | / |  |  |  |  | 0.94 | 202 |  | Kinga |  | 3 | 4
4 | 60.0 | 38 | P/FINAL | 2025-07-10 | 16:18 | 2025-07-10 | 15:57 | GIORGIANA | 2 | 59.1 | 59.2 | / | / |  |  |  |  | 0.85 | 206 |  | Kinga |  | 3 | 4
4 | 60.0 | 39 | P/FINAL | 2025-07-10 | 16:18 | 2025-07-10 | 15:57 | GIORGIANA | 2 | 60.7 | 62.1 | / | / |  |  |  | 1.782 | 2.28 | 212 |  | Kinga |  | 3 | 4
2 | 15.0 | 40 | P/FINAL | 2025-07-11 | 07:32 | 2025-07-11 | 07:04 | ROLAND | 2 | 13.2 | 13.2 | / | / |  |  |  |  | 6.43 | 199 |  | Kinga |  | 3 | 2 | 81230
2 | 15.0 | 41 | P/FINAL | 2025-07-11 | 07:33 | 2025-07-11 | 07:04 | ROLAND | 2 | 13.9 | 13.9 | / | / |  |  |  | 1.777 | 2.91 | 191 |  | Kinga |  | 3 | 2 | 81230
2 | 15.0 | 42 | P/FINAL | 2025-07-11 | 07:33 | 2025-07-11 | 07:04 | ROLAND | 2 | 14.7 | 14.6 | / | / |  |  |  |  | 0.90 | 210 |  | Kinga |  | 3 | 2 | 81230
3 | 35.0 | 43 | P/FINAL | 2025-07-11 | 07:34 | 2025-07-11 | 07:04 | ROLAND | 2 | 34.0 | 34.0 | / | / |  |  |  |  | 1.88 | 205 |  | Kinga |  | 3 | 3 | 81230
3 | 35.0 | 44 | P/FINAL | 2025-07-11 | 07:35 | 2025-07-11 | 07:04 | ROLAND | 2 | 33.8 | 33.9 | / | / |  |  |  | 1.775 | 2.38 | 211 |  | Kinga |  | 3 | 3 | 81230
3 | 35.0 | 45 | P/FINAL | 2025-07-11 | 07:35 | 2025-07-11 | 07:04 | ROLAND | 2 | 35.7 | 35.7 | / | / |  |  |  |  | 1.02 | 207 |  | Kinga |  | 3 | 3 | 81230
4 | 60.0 | 46 | P/FINAL | 2025-07-11 | 07:39 | 2025-07-11 | 07:04 | ROLAND | 2 | 55.1 | 55.1 | / | / |  |  |  |  | 2.72 | 214 |  | Kinga |  | 3 | 4 | 81230
4 | 60.0 | 47 | P/FINAL | 2025-07-11 | 07:39 | 2025-07-11 | 07:04 | ROLAND | 2 | 57.6 | 57.7 | / | / |  |  |  | 1.803 | 3.89 | 207 |  | Kinga |  | 3 | 4 | 81230
2 | 15.0 | 48 | P/FINAL | 2025-07-11 | 07:43 | 2025-07-10 | 17:30 | GIORGIANA | 2 | 15.2 | 15.0 | / | / |  |  |  |  | 6.30 | 213 |  | Kinga |  | 3 | 2 | 81230
2 | 15.0 | 49 | P/FINAL | 2025-07-11 | 07:43 | 2025-07-10 | 17:30 | GIORGIANA | 2 | 14.8 | 14.4 | / | / |  |  |  |  | 1.69 | 214 |  | Kinga |  | 3 | 2 | 81230
3 | 35.0 | 50 | P/FINAL | 2025-07-11 | 07:44 | 2025-07-10 | 17:30 | GIORGIANA | 2 | 34.8 | 34.9 | / | / |  |  |  |  | 1.95 | 204 |  | Kinga |  | 3 | 3 | 81230
3 | 35.0 | 51 | P/FINAL | 2025-07-11 | 07:44 | 2025-07-10 | 17:30 | GIORGIANA | 2 | 34.2 | 34.3 | / | / |  |  |  | 1.804 | 1.82 | 207 |  | Kinga |  | 3 | 3 | 81230
4 | 60.0 | 52 | P/FINAL | 2025-07-11 | 07:44 | 2025-07-10 | 17:30 | GIORGIANA | 2 | 58.4 | 57.9 | / | / |  |  |  |  | 2.51 | 211 |  | Kinga |  | 3 | 4 | 81230
4 | 60.0 | 53 | P/FINAL | 2025-07-11 | 07:45 | 2025-07-10 | 17:30 | GIORGIANA | 2 | 58.1 | 58.3 | / | / |  |  |  |  | 3.10 | 214 |  | Kinga |  | 3 | 4 | 81230
2 | 15.0 | 54 | P/FINAL | 2025-07-11 | 07:51 | 2025-07-10 | 19:00 | GIORGIANA | 2 | 13.8 | 13.8 | / | / |  |  |  | 1.786 | 0.78 | 215 |  | Kinga |  | 3 | 2 | 81230
2 | 15.0 | 55 | P/FINAL | 2025-07-11 | 07:52 | 2025-07-10 | 19:00 | GIORGIANA | 2 | 14.2 | 14.3 | / | / |  |  |  |  | 4.47 | 203 |  | Kinga |  | 3 | 2 | 81230
3 | 35.0 | 56 | P/FINAL | 2025-07-11 | 07:53 | 2025-07-10 | 19:00 | GIORGIANA | 2 | 36.0 | 36.1 | / | / |  |  |  |  | 0.88 | 205 |  | Kinga |  | 3 | 3 | 81230
3 | 35.0 | 57 | P/FINAL | 2025-07-11 | 07:53 | 2025-07-10 | 19:00 | GIORGIANA | 2 | 35.9 | 35.8 | / | / |  |  |  |  | 1.61 | 212 |  | Kinga |  | 3 | 3 | 81230
4 | 60.0 | 58 | P/FINAL | 2025-07-11 | 07:54 | 2025-07-10 | 19:00 | GIORGIANA | 2 | 55.2 | 55.1 | / | / |  |  |  | 1.801 | 2.70 | 218 |  | Kinga |  | 3 | 4 | 81230
4 | 60.0 | 59 | P/FINAL | 2025-07-11 | 07:55 | 2025-07-10 | 19:00 | GIORGIANA | 2 | 55.2 | 55.5 | / | / |  |  |  |  | 3.29 | 217 |  | Kinga |  | 3 | 4 | 81230
2 | 15.0 | 60 | P/FINAL | 2025-07-11 | 08:02 | 2025-07-10 | 22:20 | GIORGIANA | 2 | 14.2 | 14.1 | / | / |  |  |  |  | 3.26 | 201 |  | Kinga |  | 3 | 2
2 | 15.0 | 61 | P/FINAL | 2025-07-11 | 08:03 | 2025-07-10 | 22:20 | GIORGIANA | 2 | 14.0 | 13.9 | / | / |  |  |  |  | 6.73 | 208 |  | Kinga |  | 3 | 2
3 | 35.0 | 62 | P/FINAL | 2025-07-11 | 08:04 | 2025-07-10 | 22:20 | GIORGIANA | 2 | 35.1 | 35.0 | / | / |  |  |  | 1.798 | 1.58 | 204 |  | Kinga |  | 3 | 3
3 | 35.0 | 63 | P/FINAL | 2025-07-11 | 08:04 | 2025-07-10 | 22:20 | GIORGIANA | 2 | 34.4 | 34.9 | / | / |  |  |  |  | 2.46 | 212 |  | Kinga |  | 3 | 3
4 | 60.0 | 64 | P/FINAL | 2025-07-11 | 08:05 | 2025-07-10 | 22:20 | GIORGIANA | 2 | 58.2 | 58.3 | / | / |  |  |  |  | 2.88 | 200 |  | Kinga |  | 3 | 4
4 | 60.0 | 65 | P/FINAL | 2025-07-11 | 08:06 | 2025-07-10 | 22:20 | GIORGIANA | 2 | 58.2 | 58.3 | / | / |  |  |  | 1.762 | 18.4 | 202 |  | Kinga |  | 3 | 4
2 | 15.0 | 66 | P/FINAL | 2025-07-11 | 08:31 | 2025-07-10 | 21:15 | GIORGIANA | 2 | 13.3 | 13.4 | / | / |  |  |  |  | 0.84 | 213 |  | Kinga |  | 3 | 2
2 | 15.0 | 67 | P/FINAL | 2025-07-11 | 08:31 | 2025-07-10 | 21:15 | GIORGIANA | 2 | 13.7 | 14.0 | / | / |  |  |  | 1.758 | 0.90 | 210 |  | Kinga |  | 3 | 2
2 | 15.0 | 68 | P/FINAL | 2025-07-11 | 08:32 | 2025-07-10 | 21:15 | GIORGIANA | 2 | 15.2 | 15.2 | / | / |  |  |  |  | 1.53 | 183 |  | Kinga |  | 3 | 2
3 | 35.0 | 69 | P/FINAL | 2025-07-11 | 08:36 | 2025-07-10 | 21:15 | GIORGIANA | 2 | 35.3 | 34.6 | / | / |  |  |  |  | 1.86 | 202 |  | Kinga |  | 3 | 3
3 | 35.0 | 70 | P/FINAL | 2025-07-11 | 08:37 | 2025-07-10 | 21:15 | GIORGIANA | 2 | 34.9 | 35.3 | / | / |  |  |  | 1.769 | 1.21 | 211 |  | Kinga |  | 3 | 3
3 | 35.0 | 71 | P/FINAL | 2025-07-11 | 08:38 | 2025-07-10 | 21:15 | GIORGIANA | 2 | 35.2 | 35.3 | / | / |  |  |  |  | 2.62 | 209 |  | Kinga |  | 3 | 3
4 | 60.0 | 72 | P/FINAL | 2025-07-11 | 08:49 | 2025-07-10 | 21:15 | GIORGIANA | 2 | 55.1 | 56.0 | / | / |  |  |  |  | 1.33 | 208 |  | Kinga |  | 3 | 4
4 | 60.0 | 73 | P/FINAL | 2025-07-11 | 08:49 | 2025-07-10 | 21:15 | GIORGIANA | 2 | 57.5 | 58.0 | / | / |  |  |  |  | 1.55 | 204 |  | Kinga |  | 3 | 4
4 | 60.0 | 74 | P/FINAL | 2025-07-11 | 08:50 | 2025-07-10 | 21:15 | GIORGIANA | 2 | 56.7 | 57.0 | / | / |  |  |  | 1.751 | 1.29 | 209 |  | Kinga |  | 3 | 4
2 | 15.0 | 75 | P/FINAL | 2025-07-11 | 09:07 | 2025-07-11 | 08:41 | ROLAND | 2 | 12.6 | 12.7 | / | / |  |  |  |  | 2.26 | 185 |  | Kinga |  | 3 | 2
2 | 15.0 | 76 | P/FINAL | 2025-07-11 | 09:08 | 2025-07-11 | 08:41 | ROLAND | 2 | 12.8 | 12.9 | / | / |  |  |  |  | 0.52 | 188 |  | Kinga |  | 3 | 2
2 | 15.0 | 77 | P/FINAL | 2025-07-11 | 09:09 | 2025-07-11 | 08:41 | ROLAND | 2 | 12.8 | 12.4 | / | / |  |  |  |  | 1.17 | 186 |  | Kinga |  | 3 | 2
3 | 35.0 | 78 | P/FINAL | 2025-07-11 | 09:20 | 2025-07-11 | 08:41 | ROLAND | 2 | 31.4 | 31.4 | / | / |  |  |  | 1.765 | 3.93 | 209 |  | Kinga |  | 3 | 3
3 | 35.0 | 79 | P/FINAL | 2025-07-11 | 09:22 | 2025-07-11 | 08:41 | ROLAND | 2 | 32.3 | 32.2 | / | / |  |  |  |  | 4.39 | 212 |  | Kinga |  | 3 | 3
3 | 35.0 | 80 | P/FINAL | 2025-07-11 | 09:23 | 2025-07-11 | 08:41 | ROLAND | 2 | 33.6 | 33.4 | / | / |  |  |  |  | 3.39 | 213 |  | Kinga |  | 3 | 3
4 | 60.0 | 81 | P/FINAL | 2025-07-11 | 09:26 | 2025-07-11 | 08:41 | ROLAND | 2 | 57.6 | 57.7 | / | / |  |  |  | 1.800 | 2.17 | 205 |  | Kinga |  | 3 | 4
4 | 60.0 | 82 | P/FINAL | 2025-07-11 | 09:26 | 2025-07-11 | 08:41 | ROLAND | 2 | 59.1 | 59.3 | / | / |  |  |  |  | 1.47 | 204 |  | Kinga |  | 3 | 4
4 | 60.0 | 83 | P/FINAL | 2025-07-11 | 09:28 | 2025-07-11 | 08:41 | ROLAND | 2 | 60.3 | 60.2 | / | / |  |  |  |  | 1.59 | 217 |  | Kinga |  | 3 | 4
1 | 0.0 | 84 | P/FINAL | 2025-07-11 | 09:38 | 2025-07-11 | 08:41 | ROLAND | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 1.77 | 217 |  | Kinga |  | 3 | 1
1 | 0.0 | 85 | P/FINAL | 2025-07-11 | 09:39 | 2025-07-11 | 08:41 | ROLAND | 2 | 0.0 | 0.0 | / | / |  |  |  | 1.688 | 2.14 | 202 |  | Kinga |  | 3 | 1
2 | 15.0 | 86 | Old A | 2025-07-11 | 10:49 | 2025-07-11 | 10:40 | ROLAND | 2 | 14.5 | 14.4 | / | / |  |  |  |  | 4.42 | 203 |  | Kinga |  | 3 | 2
2 | 15.0 | 87 | Old A | 2025-07-11 | 10:50 | 2025-07-11 | 10:40 | ROLAND | 2 | 15.2 | 15.2 | / | / |  |  |  |  | 0.90 | 200 |  | Kinga |  | 3 | 2
3 | 35.0 | 88 | Old A | 2025-07-11 | 10:50 | 2025-07-11 | 10:40 | ROLAND | 2 | 36.0 | 36.8 | / | / |  |  |  |  | 0.83 | 204 |  | Kinga |  | 3 | 3
3 | 35.0 | 89 | Old A | 2025-07-11 | 10:51 | 2025-07-11 | 10:40 | ROLAND | 2 | 37.3 | 35.9 | / | / |  |  |  |  | 1.38 | 201 |  | Kinga |  | 3 | 3
4 | 60.0 | 90 | Old A | 2025-07-11 | 10:53 | 2025-07-11 | 10:40 | ROLAND | 2 | 60.8 | 59.9 | / | / | [5 empty] | 216 |  | Kinga |  | 3 | 4
4 | 60.0 | 91 | Old A | 2025-07-11 | 10:53 | 2025-07-11 | 10:40 | ROLAND | 2 | 58.6 | 58.3 | / | / |  |  |  |  | 2.57 |  |  | Kinga |  | 3 | 4
2 | 15.0 | 92 | P/FINAL | 2025-07-11 | 11:08 | 2025-07-11 | 10:47 | ROLAND | 2 | 14.9 | 15.0 | / | / |  |  |  | 1.730 | 0.76 | 197 |  | Kinga |  | 3 | 2
2 | 15.0 | 93 | P/FINAL | 2025-07-11 | 11:09 | 2025-07-11 | 10:47 | ROLAND | 2 | 14.3 | 14.0 | / | / |  |  |  |  | 2.28 | 201 |  | Kinga |  | 3 | 2
2 | 15.0 | 94 | P/FINAL | 2025-07-11 | 11:09 | 2025-07-11 | 10:47 | ROLAND | 2 | 16.1 | 16.1 | / | / |  |  |  |  | 0.99 | 202 |  | Kinga |  | 3 | 2
3 | 35.0 | 95 | P/FINAL | 2025-07-11 | 11:11 | 2025-07-11 | 10:47 | ROLAND | 2 | 36.6 | 36.5 | / | / |  |  |  | 1.667 | 1.21 | 208 |  | Kinga |  | 3 | 3
3 | 35.0 | 96 | P/FINAL | 2025-07-11 | 11:12 | 2025-07-11 | 10:47 | ROLAND | 2 | 34.0 | 34.1 | / | / |  |  |  |  | 0.56 | 198 |  | Kinga |  | 3 | 3
3 | 35.0 | 97 | P/FINAL | 2025-07-11 | 11:13 | 2025-07-11 | 10:47 | ROLAND | 2 | 34.2 | 34.3 | / | / |  |  |  |  | 1.79 | 205 |  | Kinga |  | 3 | 3
4 | 60.0 | 98 | P/FINAL | 2025-07-11 | 11:14 | 2025-07-11 | 10:47 | ROLAND | 2 | 58.8 | 58.4 | / | / |  |  |  |  | 0.97 | 211 |  | Kinga |  | 3 | 4
4 | 60.0 | 99 | P/FINAL | 2025-07-11 | 11:15 | 2025-07-11 | 10:47 | ROLAND | 2 | 56.8 | 56.5 | / | / |  |  |  | 1.714 | 1.21 | 218 |  | Kinga |  | 3 | 4
4 | 60.0 | 100 | P/FINAL | 2025-07-11 | 11:15 | 2025-07-11 | 10:47 | ROLAND | 2 | 55.7 | 55.9 | / | / |  |  |  |  | 0.86 | 203 |  | Kinga |  | 3 | 4
2 | 15.0 | 101 | P/FINAL | 2025-07-11 | 12:38 | 2025-07-11 | 12:15 | ROLAND | 2 | 14.4 | 14.7 | / | / |  |  |  | 1.757 | 1.71 | 205 |  | Kinga |  | 3 | 2 | 90220
2 | 15.0 | 102 | P/FINAL | 2025-07-11 | 12:38 | 2025-07-11 | 12:15 | ROLAND | 2 | 14.3 | 14.2 | / | / |  |  |  |  | 2.45 | 202 |  | Kinga |  | 3 | 2 | 90220
2 | 15.0 | 103 | P/FINAL | 2025-07-11 | 12:39 | 2025-07-11 | 12:15 | ROLAND | 2 | 15.8 | 15.7 | / | / |  |  |  |  | 2.90 | 200 |  | Kinga |  | 3 | 2 | 90220
3 | 35.0 | 104 | P/FINAL | 2025-07-11 | 12:41 | 2025-07-11 | 12:15 | ROLAND | 2 | 32.1 | 32.2 | / | / |  |  |  |  | 2.46 | 206 |  | Kinga |  | 3 | 3 | 90220
3 | 35.0 | 105 | P/FINAL | 2025-07-11 | 12:41 | 2025-07-11 | 12:15 | ROLAND | 2 | 37.1 | 37.3 | / | / |  |  |  | 1.759 | 0.79 | 197 |  | Kinga |  | 3 | 3 | 90220
3 | 35.0 | 106 | P/FINAL | 2025-07-11 | 12:42 | 2025-07-11 | 12:15 | ROLAND | 2 | 35.4 | 35.5 | / | / |  |  |  |  | 1.33 | 199 |  | Kinga |  | 3 | 3 | 90220
4 | 60.0 | 107 | P/FINAL | 2025-07-11 | 12:44 | 2025-07-11 | 12:15 | ROLAND | 2 | 57.3 | 57.7 | / | / |  |  |  |  | 3.58 | 208 |  | Kinga |  | 3 | 4 | 90220
4 | 60.0 | 108 | P/FINAL | 2025-07-11 | 12:45 | 2025-07-11 | 12:15 | ROLAND | 2 | 58.7 | 60.9 | / | / |  |  |  | 1.731 | 0.73 | 200 |  | Kinga |  | 3 | 4 | 90220
4 | 60.0 | 109 | P/FINAL | 2025-07-11 | 12:45 | 2025-07-11 | 12:15 | ROLAND | 2 | 58.7 | 58.1 | / | / |  |  |  |  | 1.21 | 200 |  | Kinga |  | 3 | 4 | 90220
2 | 15.0 | 110 | P/FINAL | 2025-07-11 | 15:04 | 2025-07-11 | 14:35 | GIORGIANA | 2 | 15.4 | 15.5 | / | / |  |  |  |  | 0.76 | 216 |  | Kinga |  | 3 | 2 | 90220
2 | 15.0 | 111 | P/FINAL | 2025-07-11 | 15:05 | 2025-07-11 | 14:35 | GIORGIANA | 2 | 14.8 | 14.5 | / | / |  |  |  |  | 1.01 | 198 |  | Kinga |  | 3 | 2 | 90220
2 | 15.0 | 112 | P/FINAL | 2025-07-11 | 15:06 | 2025-07-11 | 14:35 | GIORGIANA | 2 | 16.0 | 15.9 | / | / |  |  |  | 1.715 | 0.64 | 195 |  | Kinga |  | 3 | 2 | 90220
3 | 35.0 | 113 | P/FINAL | 2025-07-11 | 15:08 | 2025-07-11 | 14:35 | GIORGIANA | 2 | 37.6 | 37.8 | / | / |  |  |  |  | 0.58 | 212 |  | Kinga |  | 3 | 3 | 90220
3 | 35.0 | 114 | P/FINAL | 2025-07-11 | 15:08 | 2025-07-11 | 14:35 | GIORGIANA | 2 | 37.4 | 37.5 | / | / |  |  |  |  | 0.82 | 186 |  | Kinga |  | 3 | 3 | 90220
3 | 35.0 | 115 | P/FINAL | 2025-07-11 | 15:09 | 2025-07-11 | 14:35 | GIORGIANA | 2 | 37.0 | 36.6 | / | / |  |  |  | 1.745 | 0.71 | 210 |  | Kinga |  | 3 | 3 | 90220
4 | 60.0 | 116 | P/FINAL | 2025-07-11 | 15:11 | 2025-07-11 | 14:35 | GIORGIANA | 2 | 57.8 | 57.8 | / | / |  |  |  |  | 0.77 | 201 |  | Kinga |  | 3 | 4 | 90220
4 | 60.0 | 117 | P/FINAL | 2025-07-11 | 15:13 | 2025-07-11 | 14:35 | GIORGIANA | 2 | 55.5 | 55.6 | / | / |  |  |  | 1.744 | 0.67 | 179 |  | Kinga |  | 3 | 4 | 90220
4 | 60.0 | 118 | P/FINAL | 2025-07-11 | 15:13 | 2025-07-11 | 14:35 | GIORGIANA | 2 | 62.3 | 61.9 | / | / |  |  |  |  | 0.67 | 190 |  | Kinga |  | 3 | 4 | 90220
2 | 15.0 | 119 | P/FINAL | 2025-07-11 | 15:54 | 2025-07-11 | 15:30 | GIORGIANA | 2 | 14.1 | 14.1 | / | / |  |  |  | 1.686 | 0.60 | 205 |  | Andrea V |  | 3 | 2
2 | 15.0 | 120 | P/FINAL | 2025-07-11 | 15:55 | 2025-07-11 | 15:30 | GIORGIANA | 2 | 15.8 | 15.7 | / | / |  |  |  |  | 0.64 | 196 |  | Andrea V |  | 3 | 2
2 | 15.0 | 121 | P/FINAL | 2025-07-11 | 15:56 | 2025-07-11 | 15:30 | GIORGIANA | 2 | 14.6 | 14.5 | / | / |  |  |  |  | 0.50 | 170 |  | Andrea V |  | 3 | 2
3 | 35.0 | 122 | P/FINAL | 2025-07-11 | 15:58 | 2025-07-11 | 15:30 | GIORGIANA | 2 | 35.8 | 36.0 | / | / |  |  |  |  | 0.58 | 208 |  | Andrea V |  | 3 | 3
3 | 35.0 | 123 | P/FINAL | 2025-07-11 | 15:59 | 2025-07-11 | 15:30 | GIORGIANA | 2 | 37.8 | 37.7 | / | / |  |  |  | 1.689 | 0.60 | 201 |  | Andrea V |  | 3 | 3
3 | 35.0 | 124 | P/FINAL | 2025-07-11 | 16:00 | 2025-07-11 | 15:30 | GIORGIANA | 2 | 38.1 | 38.0 | / | / |  |  |  |  | 0.74 | 201 |  | Andrea V |  | 3 | 3
4 | 60.0 | 125 | P/FINAL | 2025-07-11 | 16:02 | 2025-07-11 | 15:30 | GIORGIANA | 2 | 55.4 | 55.5 | / | / |  |  |  |  | 0.56 | 197 |  | Andrea V |  | 3 | 4
4 | 60.0 | 126 | P/FINAL | 2025-07-11 | 16:02 | 2025-07-11 | 15:30 | GIORGIANA | 2 | 56.3 | 56.3 | / | / |  |  |  |  | 0.88 | 190 |  | Andrea V |  | 3 | 4
4 | 60.0 | 127 | P/FINAL | 2025-07-11 | 16:03 | 2025-07-11 | 15:30 | GIORGIANA | 2 | 55.2 | 55.7 | / | / |  |  |  | 1.722 | 0.61 | 189 |  | Andrea V |  | 3 | 4
1 | 0.0 | 128 | P/FINAL | 2025-07-11 | 16:06 | 2025-07-11 | 15:30 | GIORGIANA | 2 | 0.0 | 0.0 | / | / |  |  |  | 1.831 | 0.44 | 221 |  | Andrea V |  | 3 | 1
1 | 0.0 | 129 | P/FINAL | 2025-07-11 | 16:07 | 2025-07-11 | 15:30 | GIORGIANA | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 0.42 |  |  | Andrea V |  | 3 | 1
2 | 15.0 | 130 | P/FINAL | 2025-07-12 | 12:29 | 2025-07-11 | 17:33 | Georgiana | 2 | 14.2 | 14.2 | / | / |  |  |  | 1.762 | 0.63 | 200 |  | Monika |  | 3 | 2 | 90220
2 | 15.0 | 131 | P/FINAL | 2025-07-12 | 12:31 | 2025-07-11 | 17:33 | Georgiana | 2 | 15.2 | 15.2 | / | / |  |  |  |  | 0.58 | 207 |  | Monika |  | 3 | 2 | 90220
2 | 15.0 | 132 | P/FINAL | 2025-07-12 | 12:33 | 2025-07-11 | 17:33 | Georgiana | 2 | 14.9 | 14.9 | / | / |  |  |  |  | 0.62 | 200 |  | Monika |  | 3 | 2 | 90220
3 | 35.0 | 133 | P/FINAL | 2025-07-12 | 12:34 | 2025-07-11 | 17:33 | Georgiana | 2 | 36.6 | 36.6 | / | / |  |  |  | 1.742 | 0.72 | 212 |  | Monika |  | 3 | 3 | 90220
3 | 35.0 | 134 | P/FINAL | 2025-07-12 | 12:38 | 2025-07-11 | 17:33 | Georgiana | 2 | 32.0 | 32.2 | / | / |  |  |  |  | 0.63 | 197 |  | Monika |  | 3 | 3 | 90220
3 | 35.0 | 135 | P/FINAL | 2025-07-12 | 12:39 | 2025-07-11 | 17:33 | Georgiana | 2 | 34.7 | 34.8 | / | / |  |  |  |  | 0.84 | 202 |  | Monika |  | 3 | 3 | 90220
4 | 60.0 | 136 | P/FINAL | 2025-07-12 | 12:40 | 2025-07-11 | 17:33 | Georgiana | 2 | 56.4 | 56.3 | / | / |  |  |  | 1.763 | 0.73 | 200 |  | Monika |  | 3 | 4 | 90220
4 | 60.0 | 137 | P/FINAL | 2025-07-12 | 12:41 | 2025-07-11 | 17:33 | Georgiana | 2 | 59.3 | 59.4 | / | / |  |  |  |  | 0.71 | 203 |  | Monika |  | 3 | 4 | 90220
4 | 60.0 | 138 | P/FINAL | 2025-07-12 | 12:49 | 2025-07-11 | 17:33 | Georgiana | 2 | 60.0 | 60.2 | / | / |  |  |  |  | 0.80 | 208 |  | Monika |  | 3 | 4 | 90220
1 | 0.0 | 139 | P/FINAL | 2025-07-12 | 12:50 | 2025-07-11 | 17:33 | Georgiana | 2 | 0.0 | 0.0 | / | / |  |  |  | 1.740 | 0.85 | 201 |  | Monika |  | 3 | 1 |  | SW4
2 | 15.0 | 140 | P/FINAL | 2025-07-14 | 08:22 | 2025-07-14 | 06:42 | GIORGIANA | 2 | 13.6 | 13.7 | / | / |  |  |  |  | 0.58 | 199 |  | Kinga |  | 3 | 2 | 90220
2 | 15.0 | 141 | P/FINAL | 2025-07-14 | 08:23 | 2025-07-14 | 06:42 | GIORGIANA | 2 | 15.2 | 15.1 | / | / |  |  |  |  | 2.12 | 201 |  | Kinga |  | 3 | 2 | 90220
2 | 15.0 | 142 | P/FINAL | 2025-07-14 | 08:24 | 2025-07-14 | 06:42 | GIORGIANA | 2 | 14.8 | 14.9 | / | / |  |  |  | 1.784 | 1.18 | 194 |  | Kinga |  | 3 | 2 | 90220
3 | 35.0 | 143 | P/FINAL | 2025-07-14 | 08:35 | 2025-07-14 | 06:42 | GIORGIANA | 2 | 34.8 | 34.9 | / | / |  |  |  |  | 1.29 | 202 |  | Kinga |  | 3 | 3 | 90220
3 | 35.0 | 144 | P/FINAL | 2025-07-14 | 08:36 | 2025-07-14 | 06:42 | GIORGIANA | 2 | 34.9 | 35.1 | / | / |  |  |  |  | 1.64 | 202 |  | Kinga |  | 3 | 3 | 90220
3 | 35.0 | 145 | P/FINAL | 2025-07-14 | 08:37 | 2025-07-14 | 06:42 | GIORGIANA | 2 | 37.3 | 37.5 | / | / |  |  |  | 1.768 | 1.16 | 196 |  | Kinga |  | 3 | 3 | 90220
4 | 60.0 | 146 | P/FINAL | 2025-07-14 | 08:41 | 2025-07-14 | 06:42 | GIORGIANA | 2 | 55.2 | 55.7 | / | / |  |  |  |  | 2.99 | 213 |  | Kinga |  | 3 | 4 | 90220
4 | 60.0 | 147 | P/FINAL | 2025-07-14 | 08:41 | 2025-07-14 | 06:42 | GIORGIANA | 2 | 56.4 | 56.3 | / | / |  |  |  | 1.770 | 3.89 | 207 |  | Kinga |  | 3 | 4 | 90220
4 | 60.0 | 148 | P/FINAL | 2025-07-14 | 08:42 | 2025-07-14 | 06:42 | GIORGIANA | 2 | 56.0 | 56.0 | / | / |  |  |  |  | 2.73 | 198 |  | Kinga |  | 3 | 4 | 90220
2 | 15.0 | 149 | P/FINAL | 2025-07-14 | 10:22 | 2025-07-14 | 09:41 | GIORGIANA | 2 | 15.2 | 15.0 | / | / |  |  |  |  | 1.59 | 185 |  | Kinga |  | 3 | 2 | 90220
2 | 15.0 | 150 | P/FINAL | 2025-07-14 | 10:22 | 2025-07-14 | 09:41 | GIORGIANA | 2 | 14.4 | 14.4 | / | / |  |  |  | 1.775 | 1.06 | 208 |  | Kinga |  | 3 | 2 | 90220
2 | 15.0 | 151 | P/FINAL | 2025-07-14 | 10:24 | 2025-07-14 | 09:41 | GIORGIANA | 2 | 13.5 | 13.4 | / | / |  |  |  |  | 1.34 | 196 |  | Kinga |  | 3 | 2 | 90220
3 | 35.0 | 152 | P/FINAL | 2025-07-14 | 10:28 | 2025-07-14 | 09:41 | GIORGIANA | 2 | 35.6 | 35.7 | / | / |  |  |  |  | 2.43 | 207 |  | Kinga |  | 3 | 3 | 90220
3 | 35.0 | 153 | P/FINAL | 2025-07-14 | 10:28 | 2025-07-14 | 09:41 | GIORGIANA | 2 | 34.6 | 34.6 | / | / |  |  |  | 1.806 | 0.81 | 197 |  | Kinga |  | 3 | 3 | 90220
3 | 35.0 | 154 | P/FINAL | 2025-07-14 | 10:32 | 2025-07-14 | 09:41 | GIORGIANA | 2 | 36.1 | 35.9 | / | / |  |  |  |  | 0.75 | 190 |  | Kinga |  | 3 | 3 | 90220
4 | 60.0 | 155 | P/FINAL | 2025-07-14 | 10:55 | 2025-07-14 | 09:41 | GIORGIANA | 2 | 58.1 | 58.6 | / | / |  |  |  |  | 1.53 | 209 |  | Kinga |  | 3 | 4 | 90220
4 | 60.0 | 156 | P/FINAL | 2025-07-14 | 10:56 | 2025-07-14 | 09:41 | GIORGIANA | 2 | 59.4 | 59.6 | / | / |  |  |  | 1.798 | 1.27 | 207 |  | Kinga |  | 3 | 4 | 90220
4 | 60.0 | 157 | P/FINAL | 2025-07-14 | 10:58 | 2025-07-14 | 09:41 | GIORGIANA | 2 | 61.7 | 61.7 | / | / |  |  |  | 1.798 | 2.65 | 201 |  | Kinga |  | 3 | 4 | 90220
2 | 15.0 | 158 | P/FINAL | 2025-07-14 | 13:44 | 2025-07-14 | 12:51 | GIORGIANA | 2 | 14.3 | 14.3 | / | / |  |  |  |  | 1.17 | 190 |  | Kinga |  | 3 | 2 | 83024
2 | 15.0 | 159 | P/FINAL | 2025-07-14 | 13:45 | 2025-07-14 | 12:51 | GIORGIANA | 2 | 14.7 | 14.9 | / | / |  |  |  |  | 1.79 | 200 |  | Kinga |  | 3 | 2 | 83024
2 | 15.0 | 160 | P/FINAL | 2025-07-14 | 13:46 | 2025-07-14 | 12:51 | GIORGIANA | 2 | 13.4 | 13.6 | / | / |  |  |  |  | 4.56 | 186 |  | Kinga |  | 3 | 2 | 83024
3 | 35.0 | 161 | P/FINAL | 2025-07-14 | 13:50 | 2025-07-14 | 12:51 | GIORGIANA | 2 | 35.8 | 35.5 | / | / |  |  |  | 1.790 | 1.96 | 182 |  | Kinga |  | 3 | 3 | 83024
3 | 35.0 | 162 | P/FINAL | 2025-07-14 | 13:51 | 2025-07-14 | 12:51 | GIORGIANA | 2 | 37.2 | 37.1 | / | / |  |  |  |  | 0.68 | 190 |  | Kinga |  | 3 | 3 | 83024
3 | 35.0 | 163 | P/FINAL | 2025-07-14 | 13:54 | 2025-07-14 | 12:51 | GIORGIANA | 2 | 36.1 | 36.2 | / | / |  |  |  |  | 3.00 | 204 |  | Kinga |  | 3 | 3 | 83024
4 | 60.0 | 164 | P/FINAL | 2025-07-14 | 14:04 | 2025-07-14 | 12:51 | GIORGIANA | 2 | 58.9 | 58.2 | / | / |  |  |  | 1.787 | 2.41 | 195 |  | Kinga |  | 3 | 4 | 83024
4 | 60.0 | 165 | P/FINAL | 2025-07-14 | 14:05 | 2025-07-14 | 12:51 | GIORGIANA | 2 | 58.2 | 58.5 | / | / |  |  |  |  | 1.43 | 194 |  | Kinga |  | 3 | 4 | 83024
4 | 60.0 | 166 | P/FINAL | 2025-07-14 | 14:08 | 2025-07-14 | 12:51 | GIORGIANA | 2 | 56.0 | 55.2 | / | / |  |  |  |  | 1.91 | 191 |  | Kinga |  | 3 | 4 | 83024
2 | 15.0 | 167 | P/FINAL | 2025-07-14 | 14:47 | 2025-07-14 | 14:24 | ROLAND | 2 | 13.3 | 13.3 | / | / |  |  |  | 1.769 | 5.05 | 202 |  | Kinga |  | 3 | 2 | 83024
2 | 15.0 | 168 | P/FINAL | 2025-07-14 | 14:47 | 2025-07-14 | 14:24 | ROLAND | 2 | 12.4 | 12.4 | / | / |  |  |  |  | 1.56 | 186 |  | Kinga |  | 3 | 2 | 83024
2 | 15.0 | 169 | P/FINAL | 2025-07-14 | 14:48 | 2025-07-14 | 14:24 | ROLAND | 2 | 14.4 | 14.4 | / | / |  |  |  |  | 5.09 | 189 |  | Kinga |  | 3 | 2 | 83024
3 | 35.0 | 170 | P/FINAL | 2025-07-14 | 14:51 | 2025-07-14 | 14:24 | ROLAND | 2 | 33.5 | 33.1 | / | / |  |  |  | 1.780 | 1.50 | 197 |  | Kinga |  | 3 | 3 | 83024
3 | 35.0 | 171 | P/FINAL | 2025-07-14 | 14:52 | 2025-07-14 | 14:24 | ROLAND | 2 | 34.0 | 34.0 | / | / |  |  |  |  | 2.73 | 187 |  | Kinga |  | 3 | 3 | 83024
3 | 35.0 | 172 | P/FINAL | 2025-07-14 | 14:58 | 2025-07-14 | 14:24 | ROLAND | 2 | 32.7 | 32.8 | / | / |  |  |  |  | 0.96 | 204 |  | Kinga |  | 3 | 3 | 83024
4 | 60.0 | 173 | P/FINAL | 2025-07-14 | 15:04 | 2025-07-14 | 14:24 | ROLAND | 2 | 56.3 | 56.4 | / | / |  |  |  | 1.782 | 0.60 | 209 |  | Kinga |  | 3 | 4 | 83024
4 | 60.0 | 174 | P/FINAL | 2025-07-14 | 15:08 | 2025-07-14 | 14:24 | ROLAND | 2 | 58.8 | 58.9 | / | / |  |  |  |  | 2.11 | 204 |  | Kinga |  | 3 | 4 | 83024
4 | 60.0 | 175 | P/FINAL | 2025-07-14 | 15:09 | 2025-07-14 | 14:24 | ROLAND | 2 | 61.0 | 61.0 | / | / |  |  |  |  | 1.56 | 210 |  | Kinga |  | 3 | 4 | 83024
2 | 15.0 | 176 | P/FINAL | 2025-07-14 | 15:50 | 2025-07-14 | 15:29 | ROLAND | 2 | 14.9 | 14.9 | / | / |  |  |  |  | 3.95 | 211 |  | Kinga |  | 3 | 2 | 83024
2 | 15.0 | 177 | P/FINAL | 2025-07-14 | 15:51 | 2025-07-14 | 15:29 | ROLAND | 2 | 14.6 | 14.5 | / | / |  |  |  | 1.786 | 0.88 | 206 |  | Kinga |  | 3 | 2 | 83024
3 | 35.0 | 178 | P/FINAL | 2025-07-14 | 15:55 | 2025-07-14 | 15:29 | ROLAND | 2 | 31.5 | 31.7 | / | / |  |  |  |  | 0.68 | 206 |  | Kinga |  | 3 | 3 | 83024
3 | 35.0 | 179 | P/FINAL | 2025-07-14 | 15:56 | 2025-07-14 | 15:29 | ROLAND | 2 | 33.2 | 33.3 | / | / |  |  |  |  | 0.92 | 192 |  | Kinga |  | 3 | 3 | 83024
3 | 35.0 | 180 | P/FINAL | 2025-07-14 | 15:57 | 2025-07-14 | 15:29 | ROLAND | 2 | 32.2 | 32.2 | / | / |  |  |  | 1.779 | 3.50 | 198 |  | Kinga |  | 3 | 3 | 83024
4 | 60.0 | 181 | P/FINAL | 2025-07-14 | 16:05 | 2025-07-14 | 15:29 | ROLAND | 2 | 55.7 | 56.3 | / | / |  |  |  |  | 3.50 | 211 |  | Kinga |  | 3 | 4 | 83024
4 | 60.0 | 182 | P/FINAL | 2025-07-14 | 16:06 | 2025-07-14 | 15:29 | ROLAND | 2 | 59.2 | 58.8 | / | / |  |  |  |  | 3.84 | 209 |  | Kinga |  | 3 | 4 | 83024
4 | 60.0 | 183 | P/FINAL | 2025-07-14 | 16:07 | 2025-07-14 | 15:29 | ROLAND | 2 | 57.1 | 57.1 | / | / |  |  |  | 1.781 | 1.81 | 199 |  | Kinga |  | 3 | 4 | 83024
2 | 15.0 | 184 | P/FINAL | 2025-07-15 | 08:24 | 2025-07-15 | 07:45 | ELENA | 2 | 13.3 | 13.4 | / | / |  |  |  |  | 0.93 | 198 |  | Kinga |  | 3 | 2 | 83024
2 | 15.0 | 185 | P/FINAL | 2025-07-15 | 08:29 | 2025-07-15 | 07:45 | ELENA | 2 | 12.8 | 12.8 | / | / |  |  |  |  | 2.70 | 201 |  | Kinga |  | 3 | 2 | 83024
2 | 15.0 | 186 | P/FINAL | 2025-07-15 | 08:30 | 2025-07-15 | 07:45 | ELENA | 2 | 13.3 | 13.4 | / | / |  |  |  | 1.736 | 0.94 | 197 |  | Kinga |  | 3 | 2 | 83024
3 | 35.0 | 187 | P/FINAL | 2025-07-15 | 08:41 | 2025-07-15 | 07:45 | ELENA | 2 | 33.3 | 33.4 | / | / |  |  |  |  | 1.39 | 212 |  | Kinga |  | 3 | 3 | 83024
3 | 35.0 | 188 | P/FINAL | 2025-07-15 | 08:42 | 2025-07-15 | 07:45 | ELENA | 2 | 35.4 | 35.7 | / | / |  |  |  |  | 1.12 | 209 |  | Kinga |  | 3 | 3 | 83024
3 | 35.0 | 189 | P/FINAL | 2025-07-15 | 08:43 | 2025-07-15 | 07:45 | ELENA | 2 | 38.2 | 38.4 | / | / |  |  |  | 1.749 | 0.74 | 214 |  | Kinga |  | 3 | 3 | 83024
4 | 60.0 | 190 | P/FINAL | 2025-07-15 | 08:55 | 2025-07-15 | 07:45 | ELENA | 2 | 56.2 | 56.5 | / | / |  |  |  |  | 0.86 | 212 |  | Kinga |  | 3 | 4 | 83024
4 | 60.0 | 191 | P/FINAL | 2025-07-15 | 08:56 | 2025-07-15 | 07:45 | ELENA | 2 | 55.8 | 55.9 | / | / |  |  |  |  | 1.34 | 199 |  | Kinga |  | 3 | 4 | 83024
4 | 60.0 | 192 | P/FINAL | 2025-07-15 | 08:57 | 2025-07-15 | 07:45 | ELENA | 2 | 58.5 | 61.9 | / | / |  |  |  | 1.751 | 0.95 | 206 |  | Kinga |  | 3 | 4 | 83024
2 | 15.0 | 193 | P/FINAL | 2025-07-15 | 10:19 | 2025-07-15 | 09:35 | ELENA | 2 | 14.1 | 14.2 | / | / |  |  |  |  | 1.79 | 207 |  | Kinga |  | 3 | 2 | 83024
2 | 15.0 | 194 | P/FINAL | 2025-07-15 | 10:20 | 2025-07-15 | 09:35 | ELENA | 2 | 14.2 | 14.4 | / | / |  |  |  |  | 1.79 | 205 |  | Kinga |  | 3 | 2 | 83024
2 | 15.0 | 195 | P/FINAL | 2025-07-15 | 10:21 | 2025-07-15 | 09:35 | ELENA | 2 | 13.0 | 13.0 | / | / |  |  |  | 1.735 | 1.33 | 209 |  | Kinga |  | 3 | 2 | 83024
3 | 35.0 | 196 | P/FINAL | 2025-07-15 | 10:27 | 2025-07-15 | 09:35 | ELENA | 2 | 36.5 | 36.3 | / | / |  |  |  |  | 1.76 | 195 |  | Kinga |  | 3 | 3 | 83024
3 | 35.0 | 197 | P/FINAL | 2025-07-15 | 10:28 | 2025-07-15 | 09:35 | ELENA | 2 | 33.7 | 34.1 | / | / |  |  |  |  | 4.65 | 204 |  | Kinga |  | 3 | 3 | 83024
3 | 35.0 | 198 | P/FINAL | 2025-07-15 | 10:28 | 2025-07-15 | 09:35 | ELENA | 2 | 36.1 | 36.3 | / | / |  |  |  | 1.748 | 1.30 | 209 |  | Kinga |  | 3 | 3 | 83024
4 | 60.0 | 199 | P/FINAL | 2025-07-15 | 10:33 | 2025-07-15 | 09:35 | ELENA | 2 | 59.2 | 59.4 | / | / |  |  |  |  | 7.14 | 206 |  | Kinga |  | 3 | 4 | 83024
4 | 60.0 | 200 | P/FINAL | 2025-07-15 | 10:34 | 2025-07-15 | 09:35 | ELENA | 2 | 56.9 | 57.0 | / | / |  |  |  |  | 0.78 | 209 |  | Kinga |  | 3 | 4 | 83024
4 | 60.0 | 201 | P/FINAL | 2025-07-15 | 10:34 | 2025-07-15 | 09:35 | ELENA | 2 | 59.8 | 59.9 | / | / |  |  |  | 1.751 | 1.14 | 211 |  | Kinga |  | 3 | 4 | 83024
2 | 15.0 | 202 | P/FINAL | 2025-07-15 | 11:45 | 2025-07-15 | 11:13 | ELENA | 2 | 13.5 | 13.2 | / | / |  |  |  |  | 0.96 | 195 |  | Kinga |  | 3 | 2 | 83024
2 | 15.0 | 203 | P/FINAL | 2025-07-15 | 11:45 | 2025-07-15 | 11:13 | ELENA | 2 | 14.4 | 14.3 | / | / |  |  |  |  | 0.87 | 191 |  | Kinga |  | 3 | 2 | 83024
2 | 15.0 | 204 | P/FINAL | 2025-07-15 | 11:46 | 2025-07-15 | 11:13 | ELENA | 2 | 14.0 | 14.0 | / | / |  |  |  | 1.754 | 0.57 | 198 |  | Kinga |  | 3 | 2 | 83024
3 | 35.0 | 205 | P/FINAL | 2025-07-15 | 11:49 | 2025-07-15 | 11:13 | ELENA | 2 | 35.6 | 35.7 | / | / |  |  |  |  | 0.87 | 213 |  | Kinga |  | 3 | 3 | 83024
3 | 35.0 | 206 | P/FINAL | 2025-07-15 | 11:49 | 2025-07-15 | 11:13 | ELENA | 2 | 36.6 | 36.2 | / | / |  |  |  |  | 0.93 | 203 |  | Kinga |  | 3 | 3 | 83024
3 | 35.0 | 207 | P/FINAL | 2025-07-15 | 11:50 | 2025-07-15 | 11:13 | ELENA | 2 | 37.5 | 36.6 | / | / |  |  |  | 1.754 | 0.87 | 203 |  | Kinga |  | 3 | 3 | 83024
4 | 60.0 | 208 | P/FINAL | 2025-07-15 | 11:56 | 2025-07-15 | 11:13 | ELENA | 2 | 59.6 | 59.9 | / | / |  |  |  |  | 0.99 | 198 |  | Kinga |  | 3 | 4 | 83024
4 | 60.0 | 209 | P/FINAL | 2025-07-15 | 11:56 | 2025-07-15 | 11:13 | ELENA | 2 | 55.6 | 55.1 | / | / |  |  |  |  | 0.79 | 208 |  | Kinga |  | 3 | 4 | 83024
4 | 60.0 | 210 | P/FINAL | 2025-07-15 | 11:57 | 2025-07-15 | 11:13 | ELENA | 2 | 55.3 | 55.4 | / | / |  |  |  | 1.765 | 1.27 | 206 |  | Kinga |  | 3 | 4 | 83024
2 | 15.0 | 211 | P/FINAL | 2025-07-15 | 13:40 | 2025-07-15 | 12:37 | ELENA | 2 | 15.0 | 14.7 | / | / |  |  |  |  | 1.18 | 194 |  | Kinga |  | 3 | 2 | 83024
2 | 15.0 | 212 | P/FINAL | 2025-07-15 | 13:41 | 2025-07-15 | 12:37 | ELENA | 2 | 15.2 | 15.1 | / | / |  |  |  |  | 7.06 | 209 |  | Kinga |  | 3 | 2 | 83024
2 | 15.0 | 213 | P/FINAL | 2025-07-15 | 13:41 | 2025-07-15 | 12:37 | ELENA | 2 | 14.2 | 14.3 | / | / |  |  |  | 1.776 | 1.18 | 206 |  | Kinga |  | 3 | 2 | 83024
3 | 35.0 | 214 | P/FINAL | 2025-07-15 | 13:42 | 2025-07-15 | 12:37 | ELENA | 2 | 35.4 | 35.4 | / | / |  |  |  |  | 3.00 | 196 |  | Kinga |  | 3 | 3 | 83024
3 | 35.0 | 215 | P/FINAL | 2025-07-15 | 13:43 | 2025-07-15 | 12:37 | ELENA | 2 | 33.5 | 33.5 | / | / |  |  |  |  | 1.17 | 196 |  | Kinga |  | 3 | 3 | 83024
3 | 35.0 | 216 | P/FINAL | 2025-07-15 | 13:44 | 2025-07-15 | 12:37 | ELENA | 2 | 34.8 | 34.9 | / | / |  |  |  | 1.769 | 2.30 | 208 |  | Kinga |  | 3 | 3 | 83024
4 | 60.0 | 217 | P/FINAL | 2025-07-15 | 13:45 | 2025-07-15 | 12:37 | ELENA | 2 | 56.6 | 56.6 | / | / |  |  |  |  | 1.41 | 190 |  | Kinga |  | 3 | 4 | 83024
4 | 60.0 | 218 | P/FINAL | 2025-07-15 | 13:46 | 2025-07-15 | 12:37 | ELENA | 2 | 56.2 | 56.4 | / | / |  |  |  |  | 1.70 | 193 |  | Kinga |  | 3 | 4 | 83024
4 | 60.0 | 219 | P/FINAL | 2025-07-15 | 13:46 | 2025-07-15 | 12:37 | ELENA | 2 | 57.3 | 57.4 | / | / |  |  |  | 1.759 | 1.63 | 198 |  | Kinga |  | 3 | 4 | 83024
2 | 15.0 | 220 | P/FINAL | 2025-07-15 | 14:38 | 2025-07-15 | 14:17 | ROLAND | 2 | 15.1 | 15.1 | / | / |  |  |  |  | 1.31 | 190 |  | Kinga |  | 3 | 2 | 83024
2 | 15.0 | 221 | P/FINAL | 2025-07-15 | 14:39 | 2025-07-15 | 14:17 | ROLAND | 2 | 16.4 | 16.2 | / | / |  |  |  |  | 4.97 | 202 |  | Kinga |  | 3 | 2 | 83024
2 | 15.0 | 222 | P/FINAL | 2025-07-15 | 14:40 | 2025-07-15 | 14:17 | ROLAND | 2 | 15.5 | 15.6 | / | / |  |  |  | 1.781 | 4.97 | 208 |  | Kinga |  | 3 | 2 | 83024
3 | 35.0 | 223 | P/FINAL | 2025-07-15 | 14:42 | 2025-07-15 | 14:17 | ROLAND | 2 | 34.2 | 34.5 | / | / |  |  |  |  | 5.68 | 189 |  | Kinga |  | 3 | 3 | 83024
3 | 35.0 | 224 | P/FINAL | 2025-07-15 | 14:43 | 2025-07-15 | 14:17 | ROLAND | 2 | 37.2 | 36.7 | / | / |  |  |  |  | 4.82 | 199 |  | Kinga |  | 3 | 3 | 83024
3 | 35.0 | 225 | P/FINAL | 2025-07-15 | 14:44 | 2025-07-15 | 14:17 | ROLAND | 2 | 35.3 | 35.3 | / | / |  |  |  |  | 2.95 | 195 |  | Kinga |  | 3 | 3 | 83024
4 | 60.0 | 226 | P/FINAL | 2025-07-15 | 14:45 | 2025-07-15 | 14:17 | ROLAND | 2 | 60.0 | 59.8 | / | / |  |  |  | 1.754 | 1.91 | 211 |  | Kinga |  | 3 | 4 | 83024
4 | 60.0 | 227 | P/FINAL | 2025-07-15 | 14:45 | 2025-07-15 | 14:17 | ROLAND | 2 | 60.0 | 59.8 | / | / |  |  |  |  | 1.91 | 203 |  | Kinga |  | 3 | 4 | 83024
4 | 60.0 | 228 | P/FINAL | 2025-07-15 | 14:46 | 2025-07-15 | 14:17 | ROLAND | 2 | 59.0 | 59.0 | / | / |  |  |  |  | 2.68 | 198 |  | Kinga |  | 3 | 4 | 83024
4 | 60.0 | 229 | P/FINAL | 2025-07-15 | 14:46 | 2025-07-15 | 14:17 | ROLAND | 2 | 59.1 | 59.0 | / | / |  |  |  | 1.775 | 2.36 | 211 |  | Kinga |  | 3 | 4 | 83024
2 | 15.0 | 230 | P/FINAL | 2025-07-15 | 15:41 | 2025-07-15 | 15:19 | ROLAND | 2 | 14.0 | 13.8 | / | / |  |  |  |  | 5.36 | 197 |  | Kinga |  | 3 | 2
2 | 15.0 | 231 | P/FINAL | 2025-07-15 | 15:42 | 2025-07-15 | 15:19 | ROLAND | 2 | 15.5 | 15.7 | / | / |  |  |  |  | 1.19 | 199 |  | Kinga |  | 3 | 2
2 | 15.0 | 232 | P/FINAL | 2025-07-15 | 15:42 | 2025-07-15 | 15:19 | ROLAND | 2 | 14.3 | 14.4 | / | / |  |  |  |  | 5.75 | 204 |  | Kinga |  | 3 | 2
3 | 35.0 | 233 | P/FINAL | 2025-07-15 | 15:45 | 2025-07-15 | 15:19 | ROLAND | 2 | 34.4 | 34.5 | / | / |  |  |  | 1.818 | 1.62 | 209 |  | Kinga |  | 3 | 3
3 | 35.0 | 234 | P/FINAL | 2025-07-15 | 15:45 | 2025-07-15 | 15:19 | ROLAND | 2 | 33.2 | 33.3 | / | / |  |  |  |  | 4.73 | 202 |  | Kinga |  | 3 | 3
3 | 35.0 | 235 | P/FINAL | 2025-07-15 | 15:46 | 2025-07-15 | 15:19 | ROLAND | 2 | 32.3 | 32.3 | / | / |  |  |  |  | 5.73 | 186 |  | Kinga |  | 3 | 3
4 | 60.0 | 236 | P/FINAL | 2025-07-15 | 15:47 | 2025-07-15 | 15:19 | ROLAND | 2 | 58.7 | 58.2 | / | / |  |  |  |  | 2.00 | 189 |  | Kinga |  | 3 | 4
4 | 60.0 | 237 | P/FINAL | 2025-07-15 | 15:47 | 2025-07-15 | 15:19 | ROLAND | 2 | 58.7 | 58.2 | / | / |  |  |  | 1.770 | 2.00 | 211 |  | Kinga |  | 3 | 4
4 | 60.0 | 238 | P/FINAL | 2025-07-15 | 15:49 | 2025-07-15 | 15:19 | ROLAND | 2 | 55.9 | 56.1 | / | / |  |  |  |  | 2.13 | 195 |  | Kinga |  | 3 | 4
4 | 60.0 | 239 | P/FINAL | 2025-07-15 | 15:50 | 2025-07-15 | 15:19 | ROLAND | 2 | 58.8 | 58.9 | / | / |  |  |  |  | 3.45 | 198 |  | Kinga |  | 3 | 4
4 | 60.0 | 240 | P/FINAL | 2025-07-15 | 15:50 | 2025-07-15 | 15:19 | ROLAND | 2 | 60.1 | 60.2 | / | / |  |  |  | 1.732 | 3.45 | 197 |  | Kinga |  | 3 | 4
2 | 15.0 | 241 | P/FINAL | 2025-07-16 | 07:14 | 2025-07-16 | 06:43 | ELENA | 2 | 13.2 | 13.2 | / | / |  |  |  |  | 2.24 | 189 |  | Kinga |  | 3 | 2 | 83024
2 | 15.0 | 242 | P/FINAL | 2025-07-16 | 07:15 | 2025-07-16 | 06:43 | ELENA | 2 | 14.8 | 14.8 | / | / |  |  |  |  | 6.74 | 199 |  | Kinga |  | 3 | 2 | 83024
2 | 15.0 | 243 | P/FINAL | 2025-07-16 | 07:15 | 2025-07-16 | 06:43 | ELENA | 2 | 13.6 | 13.6 | / | / |  |  |  | 1.848 | 0.97 | 196 |  | Kinga |  | 3 | 2 | 83024
3 | 35.0 | 244 | P/FINAL | 2025-07-16 | 07:17 | 2025-07-16 | 06:43 | ELENA | 2 | 33.8 | 33.9 | / | / |  |  |  |  | 1.04 | 200 |  | Kinga |  | 3 | 3 | 83024
3 | 35.0 | 245 | P/FINAL | 2025-07-16 | 07:17 | 2025-07-16 | 06:43 | ELENA | 2 | 34.9 | 34.7 | / | / |  |  |  |  | 1.13 | 205 |  | Kinga |  | 3 | 3 | 83024
3 | 35.0 | 246 | P/FINAL | 2025-07-16 | 07:19 | 2025-07-16 | 06:43 | ELENA | 2 | 33.0 | 33.2 | / | / |  |  |  | 1.815 | 0.97 | 192 |  | Kinga |  | 3 | 3 | 83024
4 | 60.0 | 247 | P/FINAL | 2025-07-16 | 07:29 | 2025-07-16 | 06:43 | ELENA | 2 | 58.4 | 58.4 | / | / |  |  |  |  | 1.25 | 196 |  | Kinga |  | 3 | 4 | 83024
4 | 60.0 | 248 | P/FINAL | 2025-07-16 | 07:29 | 2025-07-16 | 06:43 | ELENA | 2 | 57.8 | 58.5 | / | / |  |  |  |  | 1.72 | 199 |  | Kinga |  | 3 | 4 | 83024
4 | 60.0 | 249 | P/FINAL | 2025-07-16 | 07:30 | 2025-07-16 | 06:43 | ELENA | 2 | 59.0 | 59.1 | / | / |  |  |  | 1.823 | 5.65 | 197 |  | Kinga |  | 3 | 4 | 83024
2 | 15.0 | 250 | P/FINAL | 2025-07-16 | 09:04 | 2025-07-16 | 08:26 | ELENA | 2 | 14.2 | 14.2 | / | / |  |  |  |  | 1.16 | 186 |  | Kinga |  | 3 | 2 | 23843
2 | 15.0 | 251 | P/FINAL | 2025-07-16 | 09:04 | 2025-07-16 | 08:26 | ELENA | 2 | 13.5 | 13.6 | / | / |  |  |  |  | 1.48 | 193 |  | Kinga |  | 3 | 2 | 23843
2 | 15.0 | 252 | P/FINAL | 2025-07-16 | 09:05 | 2025-07-16 | 08:26 | ELENA | 2 | 14.9 | 14.9 | / | / |  |  |  | 1.848 | 14.8 | 197 |  | Kinga |  | 3 | 2 | 23843
3 | 35.0 | 253 | P/FINAL | 2025-07-16 | 09:07 | 2025-07-16 | 08:26 | ELENA | 2 | 33.2 | 32.7 | / | / |  |  |  |  | 1.37 | 205 |  | Kinga |  | 3 | 3 | 23843
3 | 35.0 | 254 | P/FINAL | 2025-07-16 | 09:07 | 2025-07-16 | 08:26 | ELENA | 2 | 34.9 | 35.0 | / | / |  |  |  |  | 1.68 | 192 |  | Kinga |  | 3 | 3 | 23843
3 | 35.0 | 255 | P/FINAL | 2025-07-16 | 09:08 | 2025-07-16 | 08:26 | ELENA | 2 | 34.9 | 35.0 | / | / |  |  |  | 1.815 | 1.93 | 213 |  | Kinga |  | 3 | 3 | 23843
4 | 60.0 | 256 | P/FINAL | 2025-07-16 | 09:11 | 2025-07-16 | 08:26 | ELENA | 2 | 55.2 | 55.2 | / | / |  |  |  |  | 1.15 | 206 |  | Kinga |  | 3 | 4 | 23843
4 | 60.0 | 257 | P/FINAL | 2025-07-16 | 09:12 | 2025-07-16 | 08:26 | ELENA | 2 | 59.8 | 60.0 | / | / |  |  |  |  | 2.48 | 197 |  | Kinga |  | 3 | 4 | 23843
4 | 60.0 | 258 | P/FINAL | 2025-07-16 | 09:12 | 2025-07-16 | 08:26 | ELENA | 2 | 60.0 | 60.3 | / | / |  |  |  | 1.830 | 2.84 | 208 |  | Kinga |  | 3 | 4 | 23843
2 | 15.0 | 259 | P/FINAL | 2025-07-16 | 10:04 | 2025-07-16 | 09:36 | ELENA | 2 | 13.2 | 13.3 | / | / |  |  |  |  | 5.25 | 207 |  | Kinga |  | 3 | 2 | 23843
2 | 15.0 | 260 | P/FINAL | 2025-07-16 | 10:04 | 2025-07-16 | 09:36 | ELENA | 2 | 13.8 | 13.8 | / | / |  |  |  |  | 6.06 | 195 |  | Kinga |  | 3 | 2 | 23843
2 | 15.0 | 261 | P/FINAL | 2025-07-16 | 10:05 | 2025-07-16 | 09:36 | ELENA | 2 | 13.1 | 13.0 | / | / |  |  |  | 1.831 | 0.92 | 200 |  | Kinga |  | 3 | 2 | 23843
4 | 60.0 | 262 | P/FINAL | 2025-07-16 | 10:46 | 2025-07-16 | 09:36 | ELENA | 2 | 56.0 | 56.2 | / | / |  |  |  |  | 1.22 | 198 |  | Kinga |  | 3 | 4 | 2384309:36
4 | 60.0 | 263 | P/FINAL | 2025-07-16 | 10:46 | 2025-07-16 | 09:36 | ELENA | 2 | 58.0 | 58.4 | / | / |  |  |  |  | 5.31 | 196 |  | Kinga |  | 3 | 4 | 2384309:36
4 | 60.0 | 264 | P/FINAL | 2025-07-16 | 10:47 | 2025-07-16 | 09:36 | ELENA | 2 | 57.1 | 57.3 | / | / |  |  |  | 1.844 | 1.71 | 192 |  | Kinga |  | 3 | 4 | 2384309:36
3 | 35.0 | 265 | P/FINAL | 32025/07/16 | 10:53 | 2025-07-16 | 09:36 | ELENA | 2 | 35.2 | 34.8 | / | / |  |  |  |  | 3.30 | 198 |  | Kinga |  | 3 | 3 | 2384309:36
3 | 35.0 | 266 | P/FINAL | 32025/07/16 | 10:54 | 2025-07-16 | 09:36 | ELENA | 2 | 35.5 | 34.9 | / | / |  |  |  |  | 3.30 | 202 |  | Kinga |  | 3 | 3 | 2384309:36
3 | 35.0 | 267 | P/FINAL | 32025/07/16 | 10:54 | 2025-07-16 | 09:36 | ELENA | 2 | 34.7 | 34.4 | / | / |  |  |  | 1.812 | 2.45 | 207 |  | Kinga |  | 3 | 3 | 2384309:36
2 | 15.0 | 268 | P/FINAL | 2025-07-16 | 12:02 | 2025-07-16 | 11:45 | ELENA | 2 | 15.3 | 15.3 | / | / |  |  |  |  | 8.64 | 202 |  | Kinga |  | 3 | 2 | 23843
2 | 15.0 | 269 | P/FINAL | 2025-07-16 | 12:03 | 2025-07-16 | 11:45 | ELENA | 2 | 15.1 | 15.2 | / | / |  |  |  |  | 9.38 | 195 |  | Kinga |  | 3 | 2 | 23843
2 | 15.0 | 270 | P/FINAL | 2025-07-16 | 12:04 | 2025-07-16 | 11:45 | ELENA | 2 | 12.9 | 13.0 | / | / |  |  |  | 1.826 | 2.62 | 200 |  | Kinga |  | 3 | 2 | 23843
3 | 35.0 | 271 | P/FINAL | 2025-07-16 | 12:07 | 2025-07-16 | 11:45 | ELENA | 2 | 35.4 | 35.5 | / | / |  |  |  |  | 1.92 | 202 |  | Kinga |  | 3 | 3 | 23843
3 | 35.0 | 272 | P/FINAL | 2025-07-16 | 12:08 | 2025-07-16 | 11:45 | ELENA | 2 | 34.7 | 34.7 | / | / |  |  |  |  | 2.52 | 209 |  | Kinga |  | 3 | 3 | 23843
3 | 35.0 | 273 | P/FINAL | 2025-07-16 | 12:09 | 2025-07-16 | 11:45 | ELENA | 2 | 35.7 | 35.8 | / | / |  |  |  |  | 2.14 | 198 |  | Kinga |  | 3 | 3 | 23843
3 | 35.0 | 274 | P/FINAL | 2025-07-16 | 12:09 | 2025-07-16 | 11:45 | ELENA | 2 | 35.7 | 35.8 | / | / |  |  |  | 1.863 | 2.14 | 200 |  | Kinga |  | 3 | 3 | 23843
4 | 60.0 | 275 | P/FINAL | 2025-07-16 | 12:14 | 2025-07-16 | 11:45 | ELENA | 2 | 55.8 | 55.9 | / | / |  |  |  |  | 1.93 | 209 |  | Kinga |  | 3 | 4 | 23843
4 | 60.0 | 276 | P/FINAL | 2025-07-16 | 12:15 | 2025-07-16 | 11:45 | ELENA | 2 | 55.1 | 55.1 | / | / |  |  |  |  | 1.43 | 207 |  | Kinga |  | 3 | 4 | 23843
4 | 60.0 | 277 | P/FINAL | 2025-07-16 | 12:15 | 2025-07-16 | 11:45 | ELENA | 2 | 55.3 | 55.8 | / | / |  |  |  | 1.826 | 2.12 | 202 |  | Kinga |  | 3 | 4 | 23843
3 | 35.0 | 278 | P/FINAL | 2025-07-16 | 13:40 | 2025-07-16 | 13:00 | ELENA | 2 | 34.1 | 34.2 | / | / |  |  |  |  | 2.12 | 204 |  | Kinga |  | 3 | 3
3 | 35.0 | 279 | P/FINAL | 2025-07-16 | 13:40 | 2025-07-16 | 13:00 | ELENA | 2 | 35.8 | 35.5 | / | / |  |  |  |  | 2.43 | 198 |  | Kinga |  | 3 | 3
3 | 35.0 | 280 | P/FINAL | 2025-07-16 | 13:40 | 2025-07-16 | 13:00 | ELENA | 2 | 35.8 | 35.5 | / | / |  |  |  | 1.814 | 2.43 | 202 |  | Kinga |  | 3 | 3
3 | 35.0 | 281 | P/FINAL | 2025-07-16 | 13:40 | 2025-07-16 | 13:00 | ELENA | 2 | 36.3 | 36.4 | / | / |  |  |  |  | 2.51 | 197 |  | Kinga |  | 3 | 3
2 | 15.0 | 282 | P/FINAL | 2025-07-16 | 13:42 | 2025-07-16 | 13:00 | ELENA | 2 | 13.6 | 13.6 | / | / |  |  |  |  | 4.05 | 102 |  | Kinga |  | 3 | 2
2 | 15.0 | 283 | P/FINAL | 2025-07-16 | 13:43 | 2025-07-16 | 13:00 | ELENA | 2 | 15.2 | 15.2 | / | / |  |  |  |  | 1.44 | 211 |  | Kinga |  | 3 | 2
2 | 15.0 | 284 | P/FINAL | 2025-07-16 | 13:43 | 2025-07-16 | 13:00 | ELENA | 2 | 14.6 | 14.5 | / | / |  |  |  | 1.853 | 2.55 | 194 |  | Kinga |  | 3 | 2
4 | 60.0 | 285 | P/FINAL | 2025-07-16 | 13:45 | 2025-07-16 | 13:00 | ELENA | 2 | 57.1 | 56.6 | / | / |  |  |  |  | 4.95 | 207 |  | Kinga |  | 3 | 4
4 | 60.0 | 286 | P/FINAL | 2025-07-16 | 13:45 | 2025-07-16 | 13:00 | ELENA | 2 | 55.5 | 55.4 | / | / |  |  |  |  | 1.21 | 201 |  | Kinga |  | 3 | 4
4 | 60.0 | 287 | P/FINAL | 2025-07-16 | 13:45 | 2025-07-16 | 13:00 | ELENA | 2 | 55.5 | 55.4 | / | / |  |  |  | 1.827 | 1.21 | 198 |  | Kinga |  | 3 | 4
Standard # 3
Total Tests | Total Readings | Total Average | Selected Tests | Selected Readings | Selected Average
86 | 172 | 34.9 mg/L | 86 | 172 | 34.9 mg/L
Standard # 4
Total Tests | Total Readings | Total Average | Selected Tests | Selected Readings | Selected Average
86 | 172 | 57.8 mg/L | 86 | 172 | 57.8 mg/L
Standard # 2
Total Tests | Total Readings | Total Average | Selected Tests | Selected Readings | Selected Average
83 | 166 | 14.3 mg/L | 83 | 166 | 14.3 mg/L
Standard # 1
Total Tests | Total Readings | Total Average | Selected Tests | Selected Readings | Selected Average
5 | 10 | 0 mg/L | 5 | 10 | 0 mg/L
Specifications
Standard Value [ppm] | Target Value ÷ U [ppm] | Mean Value [ppm] | Passed
0.0 | 0.0 ÷ 2.0 | 0.0 | YES
15.0 | 12.3 ÷ 17.7 | 14.3 | YES
35.0 | 31.3 ÷ 38.7 | 34.9 | YES
60.0 | 55.1 ÷ 64.9 | 57.8 | YES
Validation Date | by