## Hanna Code
Lot | Hanna SFG Code | Recipe | Preparation Week |  | Hanna SFG Code | HI782-0
1159 | HI782-0 | CP-R80 | W 35/1 |  | Parameter Formula | NO3
Description | Exp | Line |  |  | QC Department | Registration Book | QC Type
Marine Nitrate HR Reagent | 2030-08-01 | L57 Powder |  |  | Laboratory | TABLE 3 | Production
Ref Weight. | Min (mg) | Max (mg)
206 | 196 | 216
REAGENT SET 1 | [5 empty] | REAGENT SET 2
Reagent A Lot | Reagent B Lot | Reagent C Lot | Reagent D Lot | Reagent E Lot |  | Reagent A Lot | Reagent B Lot | Reagent C Lot | Reagent D Lot | Reagent E Lot
Expiration A | Expiration B | Expiration C | Expiration D | Expiration E |  | Expiration A | Expiration B | Expiration C | Expiration D | Expiration E
Reagent Range Min | Reagent Range Max
0.0 mg/L | 75.0 mg/L
Prep. Operator | First day Prod. | Last day Prod. | Machine |  |  | Old Lot A | Old Lot B | Old Lot D
ERIKA | 2025-09-01 | 2025-09-08 | 2 HEAD |  |  | 1141
[6 empty] | Expiration A | Expiration B | Expiration D
2030-08-01
Meter
Meter 1 Family | Meter 1 Code | Meter 2 Family | Meter 2 Code
HI97115 | CTO-740 | HI97115 | CTO-730
ph Meter | Description | Turbid. Meter | Description | Spectr. Meter | Description
|  | HI98713 | CTO-810
Standard | mg/L | [5 empty] | Tolerance |  |  |  |  | pH
STD Number | STD Value | STD Min | STD Max |  | STD MR |  | Fixed | AndOr | Percentage | QCRestriction |  | pH Value | pH Min | pH Max
1 | 0.0 | 0.0 | 2.0 |  | MR101 |  | 2.0 | & | 4.9% | 100%
2 | 15.0 | 12.3 | 17.7
3 | 35.0 | 31.3 | 38.7 | [8 empty] | 1.7 | 1.5 | 1.9
4 | 60.0 | 55.1 | 64.9
Readings Table
Standard | STD Value | # | TEST | QC DATE | QC TIME | PROD. DATE | PROD. TIME | PROD. OPERATOR | HEAD | METER 1 | METER 2 | METER 3 | METER 4 | SPECTR. [ABS] | pH 1 | pH 2 | pH 3 | TURB. | WEIGHT [mg] | REAGENT SET | QC OPERATOR | CORRECTION | phNumber | STD | STD_ID | NOTE
1 | 0.0 | 1 | Old A | 2025-09-01 | 09:15 | 2025-09-01 | 08:38 | MARK |  | 0.0 | 0.0 | / | / |  |  |  |  | 1.63 |  |  | Kinga |  | 3 | 1
1 | 0.0 | 2 | Old A | 2025-09-01 | 09:16 | 2025-09-01 | 08:38 | MARK |  | 0.0 | 0.0 | / | / |  |  |  |  | 2.89 |  |  | Kinga |  | 3 | 1
2 | 15.0 | 3 | Old A | 2025-09-01 | 09:18 | 2025-09-01 | 08:38 | MARK |  | 12.9 | 13.8 | / | / |  |  |  |  | 1.01 |  |  | Kinga |  | 3 | 2 | 124346
2 | 15.0 | 4 | Old A | 2025-09-01 | 09:19 | 2025-09-01 | 08:38 | MARK |  | 12.8 | 13.4 | / | / |  |  |  |  | 1.17 |  |  | Kinga |  | 3 | 2 | 124346
3 | 35.0 | 5 | Old A | 2025-09-01 | 09:21 | 2025-09-01 | 08:38 | MARK |  | 32.0 | 32.6 | / | / |  |  |  |  | 2.25 |  |  | Kinga |  | 3 | 3 | 124346
3 | 35.0 | 6 | Old A | 2025-09-01 | 09:21 | 2025-09-01 | 08:38 | MARK |  | 31.9 | 32.1 | / | / |  |  |  |  | 4.36 |  |  | Kinga |  | 3 | 3 | 124346
4 | 60.0 | 7 | Old A | 2025-09-01 | 09:28 | 2025-09-01 | 08:38 | MARK |  | 56.6 | 56.9 | / | / |  |  |  |  | 115 |  |  | Kinga |  | 3 | 4 | 124346
4 | 60.0 | 8 | Old A | 2025-09-01 | 09:28 | 2025-09-01 | 08:38 | MARK |  | 57.7 | 57.2 | / | / |  |  |  |  | 1.37 |  |  | Kinga |  | 3 | 4 | 124346
2 | 15.0 | 9 | Valid | 2025-09-01 | 09:33 | 2025-09-01 | 08:38 | MARK |  | 14.3 | 15.1 | / | / |  |  |  |  | 6.28 |  |  | Kinga |  | 3 | 2 | 124346
2 | 15.0 | 10 | Valid | 2025-09-01 | 09:33 | 2025-09-01 | 08:38 | MARK |  | 14.3 | 15.1 | / | / |  |  |  | 1.856 | 0.89 |  |  | Kinga |  | 3 | 2 | 124346
3 | 35.0 | 11 | Valid | 22025/09/01 | 09:35 | 2025-09-01 | 08:38 | MARK |  | 35.3 | 35.7 | / | / |  |  |  | 1.869 | 1.41 |  |  | Kinga |  | 3 | 3 | 124346
3 | 35.0 | 12 | Valid | 22025/09/01 | 09:36 | 2025-09-01 | 08:38 | MARK |  | 33.6 | 34.4 | / | / |  |  |  |  | 6.91 |  |  | Kinga |  | 3 | 3 | 124346
4 | 60.0 | 13 | Valid | 22025/09/01 | 09:39 | 2025-09-01 | 08:38 | MARK |  | 59.6 | 60.2 | / | / |  |  |  | 1.860 | 1.98 |  |  | Kinga |  | 3 | 4 | 124346
4 | 60.0 | 14 | Valid | 22025/09/01 | 09:40 | 2025-09-01 | 08:38 | MARK |  | 59.6 | 60.3 | / | / |  |  |  |  | 2.35 |  |  | Kinga |  | 3 | 4 | 124346
1 | 0.0 | 15 | Valid | 22025/09/01 | 09:45 | 2025-09-01 | 08:38 | MARK |  | 0.0 | 0.0 | / | / |  |  |  |  | 1.53 |  |  | Kinga |  | 3 | 1 | 124346
1 | 0.0 | 16 | Valid | 22025/09/01 | 09:46 | 2025-09-01 | 08:38 | MARK |  | 0.0 | 0.0 | / | / |  |  |  | 1.846 | 1.40 |  |  | Kinga |  | 3 | 1 | 124346
1 | 0.0 | 17 | Old A | 2025-09-03 | 15:40 | 2025-09-03 | 15:31 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 2.51 |  |  | Kinga |  | 3 | 1 | 90827
1 | 0.0 | 18 | Old A | 2025-09-03 | 15:40 | 2025-09-03 | 15:31 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 2.47 |  |  | Kinga |  | 3 | 1 | 90827
2 | 15.0 | 19 | Old A | 2025-09-03 | 15:41 | 2025-09-03 | 15:31 | ELENA | 2 | 15.6 | 15.5 | / | / |  |  |  |  | 1.34 |  |  | Kinga |  | 3 | 2 | 90827
2 | 15.0 | 20 | Old A | 2025-09-03 | 15:41 | 2025-09-03 | 15:31 | ELENA | 2 | 15.2 | 15.1 | / | / |  |  |  |  | 6.92 |  |  | Kinga |  | 3 | 2 | 90827
3 | 35.0 | 21 | Old A | 2025-09-03 | 15:41 | 2025-09-03 | 15:31 | ELENA | 2 | 34.3 | 33.6 | / | / |  |  |  |  | 3.94 |  |  | Kinga |  | 3 | 3 | 90827
3 | 35.0 | 22 | Old A | 2025-09-03 | 15:42 | 2025-09-03 | 15:31 | ELENA | 2 | 35.2 | 34.9 | / | / |  |  |  |  | 2.98 |  |  | Kinga |  | 3 | 3 | 90827
4 | 60.0 | 23 | Old A | 2025-09-03 | 15:42 | 2025-09-03 | 15:31 | ELENA | 2 | 58.4 | 59.1 | / | / |  |  |  |  | 3.82 |  |  | Kinga |  | 3 | 4 | 90827
4 | 60.0 | 24 | Old A | 2025-09-03 | 15:42 | 2025-09-03 | 15:31 | ELENA | 2 | 58.7 | 58.5 | / | / |  |  |  |  | 1.74 |  |  | Kinga |  | 3 | 4 | 90827
1 | 0.0 | 25 | P/FINAL | 2025-09-03 | 15:52 | 2025-09-03 | 15:31 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 4.13 | 211 |  | Kinga |  | 3 | 1 | 90827
1 | 0.0 | 26 | P/FINAL | 2025-09-03 | 15:53 | 2025-09-03 | 15:31 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  | 1.896 | 4.11 | 189 |  | Kinga |  | 3 | 1 | 90827
2 | 15.0 | 27 | P/FINAL | 2025-09-03 | 15:54 | 2025-09-03 | 15:31 | ELENA | 2 | 13.9 | 14.0 | / | / |  |  |  |  | 3.15 | 222 |  | Kinga |  | 3 | 2 | 90827
2 | 15.0 | 28 | P/FINAL | 2025-09-03 | 15:55 | 2025-09-03 | 15:31 | ELENA | 2 | 14.1 | 14.0 | / | / |  |  |  |  | 3.11 | 200 |  | Kinga |  | 3 | 2 | 90827
3 | 35.0 | 29 | P/FINAL | 2025-09-03 | 15:56 | 2025-09-03 | 15:31 | ELENA | 2 | 37.1 | 36.6 | / | / |  |  |  | 1.900 | 2.66 | 199 |  | Kinga |  | 3 | 3 | 90827
3 | 35.0 | 30 | P/FINAL | 2025-09-03 | 15:56 | 2025-09-03 | 15:31 | ELENA | 2 | 35.2 | 35.1 | / | / |  |  |  |  | 3.02 | 193 |  | Kinga |  | 3 | 3 | 90827
3 | 35.0 | 31 | P/FINAL | 2025-09-03 | 15:57 | 2025-09-03 | 15:31 | ELENA | 2 | 35.4 | 35.4 | / | / |  |  |  |  | 2.38 | 210 |  | Kinga |  | 3 | 3 | 90827
2 | 15.0 | 32 | P/FINAL | 2025-09-03 | 15:58 | 2025-09-03 | 15:31 | ELENA | 2 | 14.4 | 14.3 | / | / |  |  |  | 1.852 | 2.69 | 189 |  | Kinga |  | 3 | 2 | 90827
4 | 60.0 | 33 | P/FINAL | 2025-09-03 | 15:59 | 2025-09-03 | 15:31 | ELENA | 2 | 59.6 | 59.8 | / | / |  |  |  |  | 0.76 | 199 |  | Kinga |  | 3 | 4 | 90827
4 | 60.0 | 34 | P/FINAL | 2025-09-03 | 15:59 | 2025-09-03 | 15:31 | ELENA | 2 | 58.3 | 56.1 | / | / |  |  |  |  | 0.88 | 201 |  | Kinga |  | 3 | 4 | 90827
4 | 60.0 | 35 | P/FINAL | 2025-09-03 | 16:00 | 2025-09-03 | 15:31 | ELENA | 2 | 61.8 | 60.5 | / | / |  |  |  | 1.891 | 1.01 | 209 |  | Kinga |  | 3 | 4 | 90827
2 | 15.0 | 36 | Old A | 2025-09-04 | 08:12 | 2025-09-03 | 18:55 | ELENA | 2 | 15.5 | 15.7 | / | / |  |  |  |  | 4.47 | 200 |  | Kinga |  | 3 | 2 | 124346
2 | 15.0 | 37 | Old A | 2025-09-04 | 08:12 | 2025-09-03 | 18:55 | ELENA | 2 | 16.7 | 16.7 | / | / |  |  |  | 1.891 | 1.06 | 202 |  | Kinga |  | 3 | 2 | 124346
3 | 35.0 | 38 | Old A | 2025-09-04 | 08:16 | 2025-09-03 | 18:55 | ELENA | 2 | 38.1 | 38.2 | / | / |  |  |  |  | 2.86 | 209 |  | Kinga |  | 3 | 3 | 124346
3 | 35.0 | 39 | Old A | 2025-09-04 | 08:16 | 2025-09-03 | 18:55 | ELENA | 2 | 38.1 | 38.2 | / | / |  |  |  | 1.870 | 1.29 | 194 |  | Kinga |  | 3 | 3 | 124346
4 | 60.0 | 40 | Old A | 2025-09-04 | 08:19 | 2025-09-03 | 18:55 | ELENA | 2 | 60.5 | 60.6 | / | / |  |  |  |  | 2.77 | 190 |  | Kinga |  | 3 | 4 | 124346
4 | 60.0 | 41 | Old A | 2025-09-04 | 08:20 | 2025-09-03 | 18:55 | ELENA | 2 | 60.8 | 60.5 | / | / |  |  |  | 1.894 | 2.93 | 206 |  | Kinga |  | 3 | 4 | 124346
2 | 15.0 | 42 | P/FINAL | 2025-09-04 | 08:34 | 2025-09-04 | 08:10 | ROLAND | 2 | 15.9 | 16.2 | / | / |  |  |  |  | 4.23 | 204 |  | Kinga |  | 3 | 2 | 124346
2 | 15.0 | 43 | P/FINAL | 2025-09-04 | 08:35 | 2025-09-04 | 08:10 | ROLAND | 2 | 16.6 | 16.6 | / | / |  |  |  | 1.897 | 2.63 | 200 |  | Kinga |  | 3 | 2 | 124346
2 | 15.0 | 44 | P/FINAL | 2025-09-04 | 08:35 | 2025-09-04 | 08:10 | ROLAND | 2 | 16.0 | 16.1 | / | / |  |  |  |  | 2.01 | 182 |  | Kinga |  | 3 | 2 | 124346
3 | 35.0 | 45 | P/FINAL | 2025-09-04 | 08:36 | 2025-09-04 | 08:10 | ROLAND | 2 | 37.4 | 37.5 | / | / |  |  |  |  | 2.19 | 184 |  | Kinga |  | 3 | 3 | 124346
3 | 35.0 | 46 | P/FINAL | 2025-09-04 | 08:37 | 2025-09-04 | 08:10 | ROLAND | 2 | 37.2 | 37.3 | / | / |  |  |  |  | 1.35 | 183 |  | Kinga |  | 3 | 3 | 124346
3 | 35.0 | 47 | P/FINAL | 2025-09-04 | 08:38 | 2025-09-04 | 08:10 | ROLAND | 2 | 37.9 | 38.7 | / | / |  |  |  | 1.909 | 2.32 | 200 |  | Kinga |  | 3 | 3 | 124346
4 | 60.0 | 48 | P/FINAL | 2025-09-04 | 08:39 | 2025-09-04 | 08:10 | ROLAND | 2 | 58.7 | 59.0 | / | / |  |  |  |  | 1.45 | 198 |  | Kinga |  | 3 | 4 | 124346
4 | 60.0 | 49 | P/FINAL | 2025-09-04 | 08:39 | 2025-09-04 | 08:10 | ROLAND | 2 | 59.9 | 59.9 | / | / |  |  |  |  | 3.66 | 195 |  | Kinga |  | 3 | 4 | 124346
4 | 60.0 | 50 | P/FINAL | 2025-09-04 | 08:40 | 2025-09-04 | 08:10 | ROLAND | 2 | 62.3 | 62.4 | / | / |  |  |  | 1.907 | 1.98 | 201 |  | Kinga |  | 3 | 4 | 124346
1 | 0.0 | 51 | P/FINAL | 2025-09-04 | 09:04 | 2025-09-03 | 20:38 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 2.39 | 182 |  | Kinga |  | 3 | 1 | 124346
1 | 0.0 | 52 | P/FINAL | 2025-09-04 | 09:04 | 2025-09-03 | 20:38 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  | 1.852 | 1.13 | 214 |  | Kinga |  | 3 | 1 | 124346
2 | 15.0 | 53 | P/FINAL | 2025-09-04 | 09:08 | 2025-09-03 | 20:38 | ELENA | 2 | 15.4 | 15.4 | / | / |  |  |  |  | 1.16 | 207 |  | Kinga |  | 3 | 2 | 124346
2 | 15.0 | 54 | P/FINAL | 2025-09-04 | 09:09 | 2025-09-03 | 20:38 | ELENA | 2 | 15.3 | 15.2 | / | / |  |  |  |  | 1.27 | 212 |  | Kinga |  | 3 | 2 | 124346
2 | 15.0 | 55 | P/FINAL | 2025-09-04 | 09:10 | 2025-09-03 | 20:38 | ELENA | 2 | 15.3 | 15.2 | / | / |  |  |  | 1.860 | 1.01 | 207 |  | Kinga |  | 3 | 2 | 124346
3 | 35.0 | 56 | P/FINAL | 2025-09-04 | 09:15 | 2025-09-03 | 20:38 | ELENA | 2 | 36.9 | 36.9 | / | / |  |  |  |  | 2.73 | 210 |  | Kinga |  | 3 | 3 | 124346
3 | 35.0 | 57 | P/FINAL | 2025-09-04 | 09:16 | 2025-09-03 | 20:38 | ELENA | 2 | 36.7 | 36.8 | / | / |  |  |  | 1.926 | 1.22 | 198 |  | Kinga |  | 3 | 3 | 124346
3 | 35.0 | 58 | P/FINAL | 2025-09-04 | 09:16 | 2025-09-03 | 20:38 | ELENA | 2 | 36.0 | 36.6 | / | / |  |  |  |  | 2.89 | 198 |  | Kinga |  | 3 | 3 | 124346
4 | 60.0 | 59 | P/FINAL | 2025-09-04 | 09:31 | 2025-09-03 | 20:38 | ELENA | 2 | 62.2 | 62.7 | / | / |  |  |  |  | 0.92 | 205 |  | Kinga |  | 3 | 4 | 12434620
4 | 60.0 | 60 | P/FINAL | 2025-09-04 | 09:32 | 2025-09-03 | 20:38 | ELENA | 2 | 62.1 | 62.4 | / | / |  |  |  | 1.897 | 1.05 | 204 |  | Kinga |  | 3 | 4 | 12434620
4 | 60.0 | 61 | P/FINAL | 2025-09-04 | 09:33 | 2025-09-03 | 20:38 | ELENA | 2 | 63.5 | 63.5 | / | / |  |  |  |  | 0.99 | 201 |  | Kinga |  | 3 | 4 | 12434620
1 | 0.0 | 62 | P/FINAL | 2025-09-04 | 10:19 | 2025-09-04 | 09:44 | ROLAND | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 2.12 | 183 |  | Kinga |  | 3 | 1 | 124346
1 | 0.0 | 63 | P/FINAL | 2025-09-04 | 10:19 | 2025-09-04 | 09:44 | ROLAND | 2 | 0.0 | 0.0 | / | / |  |  |  | 1.867 | 2.47 | 191 |  | Kinga |  | 3 | 1 | 124346
2 | 15.0 | 64 | P/FINAL | 2025-09-04 | 10:30 | 2025-09-04 | 09:44 | ROLAND | 2 | 14.0 | 14.1 | / | / |  |  |  |  | 0.61 | 182 |  | Kinga |  | 3 | 2 | 124346
2 | 15.0 | 65 | P/FINAL | 2025-09-04 | 10:31 | 2025-09-04 | 09:44 | ROLAND | 2 | 13.2 | 13.2 | / | / |  |  |  | 1.907 | 3.19 | 191 |  | Kinga |  | 3 | 2 | 124346
2 | 15.0 | 66 | P/FINAL | 2025-09-04 | 10:31 | 2025-09-04 | 09:44 | ROLAND | 2 | 14.5 | 14.7 | / | / |  |  |  |  | 0.70 | 189 |  | Kinga |  | 3 | 2 | 124346
3 | 35.0 | 67 | P/FINAL | 2025-09-04 | 10:37 | 2025-09-04 | 09:44 | ROLAND | 2 | 33.8 | 33.7 | / | / |  |  |  |  | 1.23 | 195 |  | Kinga |  | 3 | 3 | 124346
3 | 35.0 | 68 | P/FINAL | 2025-09-04 | 10:38 | 2025-09-04 | 09:44 | ROLAND | 2 | 31.7 | 31.8 | / | / |  |  |  | 1.934 | 1.19 | 201 |  | Kinga |  | 3 | 3 | 124346
3 | 35.0 | 69 | P/FINAL | 2025-09-04 | 10:38 | 2025-09-04 | 09:44 | ROLAND | 2 | 35.3 | 35.6 | / | / |  |  |  |  | 1.68 | 198 |  | Kinga |  | 3 | 3 | 124346
4 | 60.0 | 70 | P/FINAL | 2025-09-04 | 10:41 | 2025-09-04 | 09:44 | ROLAND | 2 | 58.8 | 58.4 | / | / |  |  |  |  | 0.82 | 191 |  | Kinga |  | 3 | 4 | 124346
4 | 60.0 | 71 | P/FINAL | 2025-09-04 | 10:42 | 2025-09-04 | 09:44 | ROLAND | 2 | 59.2 | 59.3 | / | / |  |  |  |  | 1.15 | 201 |  | Kinga |  | 3 | 4 | 124346
4 | 60.0 | 72 | P/FINAL | 2025-09-04 | 10:42 | 2025-09-04 | 09:44 | ROLAND | 2 | 60.2 | 60.6 | / | / |  |  |  | 1.916 | 1.29 | 198 |  | Kinga |  | 3 | 4 | 124346
2 | 15.0 | 73 | P/FINAL | 2025-09-04 | 12:28 | 2025-09-04 | 12:14 | ROLAND | 2 | 14.4 | 14.8 | / | / |  |  |  |  | 6.25 | 200 |  | Kinga |  | 3 | 2 | 23858
2 | 15.0 | 74 | P/FINAL | 2025-09-04 | 12:29 | 2025-09-04 | 12:14 | ROLAND | 2 | 14.6 | 14.1 | / | / |  |  |  | 1.842 | 1.19 | 187 |  | Kinga |  | 3 | 2 | 23858
2 | 15.0 | 75 | P/FINAL | 2025-09-04 | 12:29 | 2025-09-04 | 12:14 | ROLAND | 2 | 14.3 | 14.4 | / | / |  |  |  |  | 4.26 | 204 |  | Kinga |  | 3 | 2 | 23858
3 | 35.0 | 76 | P/FINAL | 2025-09-04 | 12:30 | 2025-09-04 | 12:14 | ROLAND | 2 | 34.5 | 34.7 | / | / |  |  |  |  | 0.75 | 189 |  | Kinga |  | 3 | 3 | 23858
3 | 35.0 | 77 | P/FINAL | 2025-09-04 | 12:30 | 2025-09-04 | 12:14 | ROLAND | 2 | 35.7 | 35.7 | / | / |  |  |  | 1.860 | 1.43 | 198 |  | Kinga |  | 3 | 3 | 23858
3 | 35.0 | 78 | P/FINAL | 2025-09-04 | 12:31 | 2025-09-04 | 12:14 | ROLAND | 2 | 35.6 | 35.8 | / | / |  |  |  |  | 2.25 | 210 |  | Kinga |  | 3 | 3 | 23858
4 | 60.0 | 79 | P/FINAL | 2025-09-04 | 12:32 | 2025-09-04 | 12:14 | ROLAND | 2 | 58.7 | 58.7 | / | / |  |  |  |  | 1.05 | 197 |  | Kinga |  | 3 | 4 | 23858
4 | 60.0 | 80 | P/FINAL | 2025-09-04 | 12:34 | 2025-09-04 | 12:14 | ROLAND | 2 | 60.8 | 61.1 | / | / |  |  |  |  | 1.25 | 202 |  | Kinga |  | 3 | 4 | 23858
4 | 60.0 | 81 | P/FINAL | 2025-09-04 | 12:35 | 2025-09-04 | 12:14 | ROLAND | 2 | 63.2 | 63.2 | / | / |  |  |  | 1.893 | 1.15 | 206 |  | Kinga |  | 3 | 4 | 23858
1 | 0.0 | 82 | P/FINAL | 2025-09-04 | 12:40 | 2025-09-04 | 12:14 | ROLAND | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 2.00 | 198 |  | Kinga |  | 3 | 1 | 23858
1 | 0.0 | 83 | P/FINAL | 2025-09-04 | 12:40 | 2025-09-04 | 12:14 | ROLAND | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 1.55 | 197 |  | Kinga |  | 3 | 1 | 23858
1 | 0.0 | 84 | P/FINAL | 2025-09-04 | 14:58 | 2025-09-04 | 14:35 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  | 1.855 | 0.93 | 185 |  | Kinga |  | 3 | 1 | 124346
1 | 0.0 | 85 | P/FINAL | 2025-09-04 | 14:59 | 2025-09-04 | 14:35 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 2.77 | 189 |  | Kinga |  | 3 | 1 | 124346
2 | 15.0 | 86 | P/FINAL | 2025-09-04 | 15:00 | 2025-09-04 | 14:35 | ELENA | 2 | 14.3 | 14.4 | / | / |  |  |  |  | 1.07 | 184 |  | Kinga |  | 3 | 2 | 124346
2 | 15.0 | 87 | P/FINAL | 2025-09-04 | 15:00 | 2025-09-04 | 14:35 | ELENA | 2 | 13.8 | 14.1 | / | / |  |  |  |  | 1.27 | 194 |  | Kinga |  | 3 | 2 | 124346
2 | 15.0 | 88 | P/FINAL | 2025-09-04 | 15:01 | 2025-09-04 | 14:35 | ELENA | 2 | 14.0 | 14.3 | / | / |  |  |  | 1.876 | 1.11 | 185 |  | Kinga |  | 3 | 2 | 124346
3 | 35.0 | 89 | P/FINAL | 2025-09-04 | 15:04 | 2025-09-04 | 14:35 | ELENA | 2 | 35.9 | 36.0 | / | / |  |  |  |  | 1.47 | 190 |  | Kinga |  | 3 | 3 | 124346
3 | 35.0 | 90 | P/FINAL | 2025-09-04 | 15:04 | 2025-09-04 | 14:35 | ELENA | 2 | 36.1 | 36.9 | / | / |  |  |  |  | 1.25 | 194 |  | Kinga |  | 3 | 3 | 124346
3 | 35.0 | 91 | P/FINAL | 2025-09-04 | 15:05 | 2025-09-04 | 14:35 | ELENA | 2 | 35.6 | 35.9 | / | / |  |  |  | 1.877 | 1.21 | 185 |  | Kinga |  | 3 | 3 | 124346
4 | 60.0 | 92 | P/FINAL | 2025-09-04 | 15:11 | 2025-09-04 | 14:35 | ELENA | 2 | 60.4 | 60.4 | / | / |  |  |  |  | 1.41 | 190 |  | Kinga |  | 3 | 4 | 124346
4 | 60.0 | 93 | P/FINAL | 2025-09-04 | 15:12 | 2025-09-04 | 14:35 | ELENA | 2 | 59.8 | 60.6 | / | / |  |  |  |  | 0.85 | 196 |  | Kinga |  | 3 | 4 | 124346
4 | 60.0 | 94 | P/FINAL | 2025-09-04 | 15:12 | 2025-09-04 | 14:35 | ELENA | 2 | 60.6 | 60.9 | / | / |  |  |  | 1.848 | 2.58 | 190 |  | Kinga |  | 3 | 4 | 124346
2 | 15.0 | 95 | P/FINAL | 2025-09-04 | 16:03 | 2025-09-04 | 15:48 | ELENA | 2 | 14.1 | 14.2 | / | / |  |  |  |  | 3.19 | 195 |  | Kinga |  | 3 | 2
2 | 15.0 | 96 | P/FINAL | 2025-09-04 | 16:03 | 2025-09-04 | 15:48 | ELENA | 2 | 15.1 | 15.5 | / | / |  |  |  |  | 3.24 | 194 |  | Kinga |  | 3 | 2
2 | 15.0 | 97 | P/FINAL | 2025-09-04 | 16:03 | 2025-09-04 | 15:48 | ELENA | 2 | 14.4 | 14.8 | / | / |  |  |  | 1.851 | 1.57 | 190 |  | Kinga |  | 3 | 2
3 | 35.0 | 98 | P/FINAL | 2025-09-04 | 16:07 | 2025-09-04 | 15:48 | ELENA | 2 | 32.2 | 32.0 | / | / |  |  |  |  | 1.44 | 199 |  | Kinga |  | 3 | 3
3 | 35.0 | 99 | P/FINAL | 2025-09-04 | 16:08 | 2025-09-04 | 15:48 | ELENA | 2 | 33.6 | 33.4 | / | / |  |  |  |  | 1.44 | 191 |  | Kinga |  | 3 | 3
3 | 35.0 | 100 | P/FINAL | 2025-09-04 | 16:08 | 2025-09-04 | 15:48 | ELENA | 2 | 34.7 | 34.9 | / | / |  |  |  | 1.893 | 1.55 | 198 |  | Kinga |  | 3 | 3
4 | 60.0 | 101 | P/FINAL | 2025-09-04 | 16:09 | 2025-09-04 | 15:48 | ELENA | 2 | 57.9 | 57.5 | / | / |  |  |  |  | 0.84 | 197 |  | Kinga |  | 3 | 4
4 | 60.0 | 102 | P/FINAL | 2025-09-04 | 16:09 | 2025-09-04 | 15:48 | ELENA | 2 | 58.9 | 58.3 | / | / |  |  |  |  | 0.84 | 196 |  | Kinga |  | 3 | 4
4 | 60.0 | 103 | P/FINAL | 2025-09-04 | 16:09 | 2025-09-04 | 15:48 | ELENA | 2 | 60.1 | 60.4 | / | / |  |  |  | 1.877 | 0.91 | 200 |  | Kinga |  | 3 | 4
2 | 15.0 | 104 | Old A | 2025-09-05 | 06:20 | 2025-09-05 | 06:08 | ROLAND | 2 | 15.1 | 14.9 | / | / |  |  |  |  | 2.74 | 220 |  | Kinga |  | 3 | 2 | 112146
2 | 15.0 | 105 | Old A | 2025-09-05 | 06:21 | 2025-09-05 | 06:08 | ROLAND | 2 | 15.5 | 15.4 | / | / |  |  |  |  | 2.05 | 228 |  | Kinga |  | 3 | 3 | 112146
3 | 35.0 | 106 | Old A | 2025-09-05 | 06:25 | 2025-09-05 | 06:08 | ROLAND | 2 | 35.9 | 35.9 | / | / |  |  |  |  | 1.09 | 212 |  | Kinga |  | 3 | 3 | 112146
3 | 35.0 | 107 | Old A | 2025-09-05 | 06:25 | 2025-09-05 | 06:08 | ROLAND | 2 | 36.2 | 36.0 | / | / |  |  |  |  | 2.22 | 205 |  | Kinga |  | 3 | 3 | 112146
4 | 60.0 | 108 | Old A | 2025-09-05 | 06:27 | 2025-09-05 | 06:08 | ROLAND | 2 | 55.2 | 57.4 | / | / |  |  |  |  | 1.88 | 217 |  | Kinga |  | 3 | 4 | 112146
4 | 60.0 | 109 | Old A | 2025-09-05 | 06:28 | 2025-09-05 | 06:08 | ROLAND | 2 | 64.6 | 63.5 | / | / |  |  |  |  | 2.44 | 206 |  | Kinga |  | 3 | 4 | 112146
1 | 0.0 | 110 | P/FINAL | 2025-09-05 | 07:30 | 2025-09-05 | 07:08 | ROLAND | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 3.17 | 217 |  | Kinga |  | 3 | 1 | 112146
1 | 0.0 | 111 | P/FINAL | 2025-09-05 | 07:30 | 2025-09-05 | 07:08 | ROLAND | 2 | 0.0 | 0.0 | / | / |  |  |  | 1.896 | 2.83 | 228 |  | Kinga |  | 3 | 1 | 112146
2 | 15.0 | 112 | P/FINAL | 2025-09-05 | 07:32 | 2025-09-05 | 07:08 | ROLAND | 2 | 14.6 | 14.7 | / | / |  |  |  |  | 1.46 | 06 |  | Kinga |  | 3 | 2 | 112146
2 | 15.0 | 113 | P/FINAL | 2025-09-05 | 07:32 | 2025-09-05 | 07:08 | ROLAND | 2 | 14.9 | 15.5 | / | / |  |  |  |  | 4.33 | 210 |  | Kinga |  | 3 | 2 | 112146
2 | 15.0 | 114 | P/FINAL | 2025-09-05 | 07:33 | 2025-09-05 | 07:08 | ROLAND | 2 | 15.6 | 15.4 | / | / |  |  |  | 1.872 | 4.67 | 205 |  | Kinga |  | 3 | 2 | 112146
3 | 35.0 | 115 | P/FINAL | 2025-09-05 | 07:34 | 2025-09-05 | 07:08 | ROLAND | 2 | 33.5 | 33.7 | / | / |  |  |  |  | 1.94 | 211 |  | Kinga |  | 3 | 3 | 112146
3 | 35.0 | 116 | P/FINAL | 2025-09-05 | 07:34 | 2025-09-05 | 07:08 | ROLAND | 2 | 35.1 | 35.0 | / | / |  |  |  |  | 1.25 | 219 |  | Kinga |  | 3 | 3 | 112146
3 | 35.0 | 117 | P/FINAL | 2025-09-05 | 07:35 | 2025-09-05 | 07:08 | ROLAND | 2 | 35.4 | 35.5 | / | / |  |  |  | 1.883 | 3.09 | 211 |  | Kinga |  | 3 | 3 | 112146
4 | 60.0 | 118 | P/FINAL | 2025-09-05 | 07:40 | 2025-09-05 | 07:08 | ROLAND | 2 | 56.4 | 55.9 | / | / |  |  |  |  | 2.27 | 205 |  | Kinga |  | 3 | 4 | 112146
4 | 60.0 | 119 | P/FINAL | 2025-09-05 | 07:40 | 2025-09-05 | 07:08 | ROLAND | 2 | 57.9 | 57.7 | / | / |  |  |  |  | 2.64 | 217 |  | Kinga |  | 3 | 4 | 112146
4 | 60.0 | 120 | P/FINAL | 2025-09-05 | 07:41 | 2025-09-05 | 07:08 | ROLAND | 2 | 59.2 | 59.2 | / | / |  |  |  | 1.905 | 4.38 | 210 |  | Kinga |  | 3 | 4 | 112146
2 | 15.0 | 121 | P/FINAL | 2025-09-05 | 09:36 | 2025-09-05 | 09:14 | ROLAND | 2 | 14.2 | 14.2 | / | / |  |  |  |  | 0.78 | 201 |  | Andrea V |  | 3 | 2 | 112146
2 | 15.0 | 122 | P/FINAL | 2025-09-05 | 09:39 | 2025-09-05 | 09:14 | ROLAND | 2 | 13.9 | 13.8 | / | / |  |  |  | 1.879 | 0.71 | 211 |  | Andrea V |  | 3 | 2 | 112146
2 | 15.0 | 123 | P/FINAL | 2025-09-05 | 09:41 | 2025-09-05 | 09:14 | ROLAND | 2 | 13.8 | 13.7 | / | / |  |  |  |  | 0.79 | 197 |  | Andrea V |  | 3 | 2 | 112146
3 | 35.0 | 124 | P/FINAL | 2025-09-05 | 09:44 | 2025-09-05 | 09:14 | ROLAND | 2 | 37.1 | 36.5 | / | / |  |  |  |  | 1.15 | 201 |  | Andrea V |  | 3 | 2 | 112146
3 | 35.0 | 125 | P/FINAL | 2025-09-05 | 09:47 | 2025-09-05 | 09:14 | ROLAND | 2 | 36.1 | 35.5 | / | / |  |  |  | 1.886 | 0.93 | 187 |  | Andrea V |  | 3 | 3 | 112146
3 | 35.0 | 126 | P/FINAL | 2025-09-05 | 09:47 | 2025-09-05 | 09:14 | ROLAND | 2 | 38.3 | 37.7 | / | / |  |  |  |  | 0.78 | 187 |  | Andrea V |  | 3 | 3 | 112146
4 | 60.0 | 127 | P/FINAL | 2025-09-05 | 09:50 | 2025-09-05 | 09:14 | ROLAND | 2 | 59.8 | 59.1 | / | / |  |  |  |  | 1.00 | 203 |  | Andrea V |  | 3 | 4 | 112146
4 | 60.0 | 128 | P/FINAL | 2025-09-05 | 09:50 | 2025-09-05 | 09:14 | ROLAND | 2 | 62.2 | 61.3 | / | / |  |  |  | 1.876 | 1.15 | 198 |  | Andrea V |  | 3 | 4 | 112146
4 | 60.0 | 129 | P/FINAL | 2025-09-05 | 09:52 | 2025-09-05 | 09:14 | ROLAND | 2 | 62.2 | 61.0 | / | / |  |  |  |  | 0.91 | 192 |  | Andrea V |  | 3 | 4 | 112146
1 | 0.0 | 130 | P/FINAL | 2025-09-05 | 09:54 | 2025-09-05 | 09:14 | ROLAND | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 1.88 | 195 |  | Andrea V |  | 3 | 1 | 112146
1 | 0.0 | 131 | P/FINAL | 2025-09-05 | 09:55 | 2025-09-05 | 09:14 | ROLAND | 2 | 0.0 | 0.0 | / | / |  |  |  | 1.898 | 1.88 | 207 |  | Andrea V |  | 3 | 1 | 112146
1 | 0.0 | 132 | P/FINAL | 2025-09-05 | 11:58 | 2025-09-05 | 11:44 | ROLAND | 2 | 0.0 | 0.0 | / | / |  |  |  | 1.869 | 1.72 | 198 |  | Kinga |  | 3 | 1 | 112146
1 | 0.0 | 133 | Old A | 2025-09-05 | 11:59 | 2025-09-05 | 11:44 | ROLAND | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 1.02 | 195 |  | Kinga |  | 3 | 1 | 112146
2 | 15.0 | 134 | P/FINAL | 2025-09-05 | 12:00 | 2025-09-05 | 11:44 | ROLAND | 2 | 14.7 | 14.6 | / | / |  |  |  | 1.929 | 0.80 | 191 |  | Kinga |  | 3 | 2 | 112146
2 | 15.0 | 135 | P/FINAL | 2025-09-05 | 12:01 | 2025-09-05 | 11:44 | ROLAND | 2 | 13.6 | 13.7 | / | / |  |  |  |  | 2.79 | 195 |  | Kinga |  | 3 | 2 | 112146
2 | 15.0 | 136 | P/FINAL | 2025-09-05 | 12:02 | 2025-09-05 | 11:44 | ROLAND | 2 | 14.0 | 14.1 | / | / |  |  |  |  | 4.52 | 198 |  | Kinga |  | 3 | 2 | 112146
3 | 35.0 | 137 | P/FINAL | 2025-09-05 | 12:02 | 2025-09-05 | 11:44 | ROLAND | 2 | 33.2 | 33.3 | / | / |  |  |  | 1.855 | 0.82 | 197 |  | Kinga |  | 3 | 3 | 112146
3 | 35.0 | 138 | P/FINAL | 2025-09-05 | 12:04 | 2025-09-05 | 11:44 | ROLAND | 2 | 34.5 | 34.7 | / | / |  |  |  |  | 0.93 | 195 |  | Kinga |  | 3 | 3 | 112146
3 | 35.0 | 139 | P/FINAL | 2025-09-05 | 12:04 | 2025-09-05 | 11:44 | ROLAND | 2 | 35.3 | 35.3 | / | / |  |  |  |  | 0.93 | 199 |  | Kinga |  | 3 | 3 | 112146
4 | 60.0 | 140 | P/FINAL | 2025-09-05 | 12:10 | 2025-09-05 | 11:44 | ROLAND | 2 | 57.3 | 55.9 | / | / |  |  |  | 1.917 | 0.86 | 196 |  | Kinga |  | 3 | 4 | 112146
4 | 60.0 | 141 | P/FINAL | 2025-09-05 | 12:10 | 2025-09-05 | 11:44 | ROLAND | 2 | 60.9 | 61.4 | / | / |  |  |  |  | 1.10 | 199 |  | Kinga |  | 3 | 4 | 112146
4 | 60.0 | 142 | P/FINAL | 2025-09-05 | 12:11 | 2025-09-05 | 11:44 | ROLAND | 2 | 62.6 | 62.8 | / | / |  |  |  |  | 2.08 | 201 |  | Kinga |  | 3 | 4 | 112146
1 | 0.0 | 143 | P/FINAL | 2025-09-05 | 14:58 | 2025-09-05 | 14:38 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 0.83 |  |  | Andrea V |  | 3 | 1 | 112146
1 | 0.0 | 144 | P/FINAL | 2025-09-05 | 14:58 | 2025-09-05 | 14:38 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  | 1.950 | 0.93 |  |  | Andrea V |  | 3 | 1 | 112146
2 | 15.0 | 145 | P/FINAL | 2025-09-05 | 15:05 | 2025-09-05 | 14:38 | ELENA | 2 | 14.7 | 14.7 | / | / |  |  |  |  | 0.70 | 186 |  | Andrea V |  | 3 | 2 | 112146
3 | 35.0 | 146 | P/FINAL | 2025-09-05 | 15:08 | 2025-09-05 | 14:38 | ELENA | 2 | 37.0 | 36.0 | / | / |  |  |  | 1.873 | 0.93 | 203 |  | Andrea V |  | 3 | 3 | 112146
2 | 15.0 | 147 | P/FINAL | 2025-09-05 | 15:09 | 2025-09-05 | 14:38 | ELENA | 2 | 14.5 | 14.4 | / | / |  |  |  |  | 0.73 | 191 |  | Andrea V |  | 3 | 2 | 112146
2 | 15.0 | 148 | P/FINAL | 2025-09-05 | 15:11 | 2025-09-05 | 14:38 | ELENA | 2 | 15.4 | 15.3 | / | / |  |  |  |  | 0.93 | 202 |  | Andrea V |  | 3 | 2 | 112146
3 | 35.0 | 149 | P/FINAL | 2025-09-05 | 15:13 | 2025-09-05 | 14:38 | ELENA | 2 | 38.4 | 38.3 | / | / |  |  |  | 1.857 | 0.86 | 186 |  | Andrea V |  | 3 | 3 | 112146
3 | 35.0 | 150 | P/FINAL | 2025-09-05 | 15:13 | 2025-09-05 | 14:38 | ELENA | 2 | 37.9 | 37.4 | / | / |  |  |  |  | 0.69 | 155 |  | Andrea V |  | 3 | 3 | 112146
4 | 60.0 | 151 | P/FINAL | 2025-09-05 | 15:14 | 2025-09-05 | 14:38 | ELENA | 2 | 62.5 | 60.5 | / | / |  |  |  | 1.879 | 0.90 | 200 |  | Andrea V |  | 3 | 4 | 112146
4 | 60.0 | 152 | P/FINAL | 2025-09-05 | 15:15 | 2025-09-05 | 14:38 | ELENA | 2 | 61.9 | 60.9 | / | / |  |  |  |  | 0.82 | 196 |  | Andrea V |  | 3 | 4 | 112146
4 | 60.0 | 153 | P/FINAL | 2025-09-05 | 15:16 | 2025-09-05 | 14:38 | ELENA | 2 | 59.7 | 59.0 | / | / |  |  |  |  | 0.80 | 194 |  | Andrea V |  | 3 | 4 | 112146
2 | 15.0 | 154 | P/FINAL | 2025-09-05 | 15:59 | 2025-09-05 | 15:35 | ELENA | 2 | 14.3 | 14.5 | / | / |  |  |  |  | 1.00 | 209 |  | Andrea V |  | 3 | 2
2 | 15.0 | 155 | P/FINAL | 2025-09-05 | 16:00 | 2025-09-05 | 15:35 | ELENA | 2 | 14.5 | 14.9 | / | / |  |  |  |  | 0.60 | 192 |  | Andrea V |  | 3 | 2
2 | 15.0 | 156 | P/FINAL | 2025-09-05 | 16:01 | 2025-09-05 | 15:35 | ELENA | 2 | 14.3 | 14.5 | / | / |  |  |  | 1.875 | 0.97 | 197 |  | Andrea V |  | 3 | 2
3 | 35.0 | 157 | P/FINAL | 2025-09-05 | 16:02 | 2025-09-05 | 15:35 | ELENA | 2 | 35.6 | 36.2 | / | / |  |  |  |  | 0.89 | 193 |  | Andrea V |  | 3 | 3 | 112146
3 | 35.0 | 158 | P/FINAL | 2025-09-05 | 16:03 | 2025-09-05 | 15:35 | ELENA | 2 | 35.2 | 35.5 | / | / |  |  |  | 1.898 | 0.90 | 193 |  | Andrea V |  | 3 | 3 | 112146
3 | 35.0 | 159 | P/FINAL | 2025-09-05 | 16:03 | 2025-09-05 | 15:35 | ELENA | 2 | 35.6 | 35.3 | / | / |  |  |  |  | 0.85 | 199 |  | Andrea V |  | 3 | 3 | 112146
4 | 60.0 | 160 | P/FINAL | 2025-09-05 | 16:04 | 2025-09-05 | 15:35 | ELENA | 2 | 60.7 | 61.8 | / | / |  |  |  |  | 1.20 | 198 |  | Andrea V |  | 3 | 4 | 112146
4 | 60.0 | 161 | P/FINAL | 2025-09-05 | 16:04 | 2025-09-05 | 15:35 | ELENA | 2 | 59.4 | 58.3 | / | / |  |  |  | 1.905 | 0.99 | 197 |  | Andrea V |  | 3 | 4 | 112146
4 | 60.0 | 162 | P/FINAL | 2025-09-05 | 16:05 | 2025-09-05 | 15:35 | ELENA | 2 | 61.5 | 60.4 | / | / |  |  |  |  | 1.02 | 201 |  | Andrea V |  | 3 | 4 | 112146
1 | 0.0 | 163 | P/FINAL | 2025-09-08 | 08:35 | 2025-09-08 | 07:35 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 0.84 | 198 |  | Andrea V |  | 3 | 1
1 | 0.0 | 164 | P/FINAL | 2025-09-08 | 08:37 | 2025-09-08 | 07:35 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  | 1.951 | 0.63 | 205 |  | Andrea V |  | 3 | 1
2 | 15.0 | 165 | P/FINAL | 2025-09-08 | 08:38 | 2025-09-08 | 07:35 | ELENA | 2 | 16.1 | 16.1 | / | / |  |  |  |  | 1.13 | 192 |  | Andrea V |  | 3 | 2 | 112146
2 | 15.0 | 166 | P/FINAL | 2025-09-08 | 08:39 | 2025-09-08 | 07:35 | ELENA | 2 | 16.7 | 16.6 | / | / |  |  |  | 1.879 | 0.56 | 200 |  | Andrea V |  | 3 | 2 | 112146
2 | 15.0 | 167 | P/FINAL | 2025-09-08 | 08:40 | 2025-09-08 | 07:35 | ELENA | 2 | 17.7 | 17.7 | / | / |  |  |  |  | 0.61 | 211 |  | Andrea V |  | 3 | 2 | 112146
3 | 35.0 | 168 | P/FINAL | 2025-09-08 | 08:41 | 2025-09-08 | 07:35 | ELENA | 2 | 36.8 | 36.6 | / | / |  |  |  |  | 0.98 | 206 |  | Andrea V |  | 3 | 3 | 112146
3 | 35.0 | 169 | P/FINAL | 2025-09-08 | 08:42 | 2025-09-08 | 07:35 | ELENA | 2 | 38.2 | 38.2 | / | / |  |  |  | 1.906 | 0.69 | 214 |  | Andrea V |  | 3 | 3 | 112146
3 | 35.0 | 170 | P/FINAL | 2025-09-08 | 08:46 | 2025-09-08 | 07:35 | ELENA | 2 | 38.2 | 38.3 | / | / |  |  |  | - | 1.46 | 204 |  | Andrea V |  | 3 | 3 | 112146
4 | 60.0 | 171 | P/FINAL | 2025-09-08 | 08:47 | 2025-09-08 | 07:35 | ELENA | 2 | 63.4 | 63.0 | / | / |  |  |  | 1.914 | 0.92 | 185 |  | Andrea V |  | 3 | 4 | 112146
4 | 60.0 | 172 | P/FINAL | 2025-09-08 | 08:49 | 2025-09-08 | 07:35 | ELENA | 2 | 63.5 | 63.5 | / | / |  |  |  |  | 0.95 | 207 |  | Andrea V |  | 3 | 4 | 112146
3 | 35.0 | 173 | P/FINAL | 2025-09-08 | 09:04 | 2025-09-08 | 07:35 | ELENA | 2 | 36.7 | 37.4 | / | / |  |  |  |  | 0.35 | 197 |  | Andrea V |  | 3 | 3 | 112146
2 | 15.0 | 174 | P/FINAL | 2025-09-08 | 11:04 | 2025-09-08 | 10:26 | ELENA | 2 | 14.9 | 14.8 | / | / |  |  |  |  | 0.80 | 188 |  | Andrea V |  | 3 | 2 | 112146
2 | 15.0 | 175 | P/FINAL | 2025-09-08 | 11:04 | 2025-09-08 | 10:26 | ELENA | 2 | 13.7 | 13.5 | / | / |  |  |  | 1.835 | 0.48 | 188 |  | Andrea V |  | 3 | 2 | 112146
2 | 15.0 | 176 | P/FINAL | 2025-09-08 | 11:06 | 2025-09-08 | 10:26 | ELENA | 2 | 14.9 | 14.7 | / | / |  |  |  |  | 0.66 | 201 |  | Andrea V |  | 3 | 2 | 112146
3 | 35.0 | 177 | P/FINAL | 2025-09-08 | 11:07 | 2025-09-08 | 10:26 | ELENA | 2 | 34.8 | 34.4 | / | / |  |  |  |  | 0.67 | 192 |  | Andrea V |  | 3 | 3 | 112146
3 | 35.0 | 178 | P/FINAL | 2025-09-08 | 11:08 | 2025-09-08 | 10:26 | ELENA | 2 | 35.8 | 35.4 | / | / |  |  |  | 1.884 | 1.11 | 201 |  | Andrea V |  | 3 | 3 | 112146
3 | 35.0 | 179 | P/FINAL | 2025-09-08 | 11:08 | 2025-09-08 | 10:26 | ELENA | 2 | 36.9 | 36.1 | / | / |  |  |  |  | 0.81 | 195 |  | Andrea V |  | 3 | 3 | 112146
4 | 60.0 | 180 | P/FINAL | 2025-09-08 | 11:11 | 2025-09-08 | 10:26 | ELENA | 2 | 59.9 | 59.8 | / | / |  |  |  |  | 1.48 | 198 |  | Andrea V |  | 3 | 4 | 112146
4 | 60.0 | 181 | P/FINAL | 2025-09-08 | 11:12 | 2025-09-08 | 10:26 | ELENA | 2 | 62.8 | 62.6 | / | / |  |  |  | 1.899 | 0.96 | 194 |  | Andrea V |  | 3 | 4 | 112146
4 | 60.0 | 182 | P/FINAL | 2025-09-08 | 11:13 | 2025-09-08 | 10:26 | ELENA | 2 | 59.4 | 59.6 | / | / |  |  |  |  | 0.79 | 183 |  | Andrea V |  | 3 | 4 | 112146
1 | 0.0 | 183 | P/FINAL | 2025-09-08 | 11:15 | 2025-09-08 | 10:26 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  | 1.934 | 0.64 | 193 |  | Andrea V |  | 3 | 1 | 112146
1 | 0.0 | 184 | P/FINAL | 2025-09-08 | 11:15 | 2025-09-08 | 10:26 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 0.55 | 196 |  | Andrea V |  | 3 | 1 | 112146
2 | 15.0 | 185 | Old A | 2025-09-08 | 13:32 | 2025-09-08 | 13:06 | Elena | 2 | 13.6 | 13.6 | / | / |  |  |  |  | 7.64 | 189 |  | Kinga |  | 3 | 2 | 112146
2 | 15.0 | 186 | Old A | 2025-09-08 | 13:33 | 2025-09-08 | 13:06 | Elena | 2 | 14.6 | 14.5 | / | / |  |  |  |  | 5.74 | 199 |  | Kinga |  | 3 | 2 | 112146
3 | 35.0 | 187 | Old A | 2025-09-08 | 13:33 | 2025-09-08 | 13:06 | Elena | 2 | 31.3 | 31.7 | / | / |  |  |  |  | 3.15 | 203 |  | Kinga |  | 3 | 3 | 112146
3 | 35.0 | 188 | Old A | 2025-09-08 | 13:34 | 2025-09-08 | 13:06 | Elena | 2 | 31.7 | 31.9 | / | / |  |  |  | 1.930 | 5.63 | 180 |  | Kinga |  | 3 | 3 | 112146
4 | 60.0 | 189 | Old A | 2025-09-08 | 13:34 | 2025-09-08 | 13:06 | Elena | 2 | 55.3 | 55.4 | / | / |  |  |  |  | 1.39 | 188 |  | Kinga |  | 3 | 4 | 112146
4 | 60.0 | 190 | Old A | 2025-09-08 | 13:34 | 2025-09-08 | 13:06 | Elena | 2 | 55.3 | 55.3 | / | / |  |  |  |  | 2.19 | 187 |  | Kinga |  | 3 | 4 | 112146
2 | 15.0 | 191 | P/FINAL | 2025-09-08 | 13:51 | 2025-09-08 | 13:06 | ELENA | 2 | 14.1 | 14.0 | / | / |  |  |  | 1.926 | 1.05 | 179 |  | Kinga |  | 3 | 2 | 112146
2 | 15.0 | 192 | P/FINAL | 2025-09-08 | 13:51 | 2025-09-08 | 13:06 | ELENA | 2 | 14.8 | 14.8 | / | / |  |  |  |  | 2.21 | 189 |  | Kinga |  | 3 | 2 | 112146
2 | 15.0 | 193 | P/FINAL | 2025-09-08 | 13:52 | 2025-09-08 | 13:06 | ELENA | 2 | 14.9 | 14.8 | / | / |  |  |  |  | 3.47 | 192 |  | Kinga |  | 3 | 2 | 112146
3 | 35.0 | 194 | P/FINAL | 2025-09-08 | 13:53 | 2025-09-08 | 13:06 | ELENA | 2 | 36.7 | 35.9 | / | / |  |  |  |  | 1.71 | 203 |  | Kinga |  | 3 | 3 | 112146
3 | 35.0 | 195 | P/FINAL | 2025-09-08 | 13:53 | 2025-09-08 | 13:06 | ELENA | 2 | 36.7 | 35.9 | / | / |  |  |  |  | 1.12 | 180 |  | Kinga |  | 3 | 3 | 112146
3 | 35.0 | 196 | P/FINAL | 2025-09-08 | 13:54 | 2025-09-08 | 13:06 | ELENA | 2 | 35.9 | 36.1 | / | / |  |  |  |  | 3.48 | 188 |  | Kinga |  | 3 | 3 | 112146
4 | 60.0 | 197 | P/FINAL | 2025-09-08 | 13:56 | 2025-09-08 | 13:06 | ELENA | 2 | 58.0 | 62.0 | / | / |  |  |  | 1.956 | 1.08 | 187 |  | Kinga |  | 3 | 4 | 112146
4 | 60.0 | 198 | P/FINAL | 2025-09-08 | 13:57 | 2025-09-08 | 13:06 | ELENA | 2 | 59.8 | 61.0 | / | / |  |  |  |  | 1.51 | 205 |  | Kinga |  | 3 | 4 | 112146
4 | 60.0 | 199 | P/FINAL | 2025-09-08 | 13:58 | 2025-09-08 | 13:06 | ELENA | 2 | 60.5 | 60.7 | / | / |  |  |  |  | 3.18 | 187 |  | Kinga |  | 3 | 4 | 112146
1 | 0.0 | 200 | P/FINAL | 2025-09-08 | 13:59 | 2025-09-08 | 13:06 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  | 3.020 | 1.99 | 204 |  | Kinga |  | 3 | 1 | 112146
1 | 0.0 | 201 | P/FINAL | 2025-09-08 | 13:59 | 2025-09-08 | 13:06 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 1.22 | 179 |  | Kinga |  | 3 | 1 | 112146
2 | 15.0 | 202 | P/FINAL | 2025-09-08 | 15:15 | 2025-09-08 | 14:48 | ROLAND | 2 | 15.6 | 16.5 | / | / |  |  |  |  | 2.34 | 212 |  | Kinga |  | 3 | 2 | 112146
2 | 15.0 | 203 | P/FINAL | 2025-09-08 | 15:15 | 2025-09-08 | 14:48 | ROLAND | 2 | 15.2 | 14.7 | / | / |  |  |  |  | 0.72 | 213 |  | Kinga |  | 3 | 2 | 112146
2 | 15.0 | 204 | P/FINAL | 2025-09-08 | 15:16 | 2025-09-08 | 14:48 | ROLAND | 2 | 14.8 | 16.4 | / | / |  |  |  | 1.994 | 2.74 | 185 |  | Kinga |  | 3 | 2 | 112146
3 | 35.0 | 205 | P/FINAL | 2025-09-08 | 15:17 | 2025-09-08 | 14:48 | ROLAND | 2 | 34.0 | 34.4 | / | / |  |  |  |  | 5.79 | 190 |  | Kinga |  | 3 | 3 | 112146
3 | 35.0 | 206 | P/FINAL | 2025-09-08 | 15:18 | 2025-09-08 | 14:48 | ROLAND | 2 | 34.6 | 34.4 | / | / |  |  |  |  | 1.76 | 197 |  | Kinga |  | 3 | 3 | 112146
3 | 35.0 | 207 | P/FINAL | 2025-09-08 | 15:19 | 2025-09-08 | 14:48 | ROLAND | 2 | 34.0 | 34.1 | / | / |  |  |  |  | 4.62 | 187 |  | Kinga |  | 3 | 3 | 112146
4 | 60.0 | 208 | P/FINAL | 2025-09-08 | 15:19 | 2025-09-08 | 14:48 | ROLAND | 2 | 55.6 | 55.6 | / | / |  |  |  | 1.986 | 1.25 | 190 |  | Kinga |  | 3 | 4 | 112146
4 | 60.0 | 209 | P/FINAL | 2025-09-08 | 15:21 | 2025-09-08 | 14:48 | ROLAND | 2 | 57.2 | 57.2 | / | / |  |  |  |  | 1.48 | 199 |  | Kinga |  | 3 | 4 | 112146
4 | 60.0 | 210 | P/FINAL | 2025-09-08 | 15:21 | 2025-09-08 | 14:48 | ROLAND | 2 | 58.7 | 58.9 | / | / |  |  |  | 1.957 | 1.18 | 197 |  | Kinga |  | 3 | 4 | 112146
2 | 15.0 | 211 | P/FINAL | 2025-09-08 | 16:08 | 2025-09-08 | 15:32 | ROLAND | 2 | 15.4 | 15.1 | / | / |  |  |  |  | 2.24 | 193 |  | Kinga |  | 3 | 2
2 | 15.0 | 212 | P/FINAL | 2025-09-08 | 16:08 | 2025-09-08 | 15:32 | ROLAND | 2 | 14.9 | 14.8 | / | / |  |  |  |  | 1.65 | 197 |  | Kinga |  | 3 | 2
2 | 15.0 | 213 | P/FINAL | 2025-09-08 | 16:09 | 2025-09-08 | 15:32 | ROLAND | 2 | 15.2 | 15.4 | / | / |  |  |  | 1.962 | 2.01 | 197 |  | Kinga |  | 3 | 2
3 | 35.0 | 214 | P/FINAL | 2025-09-08 | 16:09 | 2025-09-08 | 15:32 | ROLAND | 2 | 32.3 | 33.4 | / | / |  |  |  |  | 1.15 | 190 |  | Kinga |  | 3 | 3
3 | 35.0 | 215 | P/FINAL | 2025-09-08 | 16:09 | 2025-09-08 | 15:32 | ROLAND | 2 | 34.6 | 34.9 | / | / |  |  |  | 1.988 | 1.13 | 202 |  | Kinga |  | 3 | 3
3 | 35.0 | 216 | P/FINAL | 2025-09-08 | 16:09 | 2025-09-08 | 15:32 | ROLAND | 2 | 35.6 | 35.9 | / | / |  |  |  |  | 1.99 | 190 |  | Kinga |  | 3 | 3
4 | 60.0 | 217 | P/FINAL | 2025-09-08 | 16:10 | 2025-09-08 | 15:32 | ROLAND | 2 | 58.1 | 58.7 | / | / |  |  |  |  | 1.93 | 190 |  | Kinga |  | 3 | 4
4 | 60.0 | 218 | P/FINAL | 2025-09-08 | 16:10 | 2025-09-08 | 15:32 | ROLAND | 2 | 59.4 | 58.7 | / | / |  |  |  |  | 1.69 | 202 |  | Kinga |  | 3 | 4
4 | 60.0 | 219 | P/FINAL | 2025-09-08 | 16:11 | 2025-09-08 | 15:32 | ROLAND | 2 | 60.6 | 60.9 | / | / |  |  |  | 1.998 | 1.88 | 183 |  | Kinga |  | 3 | 4
2 | 15.0 | 220 | P/FINAL | 2025-09-09 | 08:28 | 2025-09-09 | 07:31 | Kinga | 2 | 15.0 | 15.0 | / | / |  |  |  |  | 4.75 | 193 |  | Kinga |  | 3 | 2 | 112146
2 | 15.0 | 221 | P/FINAL | 2025-09-09 | 08:28 | 2025-09-09 | 07:31 | Kinga | 2 | 16.7 | 16.5 | / | / |  |  |  | 1.922 | 2.08 | 208 |  | Kinga |  | 3 | 2 | 112146
2 | 15.0 | 222 | P/FINAL | 2025-09-09 | 08:29 | 2025-09-09 | 07:31 | Kinga | 2 | 14.8 | 15.5 | / | / |  |  |  |  | 2.08 | 181 |  | Kinga |  | 3 | 2 | 112146
3 | 35.0 | 223 | P/FINAL | 2025-09-09 | 08:30 | 2025-09-09 | 07:31 | Kinga | 2 | 33.0 | 33.2 | / | / |  |  |  | 1.962 | 1.45 | 199 |  | Kinga |  | 3 | 3 | 112146
3 | 35.0 | 224 | P/FINAL | 2025-09-09 | 08:31 | 2025-09-09 | 07:31 | Kinga | 2 | 34.3 | 33.5 | / | / |  |  |  |  | 2.02 | 193 |  | Kinga |  | 3 | 3 | 112146
3 | 35.0 | 225 | P/FINAL | 2025-09-09 | 08:32 | 2025-09-09 | 07:31 | Kinga | 2 | 35.5 | 35.5 | / | / |  |  |  |  | 2.09 | 177 |  | Kinga |  | 3 | 3 | 112146
4 | 60.0 | 226 | P/FINAL | 2025-09-09 | 08:32 | 2025-09-09 | 07:31 | Kinga | 2 | 57.0 | 57.3 | / | / |  |  |  | 1.931 | 2.33 | 194 |  | Kinga |  | 3 | 4 | 112146
4 | 60.0 | 227 | P/FINAL | 2025-09-09 | 08:33 | 2025-09-09 | 07:31 | Kinga | 2 | 61.2 | 60.3 | / | / |  |  |  |  | 2.99 | 194 |  | Kinga |  | 3 | 4 | 112146
1 | 0.0 | 228 | P/FINAL | 2025-09-09 | 08:34 | 2025-09-09 | 07:31 | Kinga | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 2.99 | 183 |  | Kinga |  | 3 | 1 | 112146
1 | 0.0 | 229 | P/FINAL | 2025-09-09 | 08:35 | 2025-09-09 | 07:31 | Kinga | 2 | 0.0 | 0.0 | / | / |  |  |  | 1.908 | 3.09 | 199 |  | Kinga |  | 3 | 1 | 112146
2 | 15.0 | 230 | P/FINAL | 2025-09-09 | 11:22 | 2025-09-09 | 10:45 | ELENA | 2 | 14.4 | 14.5 | / | / |  |  |  |  | 1.20 | 200 |  | Kinga |  | 3 | 2 | 112146
2 | 15.0 | 231 | P/FINAL | 2025-09-09 | 11:23 | 2025-09-09 | 10:45 | ELENA | 2 | 15.2 | 14.8 | / | / |  |  |  | 1.890 | 1.60 | 185 |  | Kinga |  | 3 | 2 | 112146
2 | 15.0 | 232 | P/FINAL | 2025-09-09 | 11:23 | 2025-09-09 | 10:45 | ELENA | 2 | 15.1 | 15.1 | / | / |  |  |  |  | 2.70 | 211 |  | Kinga |  | 3 | 2 | 112146
3 | 35.0 | 233 | P/FINAL | 2025-09-09 | 11:24 | 2025-09-09 | 10:45 | ELENA | 2 | 34.7 | 35.0 | / | / |  |  |  |  | 4.55 | 190 |  | Kinga |  | 3 | 3 | 112146
3 | 35.0 | 234 | P/FINAL | 2025-09-09 | 11:25 | 2025-09-09 | 10:45 | ELENA | 2 | 34.7 | 34.5 | / | / |  |  |  | 1.900 | 2.00 | 189 |  | Kinga |  | 3 | 3 | 112146
3 | 35.0 | 235 | P/FINAL | 2025-09-09 | 11:25 | 2025-09-09 | 10:45 | ELENA | 2 | 35.4 | 35.4 | / | / |  |  |  |  | 2.01 | 216 |  | Kinga |  | 3 | 3 | 112146
4 | 60.0 | 236 | P/FINAL | 2025-09-09 | 11:26 | 2025-09-09 | 10:45 | ELENA | 2 | 61.9 | 62.0 | / | / |  |  |  |  | 1.05 | 187 |  | Kinga |  | 3 | 4 | 112146
4 | 60.0 | 237 | P/FINAL | 2025-09-09 | 11:26 | 2025-09-09 | 10:45 | ELENA | 2 | 57.2 | 60.6 | / | / |  |  |  | 1.922 | 1.05 | 189 |  | Kinga |  | 3 | 4 | 112146
4 | 60.0 | 238 | P/FINAL | 2025-09-09 | 11:27 | 2025-09-09 | 10:45 | ELENA | 2 | 61.3 | 61.4 | / | / |  |  |  |  | 1.09 | 187 |  | Kinga |  | 3 | 4 | 112146
Standard # 1
Total Tests | Total Readings | Total Average | Selected Tests | Selected Readings | Selected Average
25 | 50 | 0 mg/L | 25 | 50 | 0 mg/L
Standard # 2
Total Tests | Total Readings | Total Average | Selected Tests | Selected Readings | Selected Average
57 | 114 | 14.9 mg/L | 57 | 114 | 14.9 mg/L
Standard # 3
Total Tests | Total Readings | Total Average | Selected Tests | Selected Readings | Selected Average
58 | 116 | 35.6 mg/L | 58 | 116 | 35.6 mg/L
Standard # 4
Total Tests | Total Readings | Total Average | Selected Tests | Selected Readings | Selected Average
55 | 110 | 60.1 mg/L | 55 | 110 | 60.1 mg/L
Specifications
Standard Value [ppm] | Target Value ÷ U [ppm] | Mean Value [ppm] | Passed
0.0 | 0.0 ÷ 2.0 | 0.0 | YES
15.0 | 12.3 ÷ 17.7 | 14.9 | YES
35.0 | 31.3 ÷ 38.7 | 35.6 | YES
60.0 | 55.1 ÷ 64.9 | 60.1 | YES
Validation Date | by