## Hanna Code
Lot | Hanna SFG Code | Recipe | Preparation Week |  | Hanna SFG Code | HI782-0
1269 | HI782-0 | CP-R80 | W 49/1 |  | Parameter Formula | NO3
Description | Exp | Line |  |  | QC Department | Registration Book | QC Type
Marine Nitrate HR Reagent | 2030-12-01 | L57 Powder |  |  | Laboratory | TABLET 3 | Production
Ref Weight. | Min (mg) | Max (mg)
206 | 196 | 216
REAGENT SET 1 | [5 empty] | REAGENT SET 2
Reagent A Lot | Reagent B Lot | Reagent C Lot | Reagent D Lot | Reagent E Lot |  | Reagent A Lot | Reagent B Lot | Reagent C Lot | Reagent D Lot | Reagent E Lot
Expiration A | Expiration B | Expiration C | Expiration D | Expiration E |  | Expiration A | Expiration B | Expiration C | Expiration D | Expiration E
Reagent Range Min | Reagent Range Max
0.0 mg/L | 75.0 mg/L
Prep. Operator | First day Prod. | Last day Prod. | Machine |  |  | Old Lot A | Old Lot B | Old Lot D
ERIKA | 2025-12-08 | 2025-12-11 | 2 HEAD |  |  | 1191
[6 empty] | Expiration A | Expiration B | Expiration D
2030-09-01
Meter
Meter 1 Family | Meter 1 Code | Meter 2 Family | Meter 2 Code
HI97115 | CTO-730 | HI97115 | CTO-740
ph Meter | Description | Turbid. Meter | Description | Spectr. Meter | Description
|  | HI98713 | CTO-797
Standard | mg/L | [5 empty] | Tolerance |  |  |  |  | pH
STD Number | STD Value | STD Min | STD Max |  | STD MR |  | Fixed | AndOr | Percentage | QCRestriction |  | pH Value | pH Min | pH Max
2 | 15.0 | 12.3 | 17.7 |  | MR101 |  | 2.0 | & | 4.9% | 100%
3 | 35.0 | 31.3 | 38.7
4 | 60.0 | 55.1 | 64.9 | [8 empty] | 1.7 | 1.5 | 1.9
1 | 0.0 | 0.0 | 2.0
Readings Table
Standard | STD Value | # | TEST | QC DATE | QC TIME | PROD. DATE | PROD. TIME | PROD. OPERATOR | HEAD | M. 1 HI97115 | M. 2 HI97115 | M. 3 Label17 | M. 4 Label17 | SPECTR. [ABS] | pH 1 | pH 2 | pH 3 | TURB. | WEIGHT [mg] | REAGENT SET | QC OPERATOR | CORRECTION | phNumber | STD | STD_ID | NOTE
2 | 15.0 | 1 | Valid | 2025-12-08 | 11:56 | 2025-12-08 | 11:32 | ERIKA |  | 19.2 | 19.2 | / | / |  |  |  |  | 1.28 |  |  | Kinga |  | 3 | 2 | 24345
2 | 15.0 | 2 | Valid | 2025-12-08 | 11:57 | 2025-12-08 | 11:32 | ERIKA |  | 17.4 | 17.4 | / | / |  |  |  | 1.850 | 2.58 |  |  | Kinga |  | 3 | 2 | 24345
2 | 15.0 | 3 | Valid | 2025-12-08 | 11:58 | 2025-12-08 | 11:32 | ERIKA |  | 19.4 | 19.0 | / | / |  |  |  |  | 3.03 |  |  | Kinga |  | 3 | 2 | 24345
3 | 35.0 | 4 | Valid | 2025-12-08 | 11:59 | 2025-12-08 | 11:32 | ERIKA |  | 45.6 | 45.7 | / | / |  |  |  |  | 2.44 |  |  | Kinga |  | 3 | 3 | 24345
3 | 35.0 | 5 | Valid | 2025-12-08 | 12:01 | 2025-12-08 | 11:32 | ERIKA |  | 43.3 | 43.5 | / | / |  |  |  | 1.853 | 2.65 |  |  | Kinga |  | 3 | 3 | 24345
4 | 60.0 | 6 | Valid | 2025-12-08 | 12:10 | 2025-12-08 | 11:32 | ERIKA |  | 65.5 | 65.2 | / | / |  |  |  |  | 2.30 |  |  | Kinga |  | 3 | 4 | 24345
4 | 60.0 | 7 | Valid | 2025-12-08 | 12:11 | 2025-12-08 | 11:32 | ERIKA |  | 62.7 | 62.9 | / | / |  |  |  | 1.919 | 2.89 |  |  | Kinga |  | 3 | 4 | 24345
2 | 15.0 | 8 | Old A | 2025-12-08 | 12:12 | 2025-12-08 | 11:32 | ERIKA |  | 15.0 | 15.7 | / | / |  |  |  | 1.926 | 1.44 |  |  | Kinga |  | 3 | 2 | 24345
3 | 35.0 | 9 | Old A | 2025-12-08 | 12:13 | 2025-12-08 | 11:32 | ERIKA |  | 37.0 | 37.2 | / | / |  |  |  |  | 1.77 |  |  | Kinga |  | 3 | 3 | 24345
4 | 60.0 | 10 | Old A | 2025-12-08 | 12:17 | 2025-12-08 | 11:32 | ERIKA |  | 55.3 | 55.2 | / | / |  |  |  |  | 3.97 |  |  | Kinga |  | 3 | 4 | 24345
2 | 15.0 | 11 | Old A | 2025-12-08 | 12:18 | 2025-12-08 | 11:32 | ERIKA |  | 14.4 | 14.3 | / | / |  |  |  |  | 2.00 |  |  | Kinga |  | 3 | 2 | 24345
3 | 35.0 | 12 | Old A | 2025-12-08 | 12:22 | 2025-12-08 | 11:32 | ERIKA |  | 34.4 | 34.5 | / | / |  |  |  |  | 1.60 |  |  | Kinga |  | 3 | 3 | 24345
4 | 60.0 | 13 | Old A | 2025-12-08 | 12:27 | 2025-12-08 | 11:32 | ERIKA |  | 55.8 | 55.9 | / | / |  |  |  |  | 1.42 |  |  | Kinga |  | 3 | 4 | 24345
1 | 0.0 | 14 | Valid | 2025-12-08 | 12:32 | 2025-12-08 | 11:32 | ERIKA |  | 0.0 | 0.0 | / | / |  |  |  | 1.877 | 5.91 |  |  | Kinga |  | 3 | 1 | 24345
1 | 0.0 | 15 | Old A | 2025-12-08 | 12:35 | 2025-12-08 | 11:32 | ERIKA |  | 0.0 | 0.0 | / | / |  |  |  |  | 1.37 |  |  | Kinga |  | 3 | 1 | 24345
1 | 0.0 | 16 | Old A | 2025-12-08 | 12:35 | 2025-12-08 | 11:32 | ERIKA |  | 0.0 | 0.0 | / | / |  |  |  | 1.886 | 0.78 |  |  | Kinga |  | 3 | 1 | 24345
3 | 35.0 | 17 | After correction | 2025-12-08 | 13:45 | 2025-12-08 | 13:00 | ERIKA |  | 38.2 | 38.2 | / | / |  |  |  |  | 2.31 | 207 |  | Kinga |  | 3 | 3 | 24345
3 | 35.0 | 18 | After correction | 2025-12-08 | 13:46 | 2025-12-08 | 13:00 | ERIKA |  | 37.5 | 37.6 | / | / |  |  |  |  | 3.04 | 205 |  | Kinga |  | 3 | 3 | 24345
3 | 35.0 | 19 | After correction | 2025-12-08 | 14:01 | 2025-12-08 | 13:00 | ERIKA |  | 38.1 | 38.1 | / | / |  |  |  |  | 13.8 | 250 |  | Kinga |  | 3 | 3 | 24345
3 | 35.0 | 20 | After correction | 2025-12-08 | 14:03 | 2025-12-08 | 13:00 | ERIKA |  | 38.0 | 38.1 | / | / |  |  |  | 1.853 | 2.53 | 241 |  | Kinga |  | 3 | 3 | 24345
2 | 15.0 | 21 | After correction | 2025-12-08 | 14:19 | 2025-12-08 | 13:00 | ERIKA |  | 15.6 | 15.6 | / | / |  |  |  |  | 11.9 | 245 |  | Kinga |  | 3 | 2 | 24345
4 | 60.0 | 22 | After correction | 2025-12-08 | 14:20 | 2025-12-08 | 13:00 | ERIKA |  | 57.9 | 57.7 | / | / |  |  |  |  | 4.39 | 245 |  | Kinga |  | 3 | 4 | 24345
2 | 15.0 | 23 | After correction | 2025-12-08 | 14:34 | 2025-12-08 | 13:00 | ERIKA |  | 15.4 | 15.4 | / | / |  |  |  |  | 2.50 | 244 |  | Kinga |  | 3 | 2 | 24345
4 | 60.0 | 24 | After correction | 2025-12-08 | 14:36 | 2025-12-08 | 13:00 | ERIKA |  | 58.4 | 58.2 | / | / |  |  |  |  | 4.75 | 248 |  | Kinga |  | 3 | 4 | 24345
2 | 15.0 | 25 | P/FINAL | 2025-12-09 | 08:55 | 2025-12-09 | 08:18 | ROLAND | 2 | 16.5 | 16.8 | / | / |  |  |  |  | 1.03 | 246 |  | Kinga |  | 3 | 2 | 24345
2 | 15.0 | 26 | P/FINAL | 2025-12-09 | 08:57 | 2025-12-09 | 08:18 | ROLAND | 2 | 17.1 | 17.2 | / | / |  |  |  | 1.803 | 1.18 | 248 |  | Kinga |  | 3 | 2 | 24345
2 | 15.0 | 27 | P/FINAL | 2025-12-09 | 08:57 | 2025-12-09 | 08:18 | ROLAND | 2 | 16.7 | 16.9 | / | / |  |  |  |  | 1.24 | 251 |  | Kinga |  | 3 | 2 | 24345
3 | 35.0 | 28 | P/FINAL | 2025-12-09 | 08:59 | 2025-12-09 | 08:18 | ROLAND | 2 | 37.6 | 37.7 | / | / |  |  |  | 1.784 | 1.57 | 245 |  | Kinga |  | 3 | 3 | 24345
3 | 35.0 | 29 | P/FINAL | 2025-12-09 | 08:59 | 2025-12-09 | 08:18 | ROLAND | 2 | 33.6 | 33.6 | / | / |  |  |  |  | 1.38 | 247 |  | Kinga |  | 3 | 3 | 24345
3 | 35.0 | 30 | P/FINAL | 2025-12-09 | 09:01 | 2025-12-09 | 08:18 | ROLAND | 2 | 36.1 | 36.2 | / | / |  |  |  |  | 1.27 | 242 |  | Kinga |  | 3 | 3 | 24345
1 | 0.0 | 31 | P/FINAL | 2025-12-09 | 09:04 | 2025-12-09 | 08:18 | ROLAND | 2 | 0.0 | 0.0 | / | / |  |  |  | 1.738 | 2.67 | 251 |  | Kinga |  | 3 | 1 | 24345
1 | 0.0 | 32 | P/FINAL | 2025-12-09 | 09:04 | 2025-12-09 | 08:18 | ROLAND | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 2.84 | 248 |  | Kinga |  | 3 | 1 | 24345
4 | 60.0 | 33 | P/FINAL | 2025-12-09 | 09:05 | 2025-12-09 | 08:18 | ROLAND | 2 | 55.8 | 55.9 | / | / |  |  |  |  | 1.38 | 251 |  | Kinga |  | 3 | 4 | 24345
4 | 60.0 | 34 | P/FINAL | 2025-12-09 | 09:06 | 2025-12-09 | 08:18 | ROLAND | 2 | 59.7 | 59.9 | / | / |  |  |  | 1.797 | 1.58 | 254 |  | Kinga |  | 3 | 4 | 24345
2 | 15.0 | 35 | Old A | 2025-12-09 | 10:24 | 2025-12-09 |  | ROLAND | 2 | 13.7 | 13.7 | / | / |  |  |  |  | 2.18 |  |  | Kinga |  | 3 | 2 | 24345
2 | 15.0 | 36 | Old A | 2025-12-09 | 10:25 | 2025-12-09 |  | ROLAND | 2 | 14.0 | 14.0 | / | / |  |  |  |  | 1.99 |  |  | Kinga |  | 3 | 2 | 24345
3 | 35.0 | 37 | Old A | 2025-12-09 | 10:28 | 2025-12-09 |  | ROLAND | 2 | 32.2 | 32.4 | / | / |  |  |  |  | 3.35 |  |  | Kinga |  | 3 | 3 | 24345
3 | 35.0 | 38 | Old A | 2025-12-09 | 10:28 | 2025-12-09 |  | ROLAND | 2 | 33.0 | 33.0 | / | / |  |  |  |  | 3.26 |  |  | Kinga |  | 3 | 3 | 24345
4 | 60.0 | 39 | Old A | 2025-12-09 | 10:58 | 2025-12-09 |  | ROLAND | 2 | 55.3 | 55.3 | / | / |  |  |  |  | 1.17 |  |  | Kinga |  | 3 | 4 | 24345
4 | 60.0 | 40 | Old A | 2025-12-09 | 10:58 | 2025-12-09 |  | ROLAND | 2 | 55.9 | 55.9 | / | / |  |  |  |  | 1.21 |  |  | Kinga |  | 3 | 4 | 24345
2 | 15.0 | 41 | P/FINAL | 2025-12-09 | 11:08 | 2025-12-09 | 10:43 | ROLAND | 2 | 16.0 | 16.1 | / | / |  |  |  | 1.895 | 1.13 | 242 |  | Kinga |  | 3 | 2 | 24345
2 | 15.0 | 42 | P/FINAL | 2025-12-09 | 11:09 | 2025-12-09 | 10:43 | ROLAND | 2 | 16.3 | 16.3 | / | / |  |  |  |  | 1.22 | 245 |  | Kinga |  | 3 | 2 | 24345
2 | 15.0 | 43 | P/FINAL | 2025-12-09 | 11:10 | 2025-12-09 | 10:43 | ROLAND | 2 | 16.9 | 17.0 | / | / |  |  |  |  | 1.22 | 244 |  | Kinga |  | 3 | 2 | 24345
3 | 35.0 | 44 | P/FINAL | 2025-12-09 | 11:13 | 2025-12-09 | 10:43 | ROLAND | 2 | 37.9 | 38.0 | / | / |  |  |  | 1.853 | 1.48 | 225 |  | Kinga |  | 3 | 3 | 24345
3 | 35.0 | 45 | P/FINAL | 2025-12-09 | 11:14 | 2025-12-09 | 10:43 | ROLAND | 2 | 38.1 | 38.2 | / | / |  |  |  |  | 1.98 | 229 |  | Kinga |  | 3 | 3 | 24345
3 | 35.0 | 46 | P/FINAL | 2025-12-09 | 11:15 | 2025-12-09 | 10:43 | ROLAND | 2 | 38.7 | 38.5 | / | / |  |  |  |  | 1.99 | 239 |  | Kinga |  | 3 | 3 | 24345
4 | 60.0 | 47 | P/FINAL | 2025-12-09 | 11:21 | 2025-12-09 | 10:43 | ROLAND | 2 | 60.1 | 60.2 | / | / |  |  |  | 1.863 | 1.60 | 247 |  | Kinga |  | 3 | 4 | 24345
4 | 60.0 | 48 | P/FINAL | 2025-12-09 | 11:25 | 2025-12-09 | 10:43 | ROLAND | 2 | 59.2 | 59.4 | / | / |  |  |  |  | 1.75 | 244 |  | Kinga |  | 3 | 4 | 24345
4 | 60.0 | 49 | P/FINAL | 2025-12-09 | 11:26 | 2025-12-09 | 10:43 | ROLAND | 2 | 60.6 | 60.7 | / | / |  |  |  |  | 1.74 | 242 |  | Kinga |  | 3 | 4 | 24345
1 | 0.0 | 50 | P/FINAL | 2025-12-09 | 11:28 | 2025-12-09 | 10:43 | ROLAND | 2 | 0.0 | 0.1 | / | / |  |  |  |  | 7.11 | 235 |  | Kinga |  | 3 | 1 | 24345
1 | 0.0 | 51 | P/FINAL | 2025-12-09 | 11:29 | 2025-12-09 | 10:43 | ROLAND | 2 | 0.0 | 0.0 | / | / |  |  |  | 1.806 | 1.96 | 238 |  | Kinga |  | 3 | 1 | 24345
2 | 15.0 | 52 | P/FINAL | 2025-12-09 | 14:35 | 2025-12-09 | 14:23 | ELENA | 2 | 16.1 | 16.1 | / | / |  |  |  |  | 4.21 | 243 |  | Kinga |  | 3 | 2 | 24459
2 | 15.0 | 53 | P/FINAL | 2025-12-09 | 14:44 | 2025-12-09 | 14:23 | ELENA | 2 | 16.3 | 16.3 | / | / |  |  |  |  | 1.33 | 240 |  | Kinga |  | 3 | 2 | 24459
2 | 15.0 | 54 | P/FINAL | 2025-12-09 | 14:48 | 2025-12-09 | 14:23 | ELENA | 2 | 12.7 | 12.9 | / | / |  |  |  | 1.825 | 3.36 | 244 |  | Kinga |  | 3 | 2 | 24459
1 | 0.0 | 55 | P/FINAL | 2025-12-09 | 14:48 | 2025-12-09 | 14:23 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 4.14 | 249 |  | Kinga |  | 3 | 1 | 24459
1 | 0.0 | 56 | P/FINAL | 2025-12-09 | 14:48 | 2025-12-09 | 14:23 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  | 1.821 | 2.91 | 242 |  | Kinga |  | 3 | 1 | 24459
4 | 60.0 | 57 | P/FINAL | 2025-12-09 | 14:51 | 2025-12-09 | 14:23 | ELENA | 2 | 57.3 | 57.2 | / | / |  |  |  |  | 1.72 | 247 |  | Kinga |  | 3 | 4 | 24459
4 | 60.0 | 58 | P/FINAL | 2025-12-09 | 14:53 | 2025-12-09 | 14:23 | ELENA | 2 | 59.0 | 59.1 | / | / |  |  |  | 1.850 | 1.92 | 245 |  | Kinga |  | 3 | 4 | 24459
4 | 60.0 | 59 | P/FINAL | 2025-12-09 | 14:57 | 2025-12-09 | 14:23 | ELENA | 2 | 61.0 | 60.0 | / | / |  |  |  |  | 2.18 | 244 |  | Kinga |  | 3 | 4 | 24459
3 | 35.0 | 60 | P/FINAL | 2025-12-09 | 14:58 | 2025-12-09 | 14:23 | ELENA | 2 | 37.2 | 37.3 | / | / |  |  |  |  | 1.40 | 234 |  | Kinga |  | 3 | 3 | 24459
3 | 35.0 | 61 | P/FINAL | 2025-12-09 | 14:59 | 2025-12-09 | 14:23 | ELENA | 2 | 36.5 | 36.7 | / | / |  |  |  | 1.829 | 1.47 | 242 |  | Kinga |  | 3 | 3 | 24459
3 | 35.0 | 62 | P/FINAL | 2025-12-09 | 15:04 | 2025-12-09 | 14:23 | ELENA | 2 | 38.0 | 38.0 | / | / |  |  |  | 1.850 | 1.60 | 240 |  | Kinga |  | 3 | 3 | 24459
2 | 15.0 | 63 | P/FINAL | 2025-12-09 | 15:50 | 2025-12-09 | 15:23 | ELENA | 2 | 16.1 | 16.1 | / | / |  |  |  |  | 1.48 | 242 |  | Kinga |  | 3 | 2 | 24459
2 | 15.0 | 64 | P/FINAL | 2025-12-09 | 15:50 | 2025-12-09 | 15:23 | ELENA | 2 | 16.5 | 16.6 | / | / |  |  |  |  | 2.71 | 245 |  | Kinga |  | 3 | 2 | 24459
2 | 15.0 | 65 | P/FINAL | 2025-12-09 | 15:51 | 2025-12-09 | 15:23 | ELENA | 2 | 13.6 | 13.7 | / | / |  |  |  | 1.787 | 1.91 | 237 |  | Kinga |  | 3 | 2 | 24459
3 | 35.0 | 66 | P/FINAL | 2025-12-09 | 15:53 | 2025-12-09 | 15:23 | ELENA | 2 | 31.3 | 31.6 | / | / |  |  |  |  | 1.25 | 245 |  | Kinga |  | 3 | 3 | 24459
3 | 35.0 | 67 | P/FINAL | 2025-12-09 | 15:54 | 2025-12-09 | 15:23 | ELENA | 2 | 32.3 | 32.4 | / | / |  |  |  |  | 1.55 | 248 |  | Kinga |  | 3 | 3 | 24459
3 | 35.0 | 68 | P/FINAL | 2025-12-09 | 15:56 | 2025-12-09 | 15:23 | ELENA | 2 | 32.9 | 32.9 | / | / |  |  |  | 1.845 | 1.33 | 242 |  | Kinga |  | 3 | 3 | 24459
4 | 60.0 | 69 | P/FINAL | 2025-12-09 | 15:57 | 2025-12-09 | 15:23 | ELENA | 2 | 62.3 | 62.3 | / | / |  |  |  | 1.867 | 3.27 | 41 |  | Kinga |  | 3 | 4 | 24459
4 | 60.0 | 70 | P/FINAL | 2025-12-09 | 15:58 | 2025-12-09 | 15:23 | ELENA | 2 | 61.2 | 61.3 | / | / |  |  |  |  | 0.94 | 245 |  | Kinga |  | 3 | 4 | 24459
4 | 60.0 | 71 | P/FINAL | 2025-12-09 | 16:00 | 2025-12-09 | 15:23 | ELENA | 2 | 59.8 | 59.9 | / | / |  |  |  |  | 0.99 | 240 |  | Kinga |  | 3 | 4 | 24459
2 | 15.0 | 72 | P/FINAL | 2025-12-09 | 17:53 | 2025-12-09 | 17:24 | ELENA | 2 | 15.5 | 15.7 | / | / |  |  |  |  | 1.16 | 243 |  | Andrea V |  | 3 | 2 | 24459
2 | 15.0 | 73 | P/FINAL | 2025-12-09 | 17:53 | 2025-12-09 | 17:24 | ELENA | 2 | 14.2 | 14.1 | / | / |  |  |  |  | 0.96 | 230 |  | Andrea V |  | 3 | 2 | 24459
2 | 15.0 | 74 | P/FINAL | 2025-12-09 | 17:54 | 2025-12-09 | 17:24 | ELENA | 2 | 17.6 | 17.5 | / | / |  |  |  | 1.806 | 1.49 | 236 |  | Andrea V |  | 3 | 2 | 24459
3 | 35.0 | 75 | P/FINAL | 2025-12-09 | 17:54 | 2025-12-09 | 17:24 | ELENA | 2 | 32.2 | 32.4 | / | / |  |  |  |  | 1.97 | 231 |  | Andrea V |  | 3 | 3 | 24459
3 | 35.0 | 76 | P/FINAL | 2025-12-09 | 17:55 | 2025-12-09 | 17:24 | ELENA | 2 | 34.1 | 34.3 | / | / |  |  |  | 1.809 | 0.84 | 240 |  | Andrea V |  | 3 | 3 | 24459
4 | 60.0 | 77 | P/FINAL | 2025-12-09 | 17:56 | 2025-12-09 | 17:24 | ELENA | 2 | 60.9 | 61.1 | / | / |  |  |  |  | 1.19 | 248 |  | Andrea V |  | 3 | 4 | 24459
4 | 60.0 | 78 | P/FINAL | 2025-12-09 | 17:57 | 2025-12-09 | 17:24 | ELENA | 2 | 60.9 | 61.1 | / | / |  |  |  | 1.839 | 1.21 | 237 |  | Andrea V |  | 3 | 4 | 24459
4 | 60.0 | 79 | P/FINAL | 2025-12-09 | 17:58 | 2025-12-09 | 17:24 | ELENA | 2 | 60.8 | 61.2 | / | / |  |  |  |  | 1.00 | 223 |  | Andrea V |  | 3 | 4 | 24459
1 | 0.0 | 80 | P/FINAL | 2025-12-09 | 18:00 | 2025-12-09 | 17:24 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 5.02 | 245 |  | Andrea V |  | 3 | 1 | 24459
1 | 0.0 | 81 | P/FINAL | 2025-12-09 | 18:01 | 2025-12-09 | 17:24 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  | 1.806 | 5.89 | 235 |  | Andrea V |  | 3 | 1 | 24459
3 | 35.0 | 82 | P/FINAL | 2025-12-09 | 18:06 | 2025-12-09 | 17:24 | ELENA | 2 | 35.5 | 35.3 | / | / |  |  |  |  | 1.07 | 235 |  | Andrea V |  | 3 | 3 | 24459
2 | 15.0 | 83 | P/FINAL | 2025-12-09 | 19:18 | 2025-12-09 | 18:51 | ELENA | 2 | 16.0 | 16.1 | / | / |  |  |  | 1.805 | 2.18 | 251 |  | Andrea V |  | 3 | 2
2 | 15.0 | 84 | P/FINAL | 2025-12-09 | 19:19 | 2025-12-09 | 18:51 | ELENA | 2 | 16.4 | 16.4 | / | / |  |  |  |  | 0.76 | 245 |  | Andrea V |  | 3 | 2
2 | 15.0 | 85 | P/FINAL | 2025-12-09 | 19:20 | 2025-12-09 | 18:51 | ELENA | 2 | 17.2 | 17.1 | / | / |  |  |  |  | 0.83 | 234 |  | Andrea V |  | 3 | 2
3 | 35.0 | 86 | P/FINAL | 2025-12-09 | 19:21 | 2025-12-09 | 18:51 | ELENA | 2 | 34.1 | 34.2 | / | / |  |  |  | 1.816 | 0.85 | 245 |  | Andrea V |  | 3 | 3
3 | 35.0 | 87 | P/FINAL | 2025-12-09 | 19:21 | 2025-12-09 | 18:51 | ELENA | 2 | 37.1 | 37.2 | / | / |  |  |  |  | 0.91 | 226 |  | Andrea V |  | 3 | 3
3 | 35.0 | 88 | P/FINAL | 2025-12-09 | 19:23 | 2025-12-09 | 18:51 | ELENA | 2 | 36.1 | 36.1 | / | / |  |  |  |  | 0.90 | 239 |  | Andrea V |  | 3 | 3
4 | 60.0 | 89 | P/FINAL | 2025-12-09 | 19:23 | 2025-12-09 | 18:51 | ELENA | 2 | 58.9 | 58.9 | / | / |  |  |  |  | 0.90 | 247 |  | Andrea V |  | 3 | 4 | 24459
4 | 60.0 | 90 | P/FINAL | 2025-12-09 | 19:24 | 2025-12-09 | 18:51 | ELENA | 2 | 63.2 | 63.3 | / | / |  |  |  |  | 1.13 | 225 |  | Andrea V |  | 3 | 4 | 24459
4 | 60.0 | 91 | P/FINAL | 2025-12-09 | 19:25 | 2025-12-09 | 18:51 | ELENA | 2 | 62.1 | 62.3 | / | / |  |  |  | 1.847 | 1.19 | 240 |  | Andrea V |  | 3 | 4 | 24459
1 | 0.0 | 92 | P/FINAL | 2025-12-09 | 19:25 | 2025-12-09 | 18:51 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 2.80 | 241 |  | Andrea V |  | 3 | 1 | 24459
1 | 0.0 | 93 | P/FINAL | 2025-12-09 | 19:26 | 2025-12-09 | 18:51 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  | 1.807 | 2.59 | 248 |  | Andrea V |  | 3 | 1 | 24459
2 | 15.0 | 94 | P/FINAL | 2025-12-09 | 20:14 | 2025-12-09 | 19:50 | ELENA | 2 | 14.5 | 14.6 | / | / |  |  |  |  | 1.13 | 244 |  | Andrea V |  | 3 | 2 | 24459
2 | 15.0 | 95 | P/FINAL | 2025-12-09 | 20:15 | 2025-12-09 | 19:50 | ELENA | 2 | 15.4 | 15.5 | / | / |  |  |  | 1.793 | 1.60 | 240 |  | Andrea V |  | 3 | 2 | 24459
2 | 15.0 | 96 | P/FINAL | 2025-12-09 | 20:15 | 2025-12-09 | 19:50 | ELENA | 2 | 16.6 | 16.7 | / | / |  |  |  |  | 0.78 | 227 |  | Andrea V |  | 3 | 2 | 24459
3 | 35.0 | 97 | P/FINAL | 2025-12-09 | 20:16 | 2025-12-09 | 19:50 | ELENA | 2 | 37.7 | 37.8 | / | / |  |  |  |  | 2.00 | 240 |  | Andrea V |  | 3 | 3 | 24459
3 | 35.0 | 98 | P/FINAL | 2025-12-09 | 20:17 | 2025-12-09 | 19:50 | ELENA | 2 | 36.4 | 36.5 | / | / |  |  |  | 1.802 | 1.10 | 240 |  | Andrea V |  | 3 | 3 | 24459
3 | 35.0 | 99 | P/FINAL | 2025-12-09 | 20:18 | 2025-12-09 | 19:50 | ELENA | 2 | 38.0 | 38.0 | / | / |  |  |  |  | 1.49 | 240 |  | Andrea V |  | 3 | 3 | 24459
4 | 60.0 | 100 | P/FINAL | 2025-12-09 | 20:18 | 2025-12-09 | 19:50 | ELENA | 2 | 61.0 | 61.2 | / | / |  |  |  |  | 1.12 | 234 |  | Andrea V |  | 3 | 4 | 24459
4 | 60.0 | 101 | P/FINAL | 2025-12-09 | 20:19 | 2025-12-09 | 19:50 | ELENA | 2 | 62.9 | 62.8 | / | / |  |  |  | 1.835 | 1.02 | 229 |  | Andrea V |  | 3 | 4 | 24459
4 | 60.0 | 102 | P/FINAL | 2025-12-09 | 20:20 | 2025-12-09 | 19:50 | ELENA | 2 | 62.0 | 62.2 | / | / |  |  |  |  | 1.05 | 236 |  | Andrea V |  | 3 | 4 | 24459
1 | 0.0 | 103 | P/FINAL | 2025-12-09 | 20:22 | 2025-12-09 | 19:50 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 1.67 | 232 |  | Andrea V |  | 3 | 1 | 24459
1 | 0.0 | 104 | P/FINAL | 2025-12-09 | 20:22 | 2025-12-09 | 19:50 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  | 1.814 | 1.37 | 250 |  | Andrea V |  | 3 | 1 | 24459
2 | 15.0 | 105 | P/FINAL | 2025-12-09 | 21:16 | 2025-12-09 | 20:46 | ELENA | 2 | 16.7 | 16.8 | / | / |  |  |  |  | 1.36 | 231 |  | Andrea V |  | 3 | 2 | 24459
2 | 15.0 | 106 | P/FINAL | 2025-12-09 | 21:16 | 2025-12-09 | 20:46 | ELENA | 2 | 15.1 | 15.3 | / | / |  |  |  | 1.761 | 1.03 | 238 |  | Andrea V |  | 3 | 2 | 24459
2 | 15.0 | 107 | P/FINAL | 2025-12-09 | 21:18 | 2025-12-09 | 20:46 | ELENA | 2 | 14.0 | 14.1 | / | / |  |  |  |  | 0.98 | 256 |  | Andrea V |  | 3 | 2 | 24459
3 | 35.0 | 108 | P/FINAL | 2025-12-09 | 21:18 | 2025-12-09 | 20:46 | ELENA | 2 | 37.8 | 37.9 | / | / |  |  |  |  | 1.66 | 226 |  | Andrea V |  | 3 | 3 | 24459
3 | 35.0 | 109 | P/FINAL | 2025-12-09 | 21:19 | 2025-12-09 | 20:46 | ELENA | 2 | 35.1 | 35.3 | / | / |  |  |  |  | 0.81 | 247 |  | Andrea V |  | 3 | 3 | 24459
3 | 35.0 | 110 | P/FINAL | 2025-12-09 | 21:20 | 2025-12-09 | 20:46 | ELENA | 2 | 33.7 | 33.9 | / | / |  |  |  | 1.791 | 0.71 | 239 |  | Andrea V |  | 3 | 3 | 24459
4 | 60.0 | 111 | P/FINAL | 2025-12-09 | 21:21 | 2025-12-09 | 20:46 | ELENA | 2 | 57.8 | 58.2 | / | / |  |  |  |  | 0.98 | 237 |  | Andrea V |  | 3 | 4 | 24459
4 | 60.0 | 112 | P/FINAL | 2025-12-09 | 21:22 | 2025-12-09 | 20:46 | ELENA | 2 | 60.6 | 60.8 | / | / |  |  |  | 1.819 | 2.36 | 239 |  | Andrea V |  | 3 | 4 | 24459
4 | 60.0 | 113 | P/FINAL | 2025-12-09 | 21:22 | 2025-12-09 | 20:46 | ELENA | 2 | 61.4 | 61.3 | / | / |  |  |  |  | 1.26 | 239 |  | Andrea V |  | 3 | 4 | 24459
1 | 0.0 | 114 | P/FINAL | 2025-12-09 | 21:24 | 2025-12-09 | 20:46 | ELENA | 2 | 0.0 | 0.2 | / | / |  |  |  |  | 2.42 | 252 |  | Andrea V |  | 3 | 1 | 24459
1 | 0.0 | 115 | P/FINAL | 2025-12-09 | 21:24 | 2025-12-09 | 20:46 | ELENA | 2 | 0.0 | 0.1 | / | / |  |  |  | 1.793 | 1.59 | 249 |  | Andrea V |  | 3 | 1 | 24459
2 | 15.0 | 116 | P/FINAL | 2025-12-10 | 08:33 | 2025-12-10 | 08:08 | ROLAND | 2 | 15.4 | 15.7 | / | / |  |  |  |  | 1.38 | 244 |  | Kinga |  | 3 | 2 | 24459
2 | 15.0 | 117 | P/FINAL | 2025-12-10 | 08:34 | 2025-12-10 | 08:08 | ROLAND | 2 | 16.5 | 16.6 | / | / |  |  |  |  | 1.05 | 244 |  | Kinga |  | 3 | 2 | 24459
2 | 15.0 | 118 | P/FINAL | 2025-12-10 | 08:35 | 2025-12-10 | 08:08 | ROLAND | 2 | 15.1 | 15.1 | / | / |  |  |  | 1.789 | 1.72 | 241 |  | Kinga |  | 3 | 2 | 24459
3 | 35.0 | 119 | P/FINAL | 2025-12-10 | 08:38 | 2025-12-10 | 08:08 | ROLAND | 2 | 36.9 | 36.8 | / | / |  |  |  |  | 1.33 | 250 |  | Kinga |  | 3 | 3 | 24459
3 | 35.0 | 120 | P/FINAL | 2025-12-10 | 08:38 | 2025-12-10 | 08:08 | ROLAND | 2 | 35.8 | 35.9 | / | / |  |  |  |  | 1.45 | 237 |  | Kinga |  | 3 | 3 | 24459
3 | 35.0 | 121 | P/FINAL | 2025-12-10 | 08:39 | 2025-12-10 | 08:08 | ROLAND | 2 | 37.6 | 37.6 | / | / |  |  |  | 1.790 | 1.46 | 241 |  | Kinga |  | 3 | 3 | 24459
4 | 60.0 | 122 | P/FINAL | 2025-12-10 | 08:43 | 2025-12-10 | 08:08 | ROLAND | 2 | 60.3 | 60.7 | / | / |  |  |  | 1.833 | 1.22 | 237 |  | Kinga |  | 3 | 4 | 24459
4 | 60.0 | 123 | P/FINAL | 2025-12-10 | 08:44 | 2025-12-10 | 08:08 | ROLAND | 2 | 58.8 | 58.9 | / | / |  |  |  |  | 1.16 | 245 |  | Kinga |  | 3 | 4 | 24459
4 | 60.0 | 124 | P/FINAL | 2025-12-10 | 08:45 | 2025-12-10 | 08:08 | ROLAND | 2 | 55.2 | 55.2 | / | / |  |  |  |  | 1.55 | 248 |  | Kinga |  | 3 | 4 | 24459
1 | 0.0 | 125 | P/FINAL | 2025-12-10 | 08:46 | 2025-12-10 | 08:08 | ROLAND | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 2.11 | 243 |  | Kinga |  | 3 | 1 | 24459
1 | 0.0 | 126 | P/FINAL | 2025-12-10 | 08:46 | 2025-12-10 | 08:08 | ROLAND | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 2.16 | 248 |  | Kinga |  | 3 | 1 | 24459
2 | 15.0 | 127 | P/FINAL | 2025-12-10 | 10:50 | 2025-12-10 | 10:27 | ROLAND | 2 | 16.8 | 16.8 | / | / |  |  |  | 1.734 | 5.92 | 250 |  | Kinga |  | 3 | 2 | 24459
2 | 15.0 | 128 | P/FINAL | 2025-12-10 | 10:51 | 2025-12-10 | 10:27 | ROLAND | 2 | 15.8 | 15.8 | / | / |  |  |  |  | 1.68 | 242 |  | Kinga |  | 3 | 2 | 24459
2 | 15.0 | 129 | P/FINAL | 2025-12-10 | 10:52 | 2025-12-10 | 10:27 | ROLAND | 2 | 16.0 | 16.1 | / | / |  |  |  |  | 1.83 | 232 |  | Kinga |  | 3 | 2 | 24459
3 | 35.0 | 130 | P/FINAL | 2025-12-10 | 10:52 | 2025-12-10 | 10:27 | ROLAND | 2 | 34.9 | 34.9 | / | / |  |  |  | 1.773 | 1.82 | 248 |  | Kinga |  | 3 | 3 | 24459
3 | 35.0 | 131 | P/FINAL | 2025-12-10 | 10:53 | 2025-12-10 | 10:27 | ROLAND | 2 | 35.8 | 35.8 | / | / |  |  |  |  | 1.58 | 239 |  | Kinga |  | 3 | 3 | 24459
3 | 35.0 | 132 | P/FINAL | 2025-12-10 | 10:54 | 2025-12-10 | 10:27 | ROLAND | 2 | 37.0 | 37.0 | / | / |  |  |  |  | 1.38 | 243 |  | Kinga |  | 3 | 3 | 24459
4 | 60.0 | 133 | P/FINAL | 2025-12-10 | 10:56 | 2025-12-10 | 10:27 | ROLAND | 2 | 59.7 | 60.4 | / | / |  |  |  |  | 1.54 | 246 |  | Kinga |  | 3 | 4 | 24459
4 | 60.0 | 134 | P/FINAL | 2025-12-10 | 10:57 | 2025-12-10 | 10:27 | ROLAND | 2 | 59.8 | 59.8 | / | / |  |  |  | 1.785 | 4.21 | 243 |  | Kinga |  | 3 | 4 | 24459
1 | 0.0 | 135 | P/FINAL | 2025-12-10 | 10:59 | 2025-12-10 | 10:27 | ROLAND | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 1.31 | 243 |  | Kinga |  | 3 | 1 | 24459
1 | 0.0 | 136 | P/FINAL | 2025-12-10 | 11:00 | 2025-12-10 | 10:27 | ROLAND | 2 | 0.0 | 0.1 | / | / |  |  |  |  | 1.72 | 249 |  | Kinga |  | 3 | 1 | 24459
2 | 15.0 | 137 | P/FINAL | 2025-12-10 | 14:00 | 2025-12-10 | 12:20 | ROLAND | 2 | 14.5 | 14.6 | / | / |  |  |  |  | 1.40 | 244 |  | Kinga |  | 3 | 2 | 24459
2 | 15.0 | 138 | P/FINAL | 2025-12-10 | 14:01 | 2025-12-10 | 12:20 | ROLAND | 2 | 15.1 | 15.2 | / | / |  |  |  |  | 1.67 | 236 |  | Kinga |  | 3 | 2 | 24459
2 | 15.0 | 139 | P/FINAL | 2025-12-10 | 14:02 | 2025-12-10 | 12:20 | ROLAND | 2 | 15.6 | 15.6 | / | / |  |  |  | 1.724 | 1.94 | 247 |  | Kinga |  | 3 | 2 | 24459
3 | 35.0 | 140 | P/FINAL | 2025-12-10 | 14:02 | 2025-12-10 | 12:20 | ROLAND | 2 | 35.1 | 35.2 | / | / |  |  |  |  | 1.64 | 250 |  | Kinga |  | 3 | 3 | 24459
3 | 35.0 | 141 | P/FINAL | 2025-12-10 | 14:04 | 2025-12-10 | 12:20 | ROLAND | 2 | 35.6 | 35.5 | / | / |  |  |  |  | 2.54 | 244 |  | Kinga |  | 3 | 3 | 24459
3 | 35.0 | 142 | P/FINAL | 2025-12-10 | 14:06 | 2025-12-10 | 12:20 | ROLAND | 2 | 36.0 | 35.6 | / | / |  |  |  | 1.696 | 1.58 | 245 |  | Kinga |  | 3 | 3 | 24459
4 | 60.0 | 143 | P/FINAL | 2025-12-10 | 14:07 | 2025-12-10 | 12:20 | ROLAND | 2 | 59.9 | 59.7 | / | / |  |  |  |  | 146 | 245 |  | Kinga |  | 3 | 4 | 24459
4 | 60.0 | 144 | P/FINAL | 2025-12-10 | 14:07 | 2025-12-10 | 12:20 | ROLAND | 2 | 59.3 | 59.5 | / | / |  |  |  | 1.701 | 3.69 | 243 |  | Kinga |  | 3 | 4 | 24459
2 | 15.0 | 145 | P/FINAL | 2025-12-10 | 14:43 | 2025-12-10 | 14:21 | ELENA | 2 | 14.0 | 14.0 | / | / |  |  |  |  | 3.69 | 244 |  | Kinga |  | 3 | 2 | 24459
2 | 15.0 | 146 | P/FINAL | 2025-12-10 | 14:43 | 2025-12-10 | 14:21 | ELENA | 2 | 16.6 | 16.5 | / | / |  |  |  | 1.800 | 2.22 | 228 |  | Kinga |  | 3 | 2 | 24459
2 | 15.0 | 147 | P/FINAL | 2025-12-10 | 14:46 | 2025-12-10 | 14:21 | ELENA | 2 | 16.3 | 16.4 | / | / |  |  |  |  | 2.22 | 245 |  | Kinga |  | 3 | 2 | 24459
3 | 35.0 | 148 | P/FINAL | 2025-12-10 | 14:47 | 2025-12-10 | 14:21 | ELENA | 2 | 31.9 | 31.9 | / | / |  |  |  | 1.807 | 2.22 | 248 |  | Kinga |  | 3 | 3 | 24459
3 | 35.0 | 149 | P/FINAL | 2025-12-10 | 14:48 | 2025-12-10 | 14:21 | ELENA | 2 | 31.6 | 31.5 | / | / |  |  |  |  | 1.35 | 241 |  | Kinga |  | 3 | 3 | 24459
3 | 35.0 | 150 | P/FINAL | 2025-12-10 | 14:48 | 2025-12-10 | 14:21 | ELENA | 2 | 32.4 | 32.6 | / | / |  |  |  |  | 1.39 | 254 |  | Kinga |  | 3 | 3 | 24459
4 | 60.0 | 151 | P/FINAL | 2025-12-10 | 14:49 | 2025-12-10 | 14:21 | ELENA | 2 | 62.7 | 62.6 | / | / |  |  |  | 1.814 | 1.35 | 254 |  | Kinga |  | 3 | 4 | 24459
4 | 60.0 | 152 | P/FINAL | 2025-12-10 | 14:49 | 2025-12-10 | 14:21 | ELENA | 2 | 63.6 | 64.1 | / | / |  |  |  |  | 1.50 | 246 |  | Kinga |  | 3 | 4 | 24459
4 | 60.0 | 153 | P/FINAL | 2025-12-10 | 14:51 | 2025-12-10 | 14:21 | ELENA | 2 | 61.2 | 61.4 | / | / |  |  |  |  | 1.56 | 241 |  | Kinga |  | 3 | 4 | 24459
1 | 0.0 | 154 | P/FINAL | 2025-12-10 | 14:53 | 2025-12-10 | 14:21 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 1.76 | 245 |  | Kinga |  | 3 | 1 | 24459
1 | 0.0 | 155 | P/FINAL | 2025-12-10 | 14:53 | 2025-12-10 | 14:21 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 1.56 | 237 |  | Kinga |  | 3 | 1 | 24459
2 | 15.0 | 156 | P/FINAL | 2025-12-10 | 15:49 | 2025-12-10 | 15:35 | ELENA | 2 | 13.8 | 13.9 | / | / |  |  |  |  | 2.76 |  |  | Kinga |  | 3 | 2 | 24459
2 | 15.0 | 157 | P/FINAL | 2025-12-10 | 15:50 | 2025-12-10 | 15:35 | ELENA | 2 | 15.2 | 15.3 | / | / |  |  |  |  | 1.72 |  |  | Kinga |  | 3 | 2 | 24459
2 | 15.0 | 158 | P/FINAL | 2025-12-10 | 15:51 | 2025-12-10 | 15:35 | ELENA | 2 | 16.2 | 16.3 | / | / |  |  |  |  | 7.23 |  |  | Kinga |  | 3 | 2 | 24459
3 | 35.0 | 159 | P/FINAL | 2025-12-10 | 15:54 | 2025-12-10 | 15:35 | ELENA | 2 | 31.3 | 31.4 | / | / |  |  |  |  | 1.09 |  |  | Kinga |  | 3 | 3 | 24459
3 | 35.0 | 160 | P/FINAL | 2025-12-10 | 15:56 | 2025-12-10 | 15:35 | ELENA | 2 | 32.7 | 32.6 | / | / |  |  |  |  | 1.08 |  |  | Kinga |  | 3 | 3 | 24459
3 | 35.0 | 161 | P/FINAL | 2025-12-10 | 15:56 | 2025-12-10 | 15:35 | ELENA | 2 | 33.4 | 33.6 | / | / |  |  |  |  | 1.12 |  |  | Kinga |  | 3 | 3 | 24459
4 | 60.0 | 162 | P/FINAL | 2025-12-10 | 15:58 | 2025-12-10 | 15:35 | ELENA | 2 | 58.5 | 58.6 | / | / |  |  |  |  | 1.30 |  |  | Kinga |  | 3 | 4 | 24459
4 | 60.0 | 163 | P/FINAL | 2025-12-10 | 16:00 | 2025-12-10 | 15:35 | ELENA | 2 | 59.4 | 59.4 | / | / |  |  |  |  | 1.38 |  |  | Kinga |  | 3 | 4 | 24459
2 | 15.0 | 164 | P/FINAL | 2025-12-10 | 17:12 | 2025-12-10 | 16:48 | ELENA | 2 | 15.2 | 15.2 | / | / |  |  |  |  | 0.79 | 240 |  | Andrea V |  | 3 | 2 | 24459
2 | 15.0 | 165 | P/FINAL | 2025-12-10 | 17:13 | 2025-12-10 | 16:48 | ELENA | 2 | 15.6 | 15.6 | / | / |  |  |  | 1.792 | 0.88 | 239 |  | Andrea V |  | 3 | 2 | 24459
2 | 15.0 | 166 | P/FINAL | 2025-12-10 | 17:13 | 2025-12-10 | 16:48 | ELENA | 2 | 16.5 | 16.6 | / | / |  |  |  |  | 1.52 | 257 |  | Andrea V |  | 3 | 2 | 24459
3 | 35.0 | 167 | P/FINAL | 2025-12-10 | 17:14 | 2025-12-10 | 16:48 | ELENA | 2 | 36.0 | 36.0 | / | / |  |  |  |  | 1.14 | 245 |  | Andrea V |  | 3 | 3 | 24459
3 | 35.0 | 168 | P/FINAL | 2025-12-10 | 17:15 | 2025-12-10 | 16:48 | ELENA | 2 | 37.0 | 37.1 | / | / |  |  |  | 1.783 | 1.03 | 231 |  | Andrea V |  | 3 | 3 | 24459
3 | 35.0 | 169 | P/FINAL | 2025-12-10 | 17:15 | 2025-12-10 | 16:48 | ELENA | 2 | 35.2 | 35.3 | / | / |  |  |  |  | 0.97 | 237 |  | Andrea V |  | 3 | 3 | 24459
4 | 60.0 | 170 | P/FINAL | 2025-12-10 | 17:16 | 2025-12-10 | 16:48 | ELENA | 2 | 56.0 | 55.4 | / | / |  |  |  |  | 0.89 | 231 |  | Andrea V |  | 3 | 4 | 24459
4 | 60.0 | 171 | P/FINAL | 2025-12-10 | 17:18 | 2025-12-10 | 16:48 | ELENA | 2 | 57.2 | 57.3 | / | / |  |  |  | 1.776 | 1.12 | 236 |  | Andrea V |  | 3 | 4 | 24459
1 | 0.0 | 172 | P/FINAL | 2025-12-10 | 17:21 | 2025-12-10 | 16:48 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 6.30 | 227 |  | Andrea V |  | 3 | 1 | 24459
1 | 0.0 | 173 | P/FINAL | 2025-12-10 | 17:21 | 2025-12-10 | 16:48 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  | 1.773 | 1.98 | 261 |  | Andrea V |  | 3 | 1 | 24459
4 | 60.0 | 174 | P/FINAL | 2025-12-10 | 17:29 | 2025-12-10 | 16:48 | ELENA | 2 | 59.1 | 59.6 | / | / |  |  |  |  | 1.38 | 241 |  | Andrea V |  | 3 | 4 | 24459
2 | 15.0 | 175 | P/FINAL | 2025-12-10 | 18:44 | 2025-12-10 | 17:56 | ELENA | 2 | 13.8 | 14.0 | / | / |  |  |  |  | 4.82 | 233 |  | Andrea V |  | 3 | 2 | 24459
2 | 15.0 | 176 | P/FINAL | 2025-12-10 | 18:45 | 2025-12-10 | 17:56 | ELENA | 2 | 16.3 | 16.3 | / | / |  |  |  |  | 0.87 | 238 |  | Andrea V |  | 3 | 2 | 24459
2 | 15.0 | 177 | P/FINAL | 2025-12-10 | 18:45 | 2025-12-10 | 17:56 | ELENA | 2 | 15.1 | 15.2 | / | / |  |  |  | 1.722 | 0.76 | 237 |  | Andrea V |  | 3 | 2 | 24459
3 | 35.0 | 178 | P/FINAL | 2025-12-10 | 18:46 | 2025-12-10 | 17:56 | ELENA | 2 | 36.0 | 36.1 | / | / |  |  |  |  | 1.16 | 231 |  | Andrea V |  | 3 | 3 | 24459
3 | 35.0 | 179 | P/FINAL | 2025-12-10 | 18:47 | 2025-12-10 | 17:56 | ELENA | 2 | 35.7 | 35.7 | / | / |  |  |  | 1.715 | 0.92 | 248 |  | Andrea V |  | 3 | 3 | 24459
3 | 35.0 | 180 | P/FINAL | 2025-12-10 | 18:48 | 2025-12-10 | 17:56 | ELENA | 2 | 37.5 | 37.6 | / | / |  |  |  |  | 1.08 | 258 |  | Andrea V |  | 3 | 3 | 24459
4 | 60.0 | 181 | P/FINAL | 2025-12-10 | 18:49 | 2025-12-10 | 17:56 | ELENA | 2 | 58.4 | 58.8 | / | / |  |  |  |  | 1.26 | 232 |  | Andrea V |  | 3 | 4 | 24459
4 | 60.0 | 182 | P/FINAL | 2025-12-10 | 18:49 | 2025-12-10 | 17:56 | ELENA | 2 | 57.8 | 59.4 | / | / |  |  |  | 1.745 | 1.26 | 245 |  | Andrea V |  | 3 | 4 | 24459
4 | 60.0 | 183 | P/FINAL | 2025-12-10 | 18:50 | 2025-12-10 | 17:56 | ELENA | 2 | 57.8 | 58.0 | / | / |  |  |  |  | 0.99 | 237 |  | Andrea V |  | 3 | 4 | 24459
1 | 0.0 | 184 | P/FINAL | 2025-12-10 | 18:53 | 2025-12-10 | 17:56 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 3.22 | 245 |  | Andrea V |  | 3 | 1 | 24459
1 | 0.0 | 185 | P/FINAL | 2025-12-10 | 18:54 | 2025-12-10 | 17:56 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  | 1.749 | 0.80 | 235 |  | Andrea V |  | 3 | 1 | 24459
2 | 15.0 | 186 | P/FINAL | 2025-12-10 | 20:06 | 2025-12-10 | 19:21 | ELENA | 2 | 14.1 | 14.3 | / | / |  |  |  |  | 1.42 | 232 |  | Andrea V |  | 3 | 2 | 24459
2 | 15.0 | 187 | P/FINAL | 2025-12-10 | 20:07 | 2025-12-10 | 19:21 | ELENA | 2 | 14.9 | 15.0 | / | / |  |  |  | 1.776 | 0.78 | 243 |  | Andrea V |  | 3 | 2 | 24459
2 | 15.0 | 188 | P/FINAL | 2025-12-10 | 20:07 | 2025-12-10 | 19:21 | ELENA | 2 | 16.2 | 16.0 | / | / |  |  |  |  | 0.99 | 244 |  | Andrea V |  | 3 | 2 | 24459
3 | 35.0 | 189 | P/FINAL | 2025-12-10 | 20:08 | 2025-12-10 | 19:21 | ELENA | 2 | 34.1 | 34.1 | / | / |  |  |  |  | 0.76 | 236 |  | Andrea V |  | 3 | 3 | 24459
3 | 35.0 | 190 | P/FINAL | 2025-12-10 | 20:09 | 2025-12-10 | 19:21 | ELENA | 2 | 32.9 | 32.8 | / | / |  |  |  |  | 0.90 | 223 |  | Andrea V |  | 3 | 3 | 24459
3 | 35.0 | 191 | P/FINAL | 2025-12-10 | 20:09 | 2025-12-10 | 19:21 | ELENA | 2 | 35.4 | 35.4 | / | / |  |  |  | 1.758 | 0.71 | 240 |  | Andrea V |  | 3 | 3 | 24459
4 | 60.0 | 192 | P/FINAL | 2025-12-10 | 20:11 | 2025-12-10 | 19:21 | ELENA | 2 | 57.3 | 58.0 | / | / |  |  |  |  | 1.31 | 245 |  | Andrea V |  | 3 | 4 | 24459
4 | 60.0 | 193 | P/FINAL | 2025-12-10 | 20:11 | 2025-12-10 | 19:21 | ELENA | 2 | 55.3 | 55.3 | / | / |  |  |  |  | 1.17 | 255 |  | Andrea V |  | 3 | 4 | 24459
4 | 60.0 | 194 | P/FINAL | 2025-12-10 | 20:12 | 2025-12-10 | 19:21 | ELENA | 2 | 56.2 | 56.1 | / | / |  |  |  |  | 1.10 | 248 |  | Andrea V |  | 3 | 4 | 24459
1 | 0.0 | 195 | P/FINAL | 2025-12-10 | 20:13 | 2025-12-10 | 19:21 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 1.82 | 246 |  | Andrea V |  | 3 | 1 | 24459
1 | 0.0 | 196 | P/FINAL | 2025-12-10 | 20:13 | 2025-12-10 | 19:21 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  | 1.762 | 4.09 | 240 |  | Andrea V |  | 3 | 1 | 24459
2 | 15.0 | 197 | P/FINAL | 2025-12-10 | 21:28 | 2025-12-10 | 20:51 | ELENA | 2 | 17.5 | 17.4 | / | / |  |  |  |  | 1.53 | 244 |  | Andrea V |  | 3 | 2 | 24459
2 | 15.0 | 198 | P/FINAL | 2025-12-10 | 21:29 | 2025-12-10 | 20:51 | ELENA | 2 | 15.8 | 15.8 | / | / |  |  |  | 1.731 | 1.14 | 241 |  | Andrea V |  | 3 | 2 | 24459
2 | 15.0 | 199 | P/FINAL | 2025-12-10 | 21:29 | 2025-12-10 | 20:51 | ELENA | 2 | 16.0 | 16.0 | / | / |  |  |  |  | 0.91 | 240 |  | Andrea V |  | 3 | 2 | 24459
3 | 35.0 | 200 | P/FINAL | 2025-12-10 | 21:30 | 2025-12-10 | 20:51 | ELENA | 2 | 36.8 | 36.8 | / | / |  |  |  |  | 1.50 | 233 |  | Andrea V |  | 3 | 3 | 24459
3 | 35.0 | 201 | P/FINAL | 2025-12-10 | 21:30 | 2025-12-10 | 20:51 | ELENA | 2 | 35.5 | 35.5 | / | / |  |  |  | 1.740 | 1.05 | 238 |  | Andrea V |  | 3 | 3 | 24459
3 | 35.0 | 202 | P/FINAL | 2025-12-10 | 21:31 | 2025-12-10 | 20:51 | ELENA | 2 | 32.5 | 32.6 | / | / |  |  |  |  | 0.80 | 252 |  | Andrea V |  | 3 | 3 | 24459
4 | 60.0 | 203 | P/FINAL | 2025-12-10 | 21:32 | 2025-12-10 | 20:51 | ELENA | 2 | 59.4 | 59.4 | / | / |  |  |  |  | 1.26 | 242 |  | Andrea V |  | 3 | 4 | 24459
4 | 60.0 | 204 | P/FINAL | 2025-12-10 | 21:33 | 2025-12-10 | 20:51 | ELENA | 2 | 58.3 | 58.4 | / | / |  |  |  | 1.749 | 1.08 | 243 |  | Andrea V |  | 3 | 4 | 24459
4 | 60.0 | 205 | P/FINAL | 2025-12-10 | 21:33 | 2025-12-10 | 20:51 | ELENA | 2 | 57.8 | 57.9 | / | / |  |  |  |  | 1.00 | 238 |  | Andrea V |  | 3 | 4 | 24459
1 | 0.0 | 206 | P/FINAL | 2025-12-10 | 21:35 | 2025-12-10 | 20:51 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 6.62 | 240 |  | Andrea V |  | 3 | 1 | 24459
1 | 0.0 | 207 | P/FINAL | 2025-12-10 | 21:35 | 2025-12-10 | 20:51 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  | 1.761 | 2.23 | 241 |  | Andrea V |  | 3 | 1 | 24459
2 | 15.0 | 208 | P/FINAL | 2025-12-11 | 08:27 | 2025-12-11 | 07:57 | ROLAND | 2 | 16.6 | 16.6 | / | / |  |  |  |  | 2.49 | 230 |  | Kinga |  | 3 | 2 | 121534
2 | 15.0 | 209 | P/FINAL | 2025-12-11 | 08:29 | 2025-12-11 | 07:57 | ROLAND | 2 | 17.6 | 17.6 | / | / |  |  |  |  | 2.88 | 220 |  | Kinga |  | 3 | 2 | 121534
2 | 15.0 | 210 | P/FINAL | 2025-12-11 | 08:29 | 2025-12-11 | 07:57 | ROLAND | 2 | 14.4 | 14.5 | / | / |  |  |  | 1.785 | 3.42 | 234 |  | Kinga |  | 3 | 2 | 121534
4 | 60.0 | 211 | P/FINAL | 2025-12-11 | 08:33 | 2025-12-11 | 07:57 | ROLAND | 2 | 59.0 | 59.0 | / | / |  |  |  |  | 1.99 | 232 |  | Kinga |  | 3 | 4 | 121534
4 | 60.0 | 212 | P/FINAL | 2025-12-11 | 08:34 | 2025-12-11 | 07:57 | ROLAND | 2 | 60.4 | 60.7 | / | / |  |  |  |  | 2.07 | 230 |  | Kinga |  | 3 | 4 | 121534
4 | 60.0 | 213 | P/FINAL | 2025-12-11 | 08:35 | 2025-12-11 | 07:57 | ROLAND | 2 | 62.8 | 63.2 | / | / |  |  |  | 1.781 | 2.02 | 245 |  | Kinga |  | 3 | 4 | 121534
3 | 35.0 | 214 | P/FINAL | 2025-12-11 | 08:42 | 2025-12-11 | 07:57 | ROLAND | 2 | 38.1 | 38.2 | / | / |  |  |  |  | 2.02 | 232 |  | Kinga |  | 3 | 3 | 121534
1 | 0.0 | 215 | P/FINAL | 2025-12-11 | 08:44 | 2025-12-11 | 07:57 | ROLAND | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 2.20 | 241 |  | Kinga |  | 3 | 1 | 121534
1 | 0.0 | 216 | P/FINAL | 2025-12-11 | 08:44 | 2025-12-11 | 07:57 | ROLAND | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 1.36 | 252 |  | Kinga |  | 3 | 1 | 121534
3 | 35.0 | 217 | P/FINAL | 2025-12-11 | 08:52 | 2025-12-11 | 07:57 | ROLAND | 2 | 37.7 | 37.9 | / | / |  |  |  | 1.789 | 2.56 | 238 |  | Kinga |  | 3 | 3 | 121534
3 | 35.0 | 218 | P/FINAL | 2025-12-11 | 08:54 | 2025-12-11 | 07:57 | ROLAND | 2 | 36.3 | 36.6 | / | / |  |  |  |  | 3.85 | 240 |  | Kinga |  | 3 | 3 | 121534
2 | 15.0 | 219 | Old A | 2025-12-11 | 10:49 | 2025-12-11 |  | ROLAND | 2 | 14.0 | 14.2 | / | / |  |  |  |  | 3.35 | 253 |  | Kinga |  | 3 | 2 | 121534
2 | 15.0 | 220 | Old A | 2025-12-11 | 10:51 | 2025-12-11 |  | ROLAND | 2 | 15.8 | 15.9 | / | / |  |  |  | 1.717 | 2.14 | 241 |  | Kinga |  | 3 | 2 | 121534
3 | 35.0 | 221 | Old A | 2025-12-11 | 10:52 | 2025-12-11 |  | ROLAND | 2 | 34.1 | 34.2 | / | / |  |  |  |  | 1.51 | 254 |  | Kinga |  | 3 | 3 | 121534
3 | 35.0 | 222 | Old A | 2025-12-11 | 10:53 | 2025-12-11 |  | ROLAND | 2 | 35.6 | 35.7 | / | / |  |  |  |  | 1.56 | 247 |  | Kinga |  | 3 | 3 | 121534
4 | 60.0 | 223 | Old A | 2025-12-11 | 10:53 | 2025-12-11 |  | ROLAND | 2 | 61.3 | 60.9 | / | / |  |  |  |  | 2.28 | 242 |  | Kinga |  | 3 | 4 | 121534
4 | 60.0 | 224 | Old A | 2025-12-11 | 10:54 | 2025-12-11 |  | ROLAND | 2 | 60.6 | 60.9 | / | / |  |  |  |  | 2.52 | 234 |  | Kinga |  | 3 | 4 | 121534
2 | 15.0 | 225 | P/FINAL | 2025-12-11 | 11:31 | 2025-12-11 | 11:05 | ROLAND | 2 | 12.6 | 12.6 | / | / |  |  |  |  | 1.71 | 245 |  | Kinga |  | 3 | 2 | 121534
2 | 15.0 | 226 | P/FINAL | 2025-12-11 | 11:32 | 2025-12-11 | 11:05 | ROLAND | 2 | 12.7 | 12.7 | / | / |  |  |  |  | 1.33 | 251 |  | Kinga |  | 3 | 2 | 121534
2 | 15.0 | 227 | P/FINAL | 2025-12-11 | 11:32 | 2025-12-11 | 11:05 | ROLAND | 2 | 12.7 | 12.7 | / | / |  |  |  | 1.713 | 1.33 | 253 |  | Kinga |  | 3 | 2 | 121534
3 | 35.0 | 228 | P/FINAL | 2025-12-11 | 11:33 | 2025-12-11 | 11:05 | ROLAND | 2 | 32.6 | 32.9 | / | / |  |  |  |  | 4.04 | 241 |  | Kinga |  | 3 | 3 | 121534
3 | 35.0 | 229 | P/FINAL | 2025-12-11 | 11:34 | 2025-12-11 | 11:05 | ROLAND | 2 | 36.3 | 36.5 | / | / |  |  |  |  | 2.92 | 254 |  | Kinga |  | 3 | 3 | 121534
3 | 35.0 | 230 | P/FINAL | 2025-12-11 | 11:42 | 2025-12-11 | 11:05 | ROLAND | 2 | 37.7 | 37.4 | / | / |  |  |  | 1.688 | 1.42 | 247 |  | Kinga |  | 3 | 3 | 121534
3 | 35.0 | 231 | P/FINAL | 2025-12-11 | 11:42 | 2025-12-11 | 11:05 | ROLAND | 2 | 37.7 | 37.4 | / | / |  |  |  |  | 1.42 | 242 |  | Kinga |  | 3 | 3 | 121534
4 | 60.0 | 232 | P/FINAL | 2025-12-11 | 11:44 | 2025-12-11 | 11:05 | ROLAND | 2 | 55.8 | 55.7 | / | / |  |  |  |  | 2.37 | 234 |  | Kinga |  | 3 | 4 | 121534
4 | 60.0 | 233 | P/FINAL | 2025-12-11 | 11:45 | 2025-12-11 | 11:05 | ROLAND | 2 | 55.5 | 55.7 | / | / |  |  |  |  | 2.37 | 245 |  | Kinga |  | 3 | 4 | 121534
4 | 60.0 | 234 | P/FINAL | 2025-12-11 | 11:45 | 2025-12-11 | 11:05 | ROLAND | 2 | 55.5 | 55.7 | / | / |  |  |  |  | 2.37 | 251 |  | Kinga |  | 3 | 4 | 121534
4 | 60.0 | 235 | P/FINAL | 2025-12-11 | 11:45 | 2025-12-11 | 11:05 | ROLAND | 2 | 55.5 | 55.7 | / | / |  |  |  | 1.705 | 2.37 | 245 |  | Kinga |  | 3 | 4 | 121534
4 | 60.0 | 236 | P/FINAL | 2025-12-11 | 11:54 | 2025-12-11 | 11:05 | ROLAND | 2 | 59.4 | 59.5 | / | / |  |  |  | 1.708 | 1.714 | 246 |  | Kinga |  | 3 | 4 | 121534
2 | 15.0 | 237 | P/FINAL | 2025-12-11 | 13:58 | 2025-12-11 | 12:51 | ROLAND | 2 | 14.1 | 14.1 | / | / |  |  |  |  | 1.91 | 241 |  | Kinga |  | 3 | 2
2 | 15.0 | 238 | P/FINAL | 2025-12-11 | 13:59 | 2025-12-11 | 12:51 | ROLAND | 2 | 14.8 | 14.9 | / | / |  |  |  |  | 2.10 | 252 |  | Kinga |  | 3 | 2
2 | 15.0 | 239 | P/FINAL | 2025-12-11 | 14:00 | 2025-12-11 | 12:51 | ROLAND | 2 | 14.4 | 14.3 | / | / |  |  |  | 1.745 | 2.13 | 238 |  | Kinga |  | 3 | 2
3 | 35.0 | 240 | P/FINAL | 2025-12-11 | 14:00 | 2025-12-11 | 12:51 | ROLAND | 2 | 35.1 | 35.1 | / | / |  |  |  |  | 2.03 | 244 |  | Kinga |  | 3 | 3
3 | 35.0 | 241 | P/FINAL | 2025-12-11 | 14:03 | 2025-12-11 | 12:51 | ROLAND | 2 | 35.0 | 35.1 | / | / |  |  |  |  | 1.55 | 240 |  | Kinga |  | 3 | 3
3 | 35.0 | 242 | P/FINAL | 2025-12-11 | 14:03 | 2025-12-11 | 12:51 | ROLAND | 2 | 35.4 | 35.5 | / | / |  |  |  | 1.735 | 2.64 | 233 |  | Kinga |  | 3 | 3
4 | 60.0 | 243 | P/FINAL | 2025-12-11 | 14:17 | 2025-12-11 | 12:51 | ROLAND | 2 | 55.9 | 55.9 | / | / |  |  |  |  | 2.87 | 242 |  | Kinga |  | 3 | 4
4 | 60.0 | 244 | P/FINAL | 2025-12-11 | 14:17 | 2025-12-11 | 12:51 | ROLAND | 2 | 62.3 | 62.4 | / | / |  |  |  |  | 2.98 | 249 |  | Kinga |  | 3 | 4
1 | 0.0 | 245 | P/FINAL | 2025-12-11 | 14:19 | 2025-12-11 | 12:51 | ROLAND | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 2.13 | 248 |  | Kinga |  | 3 | 1
1 | 0.0 | 246 | P/FINAL | 2025-12-11 | 14:19 | 2025-12-11 | 12:51 | ROLAND | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 2.13 | 251 |  | Kinga |  | 3 | 1
2 | 15.0 | 247 | P/FINAL | 2025-12-11 | 14:39 | 2025-12-11 | 14:21 | ELENA | 2 | 16.0 | 16.2 | / | / |  |  |  |  | 1.35 | 239 |  | Kinga |  | 3 | 2
2 | 15.0 | 248 | P/FINAL | 2025-12-11 | 14:39 | 2025-12-11 | 14:21 | ELENA | 2 | 15.6 | 15.7 | / | / |  |  |  | 1.741 | 3.91 | 245 |  | Kinga |  | 3 | 2
2 | 15.0 | 249 | P/FINAL | 2025-12-11 | 14:40 | 2025-12-11 | 14:21 | ELENA | 2 | 15.4 | 15.5 | / | / |  |  |  |  | 3.99 | 238 |  | Kinga |  | 3 | 2
3 | 35.0 | 250 | P/FINAL | 2025-12-11 | 14:41 | 2025-12-11 | 14:21 | ELENA | 2 | 36.1 | 36.1 | / | / |  |  |  |  | 1.54 | 248 |  | Kinga |  | 3 | 3
3 | 35.0 | 251 | P/FINAL | 2025-12-11 | 14:42 | 2025-12-11 | 14:21 | ELENA | 2 | 36.9 | 37.3 | / | / |  |  |  | 1.739 | 1.47 | 253 |  | Kinga |  | 3 | 3
3 | 35.0 | 252 | P/FINAL | 2025-12-11 | 14:42 | 2025-12-11 | 14:21 | ELENA | 2 | 37.1 | 37.2 | / | / |  |  |  |  | 1.51 | 239 |  | Kinga |  | 3 | 3
4 | 60.0 | 253 | P/FINAL | 2025-12-11 | 14:43 | 2025-12-11 | 14:21 | ELENA | 2 | 57.3 | 57.7 | / | / |  |  |  |  | 2.61 | 248 |  | Kinga |  | 3 | 4
4 | 60.0 | 254 | P/FINAL | 2025-12-11 | 14:43 | 2025-12-11 | 14:21 | ELENA | 2 | 57.3 | 57.7 | / | / |  |  |  |  | 2.61 | 247 |  | Kinga |  | 3 | 4
4 | 60.0 | 255 | P/FINAL | 2025-12-11 | 14:43 | 2025-12-11 | 14:21 | ELENA | 2 | 59.5 | 59.1 | / | / |  |  |  |  | 1.72 | 232 |  | Kinga |  | 3 | 4
4 | 60.0 | 256 | P/FINAL | 2025-12-11 | 14:44 | 2025-12-11 | 14:21 | ELENA | 2 | 61.1 | 61.3 | / | / |  |  |  | 1.770 | 1.97 | 345 |  | Kinga |  | 3 | 4
2 | 15.0 | 257 | P/FINAL | 2025-12-11 | 15:41 | 2025-12-11 | 15:26 | ELENA | 2 | 15.5 | 15.5 | / | / |  |  |  |  | 5.03 | 247 |  | Kinga |  | 3 | 2
2 | 15.0 | 258 | P/FINAL | 2025-12-11 | 15:42 | 2025-12-11 | 15:26 | ELENA | 2 | 16.1 | 16.1 | / | / |  |  |  |  | 2.71 | 247 |  | Kinga |  | 3 | 2
2 | 15.0 | 259 | P/FINAL | 2025-12-11 | 15:44 | 2025-12-11 | 15:26 | ELENA | 2 | 16.7 | 16.7 | / | / |  |  |  | 1.778 | 3.40 | 239 |  | Kinga |  | 3 | 2
3 | 35.0 | 260 | P/FINAL | 2025-12-11 | 15:45 | 2025-12-11 | 15:26 | ELENA | 2 | 33.5 | 33.2 | / | / |  |  |  |  | 1.32 | 254 |  | Kinga |  | 3 | 3
3 | 35.0 | 261 | P/FINAL | 2025-12-11 | 15:45 | 2025-12-11 | 15:26 | ELENA | 2 | 35.8 | 35.6 | / | / |  |  |  | 1.778 | 3.87 | 248 |  | Kinga |  | 3 | 3
4 | 60.0 | 262 | P/FINAL | 2025-12-11 | 15:47 | 2025-12-11 | 15:26 | ELENA | 2 | 57.6 | 56.9 | / | / |  |  |  |  | 1.62 | 245 |  | Kinga |  | 3 | 4
4 | 60.0 | 263 | P/FINAL | 2025-12-11 | 15:47 | 2025-12-11 | 15:26 | ELENA | 2 | 59.2 | 59.1 | / | / |  |  |  |  | 1.79 | 251 |  | Kinga |  | 3 | 4
4 | 60.0 | 264 | P/FINAL | 2025-12-11 | 15:48 | 2025-12-11 | 15:26 | ELENA | 2 | 60.1 | 60.4 | / | / |  |  |  |  | 1.84 | 241 |  | Kinga |  | 3 | 4
2 | 15.0 | 265 | P/FINAL | 2025-12-12 | 08:31 | 2025-12-11 | 17:50 | ELENA | 2 | 14.6 | 14.6 | / | / |  |  |  |  | 1.79 | 38 |  | Kinga |  | 3 | 2
2 | 15.0 | 266 | P/FINAL | 2025-12-12 | 08:32 | 2025-12-11 | 17:50 | ELENA | 2 | 15.2 | 15.3 | / | / |  |  |  |  | 1.81 | 238 |  | Kinga |  | 3 | 2
2 | 15.0 | 267 | P/FINAL | 2025-12-12 | 08:32 | 2025-12-11 | 17:50 | ELENA | 2 | 15.7 | 15.8 | / | / |  |  |  |  | 1.89 | 240 |  | Kinga |  | 3 | 2
3 | 35.0 | 268 | P/FINAL | 2025-12-12 | 08:33 | 2025-12-11 | 17:50 | ELENA | 2 | 38.1 | 38.1 | / | / |  |  |  |  | 2.69 | 242 |  | Kinga |  | 3 | 3
3 | 35.0 | 269 | P/FINAL | 2025-12-12 | 08:34 | 2025-12-11 | 17:50 | ELENA | 2 | 36.3 | 36.4 | / | / |  |  |  | 1.687 | 2.50 | 256 |  | Kinga |  | 3 | 3
3 | 35.0 | 270 | P/FINAL | 2025-12-12 | 08:34 | 2025-12-11 | 17:50 | ELENA | 2 | 36.1 | 36.2 | / | / |  |  |  |  | 2.83 | 247 |  | Kinga |  | 3 | 3
4 | 60.0 | 271 | P/FINAL | 2025-12-12 | 08:35 | 2025-12-11 | 17:50 | ELENA | 2 | 61.2 | 61.3 | / | / |  |  |  |  | 2.87 | 239 |  | Kinga |  | 3 | 4
4 | 60.0 | 272 | P/FINAL | 2025-12-12 | 08:36 | 2025-12-11 | 17:50 | ELENA | 2 | 59.4 | 59.5 | / | / |  |  |  |  | 2.82 | 237 |  | Kinga |  | 3 | 4
4 | 60.0 | 273 | P/FINAL | 2025-12-12 | 08:38 | 2025-12-11 | 17:50 | ELENA | 2 | 60.9 | 61.0 | / | / |  |  |  | 1.711 | 2.98 | 251 |  | Kinga |  | 3 | 4
1 | 0.0 | 274 | P/FINAL | 2025-12-12 | 08:38 | 2025-12-11 | 17:50 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 3.00 | 247 |  | Kinga |  | 3 | 1
1 | 0.0 | 275 | P/FINAL | 2025-12-12 | 08:38 | 2025-12-11 | 17:50 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  |  | 3.00 | 246 |  | Kinga |  | 3 | 1
1 | 0.0 | 276 | P/FINAL | 2025-12-12 | 08:39 | 2025-12-11 | 17:50 | ELENA | 2 | 0.0 | 0.0 | / | / |  |  |  | 1.719 | 3.01 | 242 |  | Kinga |  | 3 | 1
Standard # 2
Total Tests | Total Readings | Total Average | Selected Tests | Selected Readings | Selected Average
69 | 138 | 15.5 mg/L | 69 | 138 | 15.5 mg/L
Standard # 3
Total Tests | Total Readings | Total Average | Selected Tests | Selected Readings | Selected Average
69 | 138 | 35.6 mg/L | 69 | 138 | 35.6 mg/L
Standard # 4
Total Tests | Total Readings | Total Average | Selected Tests | Selected Readings | Selected Average
67 | 134 | 59.4 mg/L | 67 | 134 | 59.4 mg/L
Standard # 1
Total Tests | Total Readings | Total Average | Selected Tests | Selected Readings | Selected Average
35 | 70 | 0 mg/L | 35 | 70 | 0 mg/L
Specifications
Standard Value [ppm] | Target Value ÷ U [ppm] | Mean Value [ppm] | Passed
0.0 | 0.0 ÷ 2.0 | 0.0 | YES
15.0 | 12.3 ÷ 17.7 | 15.5 | YES
35.0 | 31.3 ÷ 38.7 | 35.6 | YES
60.0 | 55.1 ÷ 64.9 | 59.4 | YES
Validation Date | by
Laura