## MR Warehouse
Updated |  | 2022-09-10 | List of Reference Material from chemical QC | [20 empty] | MR.LIST Rev0.1
|  |  | Riduction EXP days | 60 | 90 | [8 empty] | 120 | [8 empty] | 2022-09-10
MR HANNA Code | MR Bottle | MR Description | MR Supplier | MR MPN | MR Lot | MR Physical State | MR Density 20°C g/mL | MR Purity (%) | MR Value | MR Unit | MR Parameter | FW MR Parameter | MR Location | MR Storage | MR Stock q.ty | MR Stock Unit | Min.QTY | Arrived | Open | Finished | Supplier EXP | MR EXP | STATUS | Note | STATUS
MR09 | A | Ammonium 1000 mg/l NH₄ traceable to SRM from NIST NH₄Cl in H₂O | MERCK | 1198120500 | HC17846912 | L | 0.99 |  | 994 | mg/L | NH4 | 18 | ricordarsi la location | 15 ÷ 25°c | 100 | dipende da colonna G se L = mL se S = g | 100 | 2022-03-18 | 2022-03-22 | Data | 2024-11-30 | 2024-08-02 | in stock |  | in stock
MR09 | B | Ammonium 1000 mg/l NH₄ traceable to SRM from NIST NH₄Cl in H₂O | MERCK | 1198120500 | HC17846912 | L | 0.99 |  | 994 | mg/L | NH4 | 18 |  | 15 ÷ 25°c | dato che si aggiorna ogni prelieveo |  | 100 | 2022-03-18 |  |  | 2024-11-30 | FALSE | in stock |  | finished
MR09 | C | Ammonium 1000 mg/l NH₄ traceable to SRM from NIST NH₄Cl in H₂O | MERCK | 1198120500 | HC17846912 | L | 0.99 |  | 994 | mg/L | NH4 | 18 |  | 15 ÷ 25°c |  |  | 100 | 2022-03-18 |  |  | 2024-11-30 | FALSE | in stock
MR09 | D | Ammonium 1000 mg/l NH₄ traceable to SRM from NIST NH₄Cl in H₂O | MERCK | 1198120500 | HC17846912 | L | 0.99 |  | 994 | mg/L | NH4 | 18 |  | 15 ÷ 25°c |  |  | 100 | 2022-03-18 |  |  | 2024-11-30 | FALSE | in stock
MR100 | A | Magnesium 10000 mg/l Mg traceable to SRM from NIST Mg(NO₃)₂ in HNO₃ 2-3% | MERCK | 1703790100 | HC03708079 |  |  |  | #REF! | mg/L | Mg | 24.3 |  | 15 ÷ 25°c |  |  | 20 | 2022-03-18 |  |  | 2024-08-31 | FALSE | in stock
MR100 | B | Magnesium 10000 mg/l Mg traceable to SRM from NIST Mg(NO₃)₂ in HNO₃ 2-3% | MERCK | 1703790100 | HC03708079 |  |  |  | #REF! | mg/L | Mg | 24.3 |  | 15 ÷ 25°c |  |  | 20 | 2022-03-18 |  |  | 2024-08-31 | FALSE | in stock
MR100 | C | Magnesium 10000 mg/l Mg traceable to SRM from NIST Mg(NO₃)₂ in HNO₃ 2-3% | MERCK | 1703790100 | HC03708079 |  |  |  | #REF! | mg/L | Mg | 24.3 |  | 15 ÷ 25°c |  |  | 20 | 2022-03-18 |  |  | 2024-08-31 | FALSE | in stock
MR101 | A | Nitrate standard traceable to SRM from NIST NaNO₃ in H₂O 1000 mg/l NO₃ Certipur | MERCK | 1198110500 | HC17263711 |  |  |  | 997 | mg/L | NO3 | 61.97 |  | 15 ÷ 25°c
MR102 | A | NitrIte standard traceable to SRM from NIST NaNO₂ in H₂O 1000 mg/l NO₂ Certipur | MERCK | 1198990500 | HC16350999 |  |  |  | 998 | mg/L | NO2 | 45.98 |  | 15 ÷ 25°c
MR103 | A | Alkalinity, CaCO3 1000 mg/L Calibration Standard CRM | MERCK | ALK1000-500ML | HC03708079 |  |  |  | 1002 | mg/L | CaCO3 | 100.0869 |  | 2 ÷ 8°c |  |  | 20 | 2022-03-18 | 2022-03-15 |  | 2024-09-30 | 2024-06-02 | in stock
MR103 | B | Alkalinity, CaCO3 1000 mg/L Calibration Standard CRM | MERCK | ALK1000-500ML | HC03708079 |  |  |  | 1002 | mg/L | CaCO3 | 100.0869 | Refrigerator 1 | 2 ÷ 8°c |  |  | 20 | 2022-03-18 |  |  | 2024-09-30 | FALSE | in stock
MR103 | C | Alkalinity, CaCO3 1000 mg/L Calibration Standard CRM | MERCK | ALK1000-500ML | HC03708079 |  |  |  | 1002 | mg/L | CaCO3 | 100.0869 |  | 2 ÷ 8°c |  |  | 20 | 2022-03-18 |  |  | 2024-09-30 | FALSE | in stock
MR104 | C | Phosphate standard solution traceable to SRM from NIST KH₂PO₄ in H₂O 1000 mg/l PO₄ | MERCK | 1198980500 | [5 empty] | mg/L | PO4 | 94.97 |  | 15 ÷ 25°c |  |  | 100 |  |  |  |  | FALSE
MR18 | A | Fluoride standard solution 1000 mg/l F traceable to SRM from NIST NaF in H₂O | FLUKA | 1198140500 | HC17401114 |  |  |  | 999 | mg/L | F | 18.998 |  | 15 ÷ 25°c |  |  | 200 | 2022-03-18 |  |  | 2024-10-31 | FALSE | in stock
MR18 | B | Fluoride standard solution 1000 mg/l F traceable to SRM from NIST NaF in H₂O | FLUKA | 1198140500 | HC17401114 |  |  |  | 999 | mg/L | F | 18.998 |  | 15 ÷ 25°c |  |  | 200 | 2022-03-18 |  |  | 2024-10-31 | FALSE | in stock
MR19 |  | Molybdenum Standard AAS TraceCERT®, 1000 mg/L Mo in hydrochloric acid | FLUKA | 67210-250ML | [5 empty] | mg/L | Mo | 95.94 |  | 15 ÷ 25°c |  |  | 50 | [5 empty] | in stock
MR19 |  | Molybdenum Standard AAS TraceCERT®, 1000 mg/L Mo in hydrochloric acid | FLUKA | 67210-250ML | [5 empty] | mg/L | Mo | 95.94 |  | 15 ÷ 25°c |  |  | 50 | [5 empty] | in stock
MR19 |  | Molybdenum Standard AAS TraceCERT®, 1000 mg/L Mo in hydrochloric acid | FLUKA | 67210-250ML | [5 empty] | mg/L | Mo | 95.94 |  | 15 ÷ 25°c |  |  | 50 | [5 empty] | in stock
MR19 |  | Molybdenum Standard AAS TraceCERT®, 1000 mg/L Mo in hydrochloric acid | FLUKA | 67210-250ML | [5 empty] | mg/L | Mo | 95.94 |  | 15 ÷ 25°c |  |  | 50 | [5 empty] | in stock
MR19 |  | Molybdenum Standard AAS TraceCERT®, 1000 mg/L Mo in hydrochloric acid | FLUKA | 67210-250ML | [5 empty] | mg/L | Mo | 95.94 |  | 15 ÷ 25°c |  |  | 50 | [5 empty] | in stock
MR19 |  | Molybdenum Standard AAS TraceCERT®, 1000 mg/L Mo in hydrochloric acid | FLUKA | 67210-250ML | [5 empty] | mg/L | Mo | 95.94 |  | 15 ÷ 25°c |  |  | 50 | [5 empty] | in stock
MR20 | A | Calcium standard traceable to SRM from NIST Ca(NO₃)₂ in HNO₃ 0.5 mol/l 1000 mg/l Ca Certipur | MERCK | 1197780500 | HC16142778 |  |  |  | #REF! | mg/L | Ca | 40.078 |  | 15 ÷ 25°c |  |  | 200 | 2022-03-18 |  |  |  | FALSE | in stock
MR20 | B | Calcium standard traceable to SRM from NIST Ca(NO₃)₂ in HNO₃ 0.5 mol/l 1000 mg/l Ca Certipur | MERCK | 1197780500 | HC16142778 |  |  |  | #REF! | mg/L | Ca | 40.078 |  | 15 ÷ 25°c |  |  | 200 | 2022-03-18 |  |  |  | FALSE | in stock
MR21 |  | Nickel Standard for ICP TraceCERT®, 1000 mg/L Ni in nitric acid | FLUKA | 28944-100ML-F | [5 empty] | mg/L | Ni | 58.6934 |  | 15 ÷ 25°c |  |  | 50 | [5 empty] | in stock
MR21 |  | Nickel Standard for ICP TraceCERT®, 1000 mg/L Ni in nitric acid | FLUKA | 28944-100ML-F | [5 empty] | mg/L | Ni | 58.6934 |  | 15 ÷ 25°c |  |  | 50 | [5 empty] | in stock
MR21 |  | Nickel Standard for ICP TraceCERT®, 1000 mg/L Ni in nitric acid | FLUKA | 28944-100ML-F | [5 empty] | mg/L | Ni | 58.6934 |  | 15 ÷ 25°c |  |  | 50 | [5 empty] | in stock
MR21 |  | Nickel Standard for ICP TraceCERT®, 1000 mg/L Ni in nitric acid | FLUKA | 28944-100ML-F | [5 empty] | mg/L | Ni | 58.6934 |  | 15 ÷ 25°c |  |  | 50 | [5 empty] | in stock
MR26 | A | Nickel Standard for ICP TraceCERT®, 10,000 mg/L Ni - 100 mL | FLUKA | 19013-100ML | BCCF4303 |  |  |  | 9994 | mg/L | Ni | 58.6934 |  | 15 ÷ 25°c |  |  | 50 | 2022-03-18 |  |  | 2025-01-01
MR33 |  | Iron Standard AAS TraceCERT®, 1000 mg/L Fe in nitric acid | FLUKA | 16596-250ML | [5 empty] | mg/L | Fe(III) | 55.84 |  | 15 ÷ 25°c |  |  | 100 | [5 empty] | in stock
MR33 |  | Iron Standard AAS TraceCERT®, 1000 mg/L Fe in nitric acid | FLUKA | 16596-250ML | [5 empty] | mg/L | Fe(III) | 55.84 |  | 15 ÷ 25°c |  |  | 100 | [5 empty] | in stock
MR33 |  | Iron Standard AAS TraceCERT®, 1000 mg/L Fe in nitric acid | FLUKA | 16596-250ML | [5 empty] | mg/L | Fe(III) | 55.84 |  | 15 ÷ 25°c |  |  | 100 | [5 empty] | in stock
MR33 |  | Iron Standard AAS TraceCERT®, 1000 mg/L Fe in nitric acid | FLUKA | 16596-250ML | [5 empty] | mg/L | Fe(III) | 55.84 |  | 15 ÷ 25°c |  |  | 100 | [5 empty] | in stock
MR33 |  | Iron Standard AAS TraceCERT®, 1000 mg/L Fe in nitric acid | FLUKA | 16596-250ML | [5 empty] | mg/L | Fe(III) | 55.84 |  | 15 ÷ 25°c |  |  | 100 | [5 empty] | in stock
MR33 |  | Iron Standard AAS TraceCERT®, 1000 mg/L Fe in nitric acid | FLUKA | 16596-250ML | [5 empty] | mg/L | Fe(III) | 55.84 |  | 15 ÷ 25°c |  |  | 100 | [5 empty] | in stock
MR33 |  | Iron Standard AAS TraceCERT®, 1000 mg/L Fe in nitric acid | FLUKA | 16596-250ML | [5 empty] | mg/L | Fe(III) | 55.84 |  | 15 ÷ 25°c |  |  | 100 | [5 empty] | in stock
MR37 |  | Manganese Standard for AAS TraceCERT®, 1000 mg/L Mn in nitric acid | FLUKA | 77036-250 | [5 empty] | mg/L | Mn | 54.94 |  | 15 ÷ 25°c |  |  | 50 | [5 empty] | in stock
MR37 |  | Manganese Standard for AAS TraceCERT®, 1000 mg/L Mn in nitric acid | FLUKA | 77036-250 | [5 empty] | mg/L | Mn | 54.94 |  | 15 ÷ 25°c |  |  | 50 | [5 empty] | in stock
MR37 |  | Manganese Standard for AAS TraceCERT®, 1000 mg/L Mn in nitric acid | FLUKA | 77036-250 | [5 empty] | mg/L | Mn | 54.94 |  | 15 ÷ 25°c |  |  | 50 | [5 empty] | in stock
MR37 |  | Manganese Standard for AAS TraceCERT®, 1000 mg/L Mn in nitric acid | FLUKA | 77036-250 | [5 empty] | mg/L | Mn | 54.94 |  | 15 ÷ 25°c |  |  | 50 | [5 empty] | in stock
MR37 |  | Manganese Standard for AAS TraceCERT®, 1000 mg/L Mn in nitric acid | FLUKA | 77036-250 | [5 empty] | mg/L | Mn | 54.94 |  | 15 ÷ 25°c |  |  | 50 | [5 empty] | in stock
MR37 |  | Manganese Standard for AAS TraceCERT®, 1000 mg/L Mn in nitric acid | FLUKA | 77036-250 | [5 empty] | mg/L | Mn | 54.94 |  | 15 ÷ 25°c |  |  | 50 | [5 empty] | in stock
MR37 |  | Manganese Standard for AAS TraceCERT®, 1000 mg/L Mn in nitric acid | FLUKA | 77036-250 | [5 empty] | mg/L | Mn | 54.94 |  | 15 ÷ 25°c |  |  | 50 | [5 empty] | in stock
MR41 |  | Aluminium standard solution traceable NIST 1000 mg/l Al Certipur® - 500 mL | MERCK | 1197700500 | [5 empty] | mg/L | Al | 26.98 |  | 15 ÷ 25°c |  |  | 100 | [5 empty] | in stock
MR43 |  | Cr VI Standard for AAS TraceCERT®, 1000 mg/L Cr in nitric acid | FLUKA | 02733-250ML | [5 empty] | mg/L | Cr(VI) | 51.996 |  | 15 ÷ 25°c |  |  | 100 | [5 empty] | in stock
MR43 |  | Cr VI Standard for AAS TraceCERT®, 1000 mg/L Cr in nitric acid | FLUKA | 02733-250ML | [5 empty] | mg/L | Cr(VI) | 51.996 |  | 15 ÷ 25°c |  |  | 100 | [5 empty] | in stock
MR43 |  | Cr VI Standard for AAS TraceCERT®, 1000 mg/L Cr in nitric acid | FLUKA | 02733-250ML | [5 empty] | mg/L | Cr(VI) | 51.996 |  | 15 ÷ 25°c |  |  | 100 | [5 empty] | in stock
MR43 |  | Cr VI Standard for AAS TraceCERT®, 1000 mg/L Cr in nitric acid | FLUKA | 02733-250ML | [5 empty] | mg/L | Cr(VI) | 51.996 |  | 15 ÷ 25°c |  |  | 100 | [5 empty] | in stock
MR43 |  | Cr VI Standard for AAS TraceCERT®, 1000 mg/L Cr in nitric acid | FLUKA | 02733-250ML | [5 empty] | mg/L | Cr(VI) | 51.996 |  | 15 ÷ 25°c |  |  | 100 | [5 empty] | in stock
MR43 |  | Cr VI Standard for AAS TraceCERT®, 1000 mg/L Cr in nitric acid | FLUKA | 02733-250ML | [5 empty] | mg/L | Cr(VI) | 51.996 |  | 15 ÷ 25°c |  |  | 100 | [5 empty] | in stock
MR43 |  | Cr VI Standard for AAS TraceCERT®, 1000 mg/L Cr in nitric acid | FLUKA | 02733-250ML | [5 empty] | mg/L | Cr(VI) | 51.996 |  | 15 ÷ 25°c |  |  | 100 | [5 empty] | in stock
MR44 |  | Copper Standard for AAS TraceCERT®, 1000 mg/L Cu in nitric acid | FLUKA | 38996-250ml | [5 empty] | mg/L | Cu | 63.546 |  | 15 ÷ 25°c |  |  | 50 | [5 empty] | in stock
MR44 |  | Copper Standard for AAS TraceCERT®, 1000 mg/L Cu in nitric acid | FLUKA | 38996-250ml | [5 empty] | mg/L | Cu | 63.546 |  | 15 ÷ 25°c |  |  | 51 | [5 empty] | in stock
MR44 |  | Copper Standard for AAS TraceCERT®, 1000 mg/L Cu in nitric acid | FLUKA | 38996-250ml | [5 empty] | mg/L | Cu | 63.546 |  | 15 ÷ 25°c |  |  | 52 | [5 empty] | in stock
MR44 |  | Copper Standard for AAS TraceCERT®, 1000 mg/L Cu in nitric acid | FLUKA | 38996-250ml | [5 empty] | mg/L | Cu | 63.546 |  | 15 ÷ 25°c |  |  | 53 | [5 empty] | in stock
MR44 |  | Copper Standard for AAS TraceCERT®, 1000 mg/L Cu in nitric acid | FLUKA | 38996-250ml | [5 empty] | mg/L | Cu | 63.546 |  | 15 ÷ 25°c |  |  | 54 | [5 empty] | in stock
MR44 |  | Copper Standard for AAS TraceCERT®, 1000 mg/L Cu in nitric acid | FLUKA | 38996-250ml | [5 empty] | mg/L | Cu | 63.546 |  | 15 ÷ 25°c |  |  | 55 | [5 empty] | in stock
MR44 |  | Copper Standard for AAS TraceCERT®, 1000 mg/L Cu in nitric acid | FLUKA | 38996-250ml | [5 empty] | mg/L | Cu | 63.546 |  | 15 ÷ 25°c |  |  | 56 | [5 empty] | in stock
MR44 |  | Copper Standard for AAS TraceCERT®, 1000 mg/L Cu in nitric acid | FLUKA | 38996-250ml | [5 empty] | mg/L | Cu | 63.546 |  | 15 ÷ 25°c |  |  | 57 | [5 empty] | in stock
MR44 |  | Copper Standard for AAS TraceCERT®, 1000 mg/L Cu in nitric acid | FLUKA | 38996-250ml | [5 empty] | mg/L | Cu | 63.546 |  | 15 ÷ 25°c |  |  | 58 | [5 empty] | in stock
MR45 | A | Silicon STD AAS TraceCERT®, 1000 mg/L Si in NaOH | FLUKA | 16259-250ML | [5 empty] | mg/L | Si | 28.0855 |  |  |  |  | 100 | [5 empty] | in stock
MR45 | B | Silicon STD AAS TraceCERT®, 1000 mg/L Si in NaOH | FLUKA | 16259-250ML | [5 empty] | mg/L | Si | 28.0855 |  |  |  |  | 100 | [5 empty] | in stock
MR45 | C | Silicon STD AAS TraceCERT®, 1000 mg/L Si in NaOH | FLUKA | 16259-250ML | [5 empty] | mg/L | Si | 28.0855 |  |  |  |  | 100 | [5 empty] | in stock
MR45 | D | Silicon STD AAS TraceCERT®, 1000 mg/L Si in NaOH | FLUKA | 16259-250ML | [5 empty] | mg/L | Si | 28.0855 |  |  |  |  | 100 | [5 empty] | in stock
MR45 | E | Silicon STD AAS TraceCERT®, 1000 mg/L Si in NaOH | FLUKA | 16259-250ML | [5 empty] | mg/L | Si | 28.0855 |  | 2 ÷ 8°c |  |  | 100 | [5 empty] | in stock
MR50 | A | COD Standard Solution 50000 mg/L, CRM traceable to NIST | MERCK | 1250350100 | HC15713535 |  |  |  | 48980 | mg/L | COD | 1 |  | 15 ÷ 25°c |  |  | 50 | 2022-03-18 | 2022-03-14 |  | 2024-03-31 | 2023-12-02 | in stock
MR50 | B | COD Standard Solution 50000 mg/L, CRM traceable to NIST | MERCK | 1250350100 | HC15713535 |  |  |  | 48980 | mg/L | COD | 1 |  | 15 ÷ 25°c |  |  | 50 | 2022-03-18 |  |  | 2024-03-31 | FALSE | in stock
MR60 | A | Calcium standard 10,000 mg/L Ca TraceCERT in nitric acid | Fluka | 94458-100ML | BCCG7647 |  |  |  | 9999 | mg/L | Ca | 40.078 |  | 15 ÷ 25°c |  |  | 50 | 2022-03-18 |  |  | 2025-10-01 | FALSE | in stock
MR60 | B | Calcium standard 10,000 mg/L Ca TraceCERT in nitric acid | Fluka | 94458-100ML | BCCG7647 |  |  |  | 9999 | mg/L | Ca | 40.078 |  | 15 ÷ 25°c |  |  | 50 | 2022-03-18 |  |  | 2025-10-01 | FALSE | in stock
MR62 |  | Zinc Standard for AAS TraceCERT®, 1000 mg/L Zn in nitric acid | Fluka | 18827-250ML | [5 empty] | mg/L | Zn | 65.409 |  |  |  |  | 100 | [5 empty] | in stock
MR62 |  | Zinc Standard for AAS TraceCERT®, 1000 mg/L Zn in nitric acid | Fluka | 18827-250ML | [5 empty] | mg/L | Zn | 65.409 |  |  |  |  | 100 | [5 empty] | in stock
MR62 |  | Zinc Standard for AAS TraceCERT®, 1000 mg/L Zn in nitric acid | Fluka | 18827-250ML | [5 empty] | mg/L | Zn | 65.409 |  |  |  |  | 100 | [5 empty] | in stock
MR67 | A | Potassium permanganate 0.002 mol (0,01N) | MERCK | 1603221000 | HC15731422 |  |  |  | 0.002 | mg/L | KMn04 | 63.1772 |  | 15 ÷ 25°c |  |  | 400 | 2022-03-18 | 2022-03-14 |  | 2024-04-30 | 2024-01-01 | in stock
MR67 | B | Potassium permanganate 0.002 mol (0,01N) | MERCK | 1603221000 | HC15731422 |  |  |  | 0.002 | mg/L | KMn04 | 63.1772 |  | 15 ÷ 25°c |  |  | 400 | 2022-03-18 |  |  | 2024-04-30 | FALSE | in stock
MR67 | C | Potassium permanganate 0.002 mol (0,01N) | MERCK | 1603221000 | HC15731422 |  |  |  | 0.002 | mg/L | KMn04 | 63.1772 |  | 15 ÷ 25°c |  |  | 400 | 2022-03-18 |  |  | 2024-04-30 | FALSE | in stock
MR74 |  | Cadmium ion standard solution | Fluka | 20893 | [5 empty] | mg/L | [6 empty] | 100 | [5 empty] | in stock
MR74 |  | Cadmium ion standard solution | Fluka | 20893 | [5 empty] | mg/L | [6 empty] | 100 | [5 empty] | in stock
MR75 |  | Olive Oil | Sigma | O1514 | [5 empty] | mg/L | [6 empty] | 500
MR75 |  | Olive Oil | Sigma | O1514 | [5 empty] | mg/L | [6 empty] | 500
MR75 |  | Olive Oil | Sigma | O1514 | [5 empty] | mg/L | [6 empty] | 500
MR75 |  | Olive Oil | Sigma | O1514 | [5 empty] | mg/L | [6 empty] | 500
MR75 |  | Olive Oil | Sigma Aldrich | 75343 | [5 empty] | mg/L |  |  |  | 15 ÷ 25°c |  |  | 1L | [5 empty] | in stock
MR77 |  | Magnesium ion solution for ISE | Fluka | 63031 | [5 empty] | mg/L |  |  |  | 2 ÷ 8°c |  |  | 100 | [5 empty] | in stock
MR77 |  | Magnesium ion solution for ISE | Fluka | 63031 | [5 empty] | mg/L | [6 empty] | 100 | [5 empty] | in stock
MR77 |  | Magnesium ion solution for ISE | Fluka | 63031 | [5 empty] | mg/L | [6 empty] | 100 | [5 empty] | in stock
MR82 |  | Silicon ICP standard traceable to SRM from NIST (SiO₂ in NaOH 4%) 10000 mg/l Si | MERCK | 1703860100 | [5 empty] | mg/L | Si | 28.0855 |  |  |  |  | 100 | [5 empty] | in stock
MR82 |  | Silicon ICP standard traceable to SRM from NIST (SiO₂ in NaOH 4%) 10000 mg/l Si | MERCK | 1703860100 | [5 empty] | mg/L | Si | 28.0855 |  |  |  |  | 100 | [5 empty] | in stock
MR82 |  | Silicon ICP standard traceable to SRM from NIST (SiO₂ in NaOH 4%) 10000 mg/l Si | MERCK | 1703860100 | [5 empty] | mg/L | Si | 28.0855 |  |  |  |  | 100 | [5 empty] | in stock
MR82 |  | Silicon ICP standard traceable to SRM from NIST (SiO₂ in NaOH 4%) 10000 mg/l Si | MERCK | 1703860100 | [5 empty] | mg/L | Si | 28.0855 |  |  |  |  | 100 | [5 empty] | in stock
MR82 |  | Silicon ICP standard traceable to SRM from NIST (SiO₂ in NaOH 4%) 10000 mg/l Si | MERCK | 1703860100 | [5 empty] | mg/L | Si | 28.0855 |  |  |  |  | 100 | [5 empty] | in stock
MR82 |  | Silicon ICP standard traceable to SRM from NIST (SiO₂ in NaOH 4%) 10000 mg/l Si | MERCK | 1703860100 | [5 empty] | mg/L | Si | 28.0855 |  |  |  |  | 100 | [5 empty] | in stock
MR89 |  | N,N- Diethylhydroxylamine | Sigma Aldrich | 471593 | [5 empty] | mg/L |  |  |  | 2 ÷ 8°c |  |  | 100 | [5 empty] | in stock
MR91 | A | Chromium 1000 mg/l Cr traceable to SRM from NIST Cr(NO₃)₃ in HNO₃ 0.5 mol/l | Merck | 1197790100 | HC03326679 |  |  |  | 986 | mg/L | Cr III | 51.996 |  | 15 ÷ 25°c |  |  | 100 | 2022-03-18 |  |  | 2024-07-31 | FALSE | in stock
MR91 | B | Chromium 1000 mg/l Cr traceable to SRM from NIST Cr(NO₃)₃ in HNO₃ 0.5 mol/l | Merck | 1197790100 | HC03326679 |  |  |  | 986 | mg/L | Cr III | 51.996 |  | 15 ÷ 25°c |  |  | 100 | 2022-03-18 |  |  | 2024-07-31 | FALSE | in stock
MR93 |  | Magnesium standard traceable to SRM from NIST Mg(NO₃)₂ in HNO₃ 0.5 mol/l 1000 mg/l Mg Certipur | Sigma Aldrich | 1197880500 | [5 empty] | mg/L | Mg | 24.305 |  | 15 ÷ 25°c |  |  | 100 |  |  |  |  | FALSE | in stock | For Sea water /NEW CODE
MR93 | A | Buffer solution pH 12.00 | Certipur Merck | 1.99022.0001 | [9 empty] | 15 ÷ 25°c |  |  | 300 | [5 empty] | in stock
MR93 | B | Buffer solution pH 12.454 | VWR Chemicals | 84587.26 | [9 empty] | 15 ÷ 25°c |  |  | 300 | [5 empty] | in stock
MR05 |  | Calcium carbonate BioUltra ≥99.0% (KT) - 250 gr | Merck | 21061 |  |  |  | 99% |  | g | CaCO3 | 100.0869 |  | 15 ÷ 25°c |  |  | 300 | [5 empty] | in stock
Formulas:
MR EXP = IF([Open]>0,[Supplier EXP]-O2)

## QC Soft Data Base
UPDATE WK | Code | Description | Line | Recipe | QCMethod | MeterFamily1 | MeterFamily2 | ParameterMethod | ParameterFormula | MeasurementUnit | RangeMin | RangeMax | Decimal | Fixed | AndOr | Percentage | QCRestriction | STDMR | STD1Value | STD1Min | STD1Max | STD2Value | STD2Min | STD2Max | STD3Value | STD3Min | STD3Max | STD4Value | STD4Min | STD4Max | STD5Value | STD5Min | STD5Max | STD6Value | STD6Min | STD6Max | pH1Value | pH1Min | pH1Max | pH2Value | pH2Min | pH2Max | pH3Value | pH3Min | pH3Max | WeightValue | WeightMin | WeightMax | Certified | Note
WK47/2021 | BISULFITE/N | Sodium Metabisulfite Reagent | L57 Powder | R001 | CTK | HI801 | HI83399 | Nitrogen Total LR(N) | Nitrogen | mg/L | 0 | 25 | 1 | 1 | or | 5% | 100% | MR09 | 0 | 0 | 1 | 20 | 19 | 22 | / | / | / | / | / | / | / | / | / | / | / | / | 4.3 | 4 | 4.5 | / | / | / | / | / | / | 175 | 150 | 200
WK47/2021 | BUFFER REAG | Buffer Reagent | L56 CTK | B004 | colorimetric | / | / | Calcium | Calcium | mg/L | 0 | 400 | 0 | 10 | & | 5% | 100% | MR20 | 0 | 0 | 10 | 80 | 66 | 94 | 320 | 294 | 346 | / | / | / | / | / | / | / | / | / | [6 empty] | / | / | / | / | / | /
WK47/2021 | CARBON | Activated Carbon | L57 Powder | R002 | CTK | / | / | / | / | / | / | / | / | 0 | / | 0% | / |  | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | 8 | 7.9 | 8.2 | / | / | / | / | / | / | 230 | 210 | 250
WK47/2021 | DPD1-0 | Free Chlorine Reagent | L56 CTK | B136 | colorimetric | HI801 | HI83399 | Free Chlorine (liquid) | Chlorine | mg/L | 0 | 5 | 2 | 0.03 | & | 3% | 100% | MR67 | 0 | 0 | 0.03 | 0.5 | 0.46 | 0.54 | 2 | 1.91 | 2.09 | / | / | / | / | / | / | / | / | / | 0.2 | 0 | 0.4 | 6 | 5.7 | 6.3 | / | / | / | / | / | /
WK47/2021 | DPD2-0 | Free Chlorine Reagent 2 | L56 CTK | B137 | colorimetric | HI801 | HI83399 | Free Chlorine (liquid) | Chlorine | mg/L | 0 | 5 | 2 | 0.03 | & | 3% | 100% | MR67 | 0 | 0 | 0.03 | 0.5 | 0.46 | 0.54 | 2 | 1.91 | 2.09 | / | / | / | / | / | / | / | / | / | 0.2 | 0 | 0.4 | 6 | 5.7 | 6.3 | / | / | / | / | / | /
WK47/2021 | HI38000A-0 | Sulfate Reagent A | L57 Powder | R014 | CTK | / | / | Sulfate Test Kit | Sulphate | mg/L | 20 | 100 | 0 | 5 | or | 0% | 100% | MR32 | 20 | 15 | 25 | 80 | 75 | 90 | / | / | / | / | / | / | / | / | / | / | / | / | 1.8 | 1.7 | 2 | / | / | / | / | / | / | 210 | 195 | 225
WK47/2021 | HI38001A-0 | Sulfate LR Reagent A | L57 Powder | R014 | CTK | / | / | Sulfate Test Kit LR | Sulphate | mg/L | 100 | 1000 | 0 | 10 | or | 0% | 100% | MR32 | 100 | 90 | 110 | 800 | 790 | 810 | / | / | / | / | / | / | / | / | / | / | / | / | 1.8 | 1.7 | 2 | / | / | / | / | / | / | 210 | 195 | 225
WK47/2021 | HI38001A-0 | Sulfate HR Reagent A | L57 Powder | R014 | CTK | / | / | Sulfate Test Kit HR | Sulphate | mg/L | 1000 | 10000 | 0 | 100 | or | 0% | 100% | MR32 | 8000 | 7900 | 8100 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | 1.8 | 1.7 | 2 | / | / | / | / | / | / | 210 | 195 | 225
WK47/2021 | HI38050-0 | Nitrate Reagent | L57 Powder | R045 | CTK | / | / | Nitrate Test Kit | Nitrate | mg/L | 0 | 50 | 1 | 1 | or | 0% | 100% | MR40 | / | / | / | 10 | 9 | 11 | 40 | 39 | 41 | / | / | / | / | / | / | / | / | / | 2.7 | 2.5 | 3 | / | / | / | / | / | / | 250 | 230 | 270
WK47/2021 | HI38067A-0 | Silica Reagent A | L56 CTK | B145 | CTK | / | / | Silica High Range Test Kit (SiO2) LR | Silica Oxide | mg/L | 0 | 40 | 0 | 1 | or | 0% | / | MR82 | 0 | 0 | 1 | 21 | 20 | 22 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI38067B-0 | Silica High Range Reagent B | L57 Powder | R014 | CTK | / | / | Silica High Range Test Kit (SiO2) LR | Silica Oxide | mg/L | 0 | 40 | 0 | 1 | or | 0% | 100% | MR82 | 0 | 0 | 1 | 21 | 20 | 22 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | 150 | 135 | 165
WK47/2021 | HI38067C-0 | Silica High Range Reagent C | L57 Powder | R015 | CTK | / | / | Silica High Range Test Kit (SiO2) LR | Silica Oxide | mg/L | 0 | 40 | 0 | 1 | or | 0% | 100% | MR82 | 0 | 0 | 1 | 21 | 20 | 22 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | 260 | 230 | 290
WK47/2021 | HI38082B-0 | Potassium Reagent B | L57 Powder | R016 | CTK | / | / | Potassium Test Kit | Potassium |  |  |  | 1 | 0 | / | 0% | 100% | MR30 | 0 | 0 | 0 | 15 | 15 | 15 | 40 | 40 | 40 | / | / | / | / | / | / | / | / | / | 13 | 12.8 | 13.2 | 6 | 5.9 | 6.2 | / | / | / | 190 | 175 | 205
WK47/2021 | HI3833-0 | Phosphate Reagent | L57 Powder | R029 | CTK | / | / | Phosphate Test Kit | Potassium | mg/L | 0 | 5 | 0 | 1 | or | 0% | 100% | MR30 | / | / | / | 2 | 1 | 3 | 4 | 3 | 5 | / | / | / | / | / | / | / | / | / | 1.5 | 1.2 | 1.6 | / | / | / | / | / | / | 150 | 135 | 165
WK47/2021 | HI3834-0 | Iron Reagent | L57 Powder | R041 | CTK | / | / | Iron Test Kit | Iron | mg/L | 0 | 5 | 0 | 1 | or | 0% | 100% | MR33 | / | / | / | 2 | 1 | 3 | 4 | 3 | 5 | / | / | / | / | / | / | / | / | / | 5.2 | 4 | 6.5 | / | / | / | / | / | / | 125 | 110 | 140
WK47/2021 | HI38434B | Bleach Reagent B | L57 Powder | R003 | CTK | / | / | Chlorine (Cl2) | Chlorine | g/L | 50 | 150 | 0 | 5 | or | 0% | 100% | MR67 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | 1 | 0.8 | 1.2 | / | / | / | / | / | / | 300 | 280 | 320
WK47/2021 | HI3843B-0 | Bleach Reagent B | L57 Powder | R003 | CTK | / | / | Chlorine (Cl2) | Chlorine | g/L | 50 | 150 | 0 | 5 | or | 0% | 100% | MR67 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | 1 | 0.8 | 1.2 | / | / | / | / | / | / | 300 | 280 | 320
WK47/2021 | HI3845A-0 | Chromium MR & HR Reagent A | L57 Powder | R003 | CTK | / | / | Chromium (VI) HR | Chromium VI | mg/L | 0 | 1000 | 0 | 50 | or | 0% | 100% | MR43 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | 1 | 0.8 | 1.2 | / | / | / | / | / | / | 160 | 145 | 175
WK47/2021 | HI3845A-0 | Chromium MR & HR Reagent A | L57 Powder | R003 | CTK | / | / | Chromium (VI) MR | Chromium VI | mg/L | 0 | 100 | 0 | 5 | or | 0% | 100% | MR43 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | 1 | 0.8 | 1.2 | / | / | / | / | / | / | 160 | 145 | 175
WK47/2021 | HI3846-0 | Chromium Reagent | L57 Powder | R043 | CTK | / | / | Chromium Test Kit | Chromium | mg/L | 0 | 1 | 1 | 0.2 | or | 0% | 100% | MR43 | / | / | / | 0.2 | 0 | 0.4 | 0.6 | 0.4 | 0.8 | / | / | / | / | / | / | / | / | / | 1.3 | 1.1 | 1.6 | / | / | / | / | / | / | 150 | 140 | 160
WK47/2021 | HI3847-0 | Copper Reagent | L57 Powder | R026 | CTK | / | / | Copper Test Kit | Copper | mg/L | 0 | 2.5 | 1 | 0.5 | or | 0% | 100% | MR44 | / | / | / | 0.5 | 0 | 1 | 2 | 1.5 | 2.5 | / | / | / | / | / | / | / | / | / | 6.4 | 6.2 | 6.7 | / | / | / | / | / | / | 140 | 130 | 150
WK47/2021 | HI3855A-0 | Cyanide Reagent A | L57 Powder | R036 | CTK | / | / | Cyanide Test Kit LR | Cyanide | mg/L | 0 | 0.15 | 2 | 0.01 | or | 0% | 100% | CM327 | / | / | / | 0.08 | 0.07 | 0.09 | 0.14 | 0.13 | 0.15 | / | / | / | / | / | / | / | / | / | 6.8 | 6.7 | 7 | / | / | / | 6.4 | 6.2 | 6.8 | 17005 | 16805 | 17205
WK47/2021 | HI3855B-0 | Cyanide Reagent B | L57 Powder | R039 | CTK | / | / | Cyanide Test Kit LR | Cyanide | mg/L | 0 | 0.15 | 2 | 0.01 | or | 0% | 100% | CM327 | / | / | / | 0.08 | 0.07 | 0.09 | 0.14 | 0.13 | 0.15 | / | / | / | / | / | / | / | / | / | 6.8 | 6.7 | 7 | / | / | / | 6.4 | 6.2 | 6.8 | 200 | 190 | 210
WK47/2021 | HI3855C-0 | Cyanide Reagent C | L57 Powder | R040 | CTK | / | / | Cyanide Test Kit LR | Cyanide | mg/L | 0 | 0.15 | 2 | 0.01 | or | 0% | 100% | CM327 | / | / | / | 0.08 | 0.07 | 0.09 | 0.14 | 0.13 | 0.15 | / | / | / | / | / | / | / | / | / | 6.8 | 6.7 | 7 | / | / | / | 6.4 | 6.2 | 6.8 | 95 | 85 | 105
WK47/2021 | HI3856-0 | Silica High Range Reagent A | L57 Powder | R026 | CTK | / | / | Copper Ultra Low Range Test Kit | Copper | mg/L | 0 | 0.25 | 2 | 0.05 | or | 0% | 100% | MR44 | / | / | / | 0.1 | 0.05 | 0.15 | 0.2 | 0.15 | 0.25 | / | / | / | / | / | / | / | / | / | 6.4 | 6.2 | 6.7 | / | / | / | / | / | / | 240 | 225 | 255
WK47/2021 | HI3859B-0 | Glycol Reagent B | L57 Powder | R004 | CTK | / | / | Glycol Test Kit | Glycol | mg/L | 30 |  | 0 | 0 | / | 0% | 100% |  | 25 | 25 | 25 | 50 | 50 | 50 | 100 | 100 | 100 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | 170 | 150 | 190
WK47/2021 | HI3859C-0 | Glycol Reagent C | L57 Powder | R005 | CTK | / | / | Glycol Test Kit | Glycol | mg/L | 30 |  | 0 | 0 | / | 0% | 100% |  | 25 | 25 | 25 | 50 | 50 | 50 | 100 | 100 | 100 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | 200 | 180 | 220
WK47/2021 | HI3864A-0 | Phenol Reagent A | L57 Powder | R006 | CTK | / | / | Phenols Test Kit HR | Phenols | mg/L | 0.5 | 5 | 2 | 0.1 | or | 0% | 100% |  | 0.5 | 0.4 | 0.6 | 2.5 | 2.4 | 2.6 | 4.5 | 4.4 | 4.5 | / | / | / | / | / | / | / | / | / | 11 | 10.9 | 11.2 | 11 | 10.9 | 11.2 | / | / | / | 150 | 140 | 160
WK47/2021 | HI3864B-0 | Phenol Reagent B | L57 Powder | R007 | CTK | / | / | Phenols Test Kit HR | Phenols | mg/L | 0.5 | 5 | 2 | 0.1 | or | 0% | 100% |  | 0.5 | 0.4 | 0.6 | 2.5 | 2.4 | 2.6 | 4.5 | 4.4 | 4.5 | / | / | / | / | / | / | / | / | / | 11 | 10.9 | 11.2 | 11 | 10.9 | 11.2 | / | / | / | 155 | 145 | 165
WK47/2021 | HI3873-0 | Nitrite Reagent | L57 Powder | R031 | CTK | / | / | Nitrite Test Kit | Nitrite | mg/L | 0 | 1 | 1 | 0.2 | or | 0% | 100% | MR13 | / | / | / | 0.2 | 0 | 0.4 | 0.8 | 0.6 | 1 | / | / | / | / | / | / | / | / | / | 3 | 2.9 | 3.1 | / | / | / | / | / | / | 270 | 260 | 280
WK47/2021 | HI3874-0 | Nitrate Reagent | L57 Powder | R045 | CTK | / | / | Nitrate Test Kit | Nitrate | mg/L | 0 | 50 | 0 | 10 | or | 0% | 100% | MR40 | / | / | / | 20 | 10 | 30 | 40 | 30 | 50 | / | / | / | / | / | / | / | / | / | 2.7 | 2.5 | 3 | / | / | / | / | / | / | 180 | 165 | 195
WK47/2021 | HI3895K-0 | Potassium Reagent | L57 Powder | R008 | CTK | / | / | Potassium Test | Potassium | mg/L |  |  | 1 | 0 | / | 0% | 100% | MR30 | 16 | 16 | 16 | 36 | 36 | 36 | 72 | 72 | 72 | / | / | / | / | / | / | / | / | / | 11.7 | 11.5 | 11.8 | / | / | / | / | / | / | 105 | 95 | 115
WK47/2021 | HI3895N-0 | Nitrogen Reagent | L57 Powder | R009 | CTK | / | / | Nitrogen Test | Nitrogen | mg/L |  |  | 1 | 0 | / | 0% | 100% | MR09 | 1.6 | 1.6 | 1.6 | 3.2 | 3.2 | 3.2 | 6.5 | 6.5 | 6.5 | / | / | / | / | / | / | / | / | / | 1.9 | 1.8 | 2.1 | / | / | / | / | / | / | 290 | 270 | 310
WK47/2021 | HI3895P-0 | Phosphorus Reagent | L57 Powder | R029 | CTK | / | / | Phosphorus Test | Phosphorus | mg/L |  |  | 1 | 0 | / | 0% | 100% | MR24 | 0.7 | 0.7 | 0.7 | 1.3 | 1.3 | 1.3 | 2.3 | 2.3 | 2.3 | / | / | / | / | / | / | / | / | / | 1.5 | 1.2 | 1.6 | / | / | / | / | / | / | 80 | 65 | 95
WK47/2021 | HI3895PH-0 | pH Reagent | L57 Powder | R010 | CTK | / | / | pH Test | pH | pH | 4 | 9 | / | 0 | / | 0% | / | pH 4 - 6 - 8 | 4 | 4 | 4 | 6 | 6 | 6 | 8 | 8 | 8 | / | / | / | / | / | / | / | / | / | 8.6 | 8.5 | 8.8 | / | / | / | / | / | / | 305 | 290 | 320
WK47/2021 | HI3896K-0 | Potassium Reagent | L57 Powder | R008 | CTK | / | / | Potassium Test | Potassium | mg/L | 0 | 72 | 0 | 0 | / | 0% | 100% | MR30 | 16 | 16 | 16 | 36 | 36 | 36 | 72 | 72 | 72 | / | / | / | / | / | / | / | / | / | 11.7 | 11.5 | 11.8 | / | / | / | / | / | / | 105 | 95 | 115
WK47/2021 | HI3896N-0 | Nitrogen Reagent | L57 Powder | R009 | CTK | / | / | Nitrogen Test | Nitrogen | mg/L | 0 | 7 | 0 | 0 | / | 0% | 100% | MR09 | 2 | 2 | 2 | 3 | 3 | 3 | 7 | 7 | 7 | / | / | / | / | / | / | / | / | / | 1.9 | 1.8 | 2.1 | / | / | / | / | / | / | 290 | 270 | 310
WK47/2021 | HI3896P-0 | Phosphorus Reagent | L57 Powder | R029 | CTK | / | / | Phosphorus Test | Phosphorus | mg/L | 0 | 2 | 0 | 0 | / | 0% | 100% | MR24 | 0 | 0 | 0 | 1 | 1 | 1 | 2 | 2 | 2 | / | / | / | / | / | / | / | / | / | 1.5 | 1.2 | 1.6 | / | / | / | / | / | / | 80 | 65 | 95
WK47/2021 | HI4000-47-4 | pH 4 Buffer with Chloride Ion Background | L57 Powder | R011 | [11 empty] | % | 100% |  | / | / | / |  | 0 |  |  |  |  | / | / | / | / | / | / | / | / | /
WK47/2021 | HI4000-47-7 | pH 7 Buffer with Chloride Ion Background | L57 Powder | R012 | [11 empty] | % | 100% |  | / | / | / |  | 0 |  |  |  |  | / | / | / | / | / | / | / | / | /
WK47/2021 | HI700A-0 | Ammonia Low Range Reagent A | L56 CTK | B134 | colorimetric | HI700 | / | Ammonia Low Range | N-Ammonia | mg/L | 0 | 3 | 2 | 0.05 | & | 5% | 100% | MR09 | 0 | 0 | 0.05 | 0.6 | 0.52 | 0.68 | 2.4 | 2.23 | 2.57 | / | / | / | / | / | / | / | / | / | 12.3 | 12 | 12.6 | 13 | 12.8 | 13.3 | / | / | / | / | / | /
WK47/2021 | HI700B-0 | Ammonia Low Range Reagent B | L56 CTK | B135 | colorimetric | HI700 | / | Ammonia Low Range | N-Ammonia | mg/L | 0 | 3 | 2 | 0.05 | & | 5% | 100% | MR09 | 0 | 0 | 0.05 | 0.6 | 0.52 | 0.68 | 2.4 | 2.23 | 2.57 | / | / | / | / | / | / | / | / | / | 12.3 | 12 | 12.6 | 13 | 12.8 | 13.3 | / | / | / | / | / | /
WK47/2021 | HI701-0 | Free Chlorine Reagent | L57 Powder | R025 | colorimetric | HI701 | / | Free Chlorine | Chlorine | mg/L | 0 | 2.5 | 2 | 0.03 | & | 3% | 100% | MR67 | 0 | 0 | 0.03 | 0.5 | 0.46 | 0.54 | 2 | 1.91 | 2.09 | / | / | / | / | / | / | / | / | / | 6 | 5.7 | 6.3 | / | / | / | / | / | / | 120 | 110 | 130
WK47/2021 | HI7014-0 | Free Chlorine Reagent | L57 Powder | R025 | colorimetric | HI7014 | / | Free Chlorine | Chlorine | mg/L | 0 | 2.5 | 2 | 0.03 | & | 3% | 100% | MR67 | 0 | 0 | 0.03 | 0.5 | 0.46 | 0.54 | 2 | 1.91 | 2.09 | / | / | / | / | / | / | / | / | / | 6 | 5.7 | 6.3 | / | / | / | / | / | / | 120 | 110 | 130
WK47/2021 | HI702-0 | Copper High Range Reagent | L57 Powder | R026 | colorimetric | HI702 | / | Copper High Range (Cu) | Copper | mg/L | 0 | 5 | 2 | 0.05 | & | 5% | 100% | MR44 | 0 | 0 | 0.05 | 1 | 0.9 | 1.1 | 4 | 3.75 | 4.25 | / | / | / | / | / | / | / | / | / | 6.4 | 6.2 | 6.7 | / | / | / | / | / | / | 270 | 250 | 290
WK47/2021 | HI7024-0 | Copper High Range Reagent | L57 Powder | R026 | colorimetric | HI7024 | / | Copper High Range (Cu) | Copper | mg/L | 0 | 5 | 2 | 0.05 | & | 5% | 100% | MR44 | 0 | 0 | 0.05 | 1 | 0.9 | 1.1 | 4 | 3.75 | 4.25 | / | / | / | / | / | / | / | / | / | 6.4 | 6.2 | 6.7 | / | / | / | / | / | / | 270 | 250 | 290
WK47/2021 | HI70404-0 | Potassium Iodide | L57 Powder | R024 | CTK | / | / | / | / |  |  |  | / | 0 | / | 0% | / |  | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | 7 | 6.8 | 7.2 | / | / | / | / | / | / | 270 | 260 | 280
WK47/2021 | HI70452-0 | Free and Total Chlorine DPD Reagent | L57 Powder | R020 | CTK | / | / | Free Chlorine | Chlorine |  |  |  | 1 | 0 | / | 0% | 100% | MR67 | 0 | 0 | 0 | 0.5 | 0.5 | 0.5 | 2 | 2 | 2 | / | / | / | / | / | / | / | / | / | 6 | 5.7 | 6.3 | / | / | / | / | / | / | 660 | 620 | 700
WK47/2021 | HI705A-0 | Silica Reagent A | L56 CTK | B145 | colorimetric | HI705 | / | Silica (SiO2) | Silica Oxide | mg/L | 0 | 2 | 2 | 0.03 | & | 5% | 100% | MR45 | 0 | 0 | 0.03 | 0.4 | 0.35 | 0.45 | 1.6 | 1.49 | 1.71 | / | / | / | / | / | / | / | / | / | 0.6 | 0.4 | 0.8 | 2.4 | 2.3 | 2.6 | 1.4 | 1.2 | 1.6 | / | / | /
WK47/2021 | HI705B-0 | Silica Reagent B | L57 Powder | R014 | colorimetric | HI705 | / | Silica (SiO2) | Silica Oxide | mg/L | 0 | 2 | 2 | 0.03 | & | 5% | 100% | MR45 | 0 | 0 | 0.03 | 0.4 | 0.35 | 0.45 | 1.6 | 1.49 | 1.71 | / | / | / | / | / | / | / | / | / | 0.6 | 0.4 | 0.8 | 2.4 | 2.3 | 2.6 | 1.4 | 1.2 | 1.6 | 60 | 50 | 70
WK47/2021 | HI705C-0 | Silica Reagent C | L57 Powder | R030 | colorimetric | HI705 | / | Silica (SiO2) | Silica Oxide | mg/L | 0 | 2 | 2 | 0.03 | & | 5% | 100% | MR45 | 0 | 0 | 0.03 | 0.4 | 0.35 | 0.45 | 1.6 | 1.49 | 1.71 | / | / | / | / | / | / | / | / | / | 0.6 | 0.4 | 0.8 | 2.4 | 2.3 | 2.6 | 1.4 | 1.2 | 1.6 | 85 | 75 | 95
WK47/2021 | HI706AS | Phosphorus High Range Reagent A | L56 CTK | B146 | colorimetric | HI706 | / | Phosphorus High Range (P) | Phosphorus | mg/L | 0 | 15 | 1 | 0.3 | & | 5% | 100% | MR24 | 0 | 0 | 0.3 | 3.1 | 2.6 | 3.6 | 12 | 11.1 | 12.9 | / | / | / | / | / | / | / | / | / | 0.6 | 0.4 | 0.8 | 1.4 | 1.2 | 1.6 | / | / | / | / | / | /
WK47/2021 | HI706B-0 | Phosphorus High Range Reagent B | L57 Powder | R030 | colorimetric | HI706 | / | Phosphorus High Range (P) | Phosphorus | mg/L | 0 | 15 | 1 | 0.3 | & | 5% | 100% | MR24 | 0 | 0 | 0.3 | 3.1 | 2.6 | 3.6 | 12 | 11.1 | 12.9 | / | / | / | / | / | / | / | / | / | 0.6 | 0.4 | 0.8 | 1.4 | 1.2 | 1.6 | / | / | / | 170 | 160 | 180
WK47/2021 | HI707-0 | Nitrite Low Range Reagent | L57 Powder | R031 | colorimetric | HI707 | / | Nitrite Low Range(NO2) | N-Nitrite | µg/L | 0 | 600 | 0 | 20 | & | 5% | 100% | MR13 | 0 | 0 | 20 | 120 | 94 | 145 | 480 | 436 | 524 | / | / | / | / | / | / | / | / | / | 3 | 2.9 | 3.1 | / | / | / | / | / | / | 285 | 275 | 295
WK47/2021 | HI708-0 | Nitrite High Range Reagent | L57 Powder | R032 | colorimetric | HI708 | / | Nitrite High Range(NO2) | Nitrite | mg/L | 0 | 150 | 0 | 3 | & | 5% | 100% | MR13 | 0 | 0 | 3 | 30 | 26 | 35 | 120 | 111 | 129 | / | / | / | / | / | / | / | / | / | 1.5 | 1.3 | 1.7 | / | / | / | / | / | / | 325 | 300 | 350
WK47/2021 | HI709A-0 | Manganese High Range Reagent A | L57 Powder | R033 | colorimetric | HI709 | / | Manganese High Range (Mn) | Manganese | mg/L | 0 | 20 | 1 | 0.2 | & | 5% | 100% | MR37 | 0 | 0 | 0.2 | 4 | 3.6 | 4.4 | 16 | 15 | 17 | / | / | / | / | / | / | / | / | / | 5.7 | 5 | 6.6 | 5.7 | 5 | 6.6 | / | / | / | 150 | 140 | 160
WK47/2021 | HI709B-0 | Manganese High Range Reagent B | L57 Powder | R034 | colorimetric | HI709 | / | Manganese High Range (Mn) | Manganese | mg/L | 0 | 20 | 1 | 0.2 | & | 5% | 100% | MR37 | 0 | 0 | 0.2 | 4 | 3.6 | 4.4 | 16 | 15 | 17 | / | / | / | / | / | / | / | / | / | 5.7 | 5 | 6.6 | 5.7 | 5 | 6.6 | / | / | / | 105 | 95 | 115
WK47/2021 | HI711-0 | Total Chlorine Reagent | L57 Powder | R035 | colorimetric | HI711 | / | Total Chlorine | Chlorine | mg/L | 0 | 3.5 | 2 | 0.03 | & | 3% | 100% | MR67 | 0 | 0 | 0.03 | 0.5 | 0.46 | 0.54 | 2.8 | 2.69 | 2.91 | / | / | / | / | / | / | / | / | / | 6 | 5.7 | 6.3 | / | / | / | / | / | / | 150 | 135 | 165
WK47/2021 | HI713-0 | Phosphate Reagent | L57 Powder | R029 | colorimetric | HI713 | / | Phosphate Low Range(PO4) | Phosphate | mg/L | 0 | 2.5 | 2 | 0.04 | & | 4% | 100% | MR24 | 0 | 0 | 0.04 | 0.5 | 0.44 | 0.56 | 2 | 1.88 | 2.12 | / | / | / | / | / | / | / | / | / | 1.5 | 1.2 | 1.6 | / | / | / | / | / | / | 180 | 165 | 195
WK47/2021 | HI7134-0 | Phosphate Low Range Reagent | L57 Powder | R029 | colorimetric | HI7134 | / | Phosphate Low Range(PO4) | Phosphate | mg/L | 0 | 2.5 | 2 | 0.04 | & | 4% | 100% | MR24 | 0 | 0 | 0.04 | 0.5 | 0.44 | 0.56 | 2 | 1.88 | 2.12 | / | / | / | / | / | / | / | / | / | 1.5 | 1.2 | 1.6 | / | / | / | / | / | / | 180 | 165 | 195
WK47/2021 | HI715A-0 | Ammonia Mid-range Reagent A | L56 CTK | B134 | colorimetric | HI715 | / | Ammonia Medium Range | N-Ammonia | mg/L | 0 | 9.99 | 2 | 0.05 | & | 5% | 100% | MR09 | 0 | 0 | 0.05 | 2.4 | 2.23 | 2.52 | 8 | 7.55 | 8.45 | / | / | / | / | / | / | / | / | / | 12.3 | 12 | 12.6 | 13 | 12.8 | 13.3 | / | / | / | / | / | /
WK47/2021 | HI715B-0 | Nessler Reagent | L56 CTK | B135 | colorimetric | HI715 | / | Ammonia Medium Range | N-Ammonia | mg/L | 0 | 9.99 | 2 | 0.05 | & | 5% | 100% | MR09 | 0 | 0 | 0.05 | 2.4 | 2.23 | 2.52 | 8 | 7.55 | 8.45 | / | / | / | / | / | / | / | / | / | 12.3 | 12 | 12.6 | 13 | 12.8 | 13.3 | / | / | / | / | / | /
WK47/2021 | HI716-0 | Bromine Reagent | L57 Powder | R035 | colorimetric | HI716 | / | Bromine | Bromine | mg/L | 0 | 8 | 2 | 0.08 | & | 5% | 100% | MR67 | 0 | 0 | 0.08 | 1.6 | 1.44 | 1.76 | 6.4 | 6 | 6.8 | / | / | / | / | / | / | / | / | / | 6.1 | 5.7 | 6.3 | / | / | / | / | / | / | 150 | 135 | 165
WK47/2021 | HI717AS | Phosphate HR Reagent A | L56 CTK | B146 | colorimetric | HI717 | / | Phosphate High Range(PO4) | Phosphate | mg/L | 0 | 30 | 1 | 1 | & | 5% | 100% | MR24 | 0 | 0 | 1 | 6 | 4.7 | 7.2 | 24 | 21.8 | 26.2 | / | / | / | / | / | / | / | / | / | 0.6 | 0.4 | 0.8 | 1.4 | 1.2 | 1.6 | / | / | / | / | / | /
WK47/2021 | HI717B-0 | Phosphate HR Reagent B | L57 Powder | R030 | colorimetric | HI717 | / | Phosphate High Range(PO4) | Phosphate | mg/L | 0 | 30 | 1 | 1 | & | 5% | 100% | MR24 | 0 | 0 | 1 | 6 | 4.7 | 7.2 | 24 | 21.8 | 26.2 | / | / | / | / | / | / | / | / | / | 0.6 | 0.4 | 0.8 | 1.4 | 1.2 | 1.6 | / | / | / | 170 | 160 | 180
WK47/2021 | HI718-0 | Iodine Reagent | L57 Powder | R035 | colorimetric | HI718 | / | Iodine | Iodine | mg/L | 0 | 12.5 | 1 | 0.1 | & | 5% | 100% | MR67 | 0 | 0 | 0.1 | 2.5 | 2.3 | 2.7 | 10 | 9.4 | 10.6 | / | / | / | / | / | / | / | / | / | 6.1 | 5.7 | 6.3 | / | / | / | / | / | / | 150 | 135 | 165
WK47/2021 | HI719A-0 | Magnesium Hardness Reagent A | L56 CTK | B149 | colorimetric | HI719 | / | Magnesium and Total Hardness Reagents | CaCO3 | mg/L | 0 | 2 | 2 | 0.2 | & | 5% | 100% | MR28 | 0 | 0 | 0.41 | 0.8 | 0.18 |  | 1.6 | -0.2 | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI719B-0 | Magnesium Hardness Reagent B | L56 CTK | B150 | colorimetric | HI719 | / | Magnesium and Total Hardness Reagents | CaCO3 | mg/L | 0 | 2 | 2 | 0.2 | & | 5% | 100% | MR28 | 0 | 0 | 0.41 | 0.8 | 0.18 |  | 1.6 | -0.2 | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI719C-0 | Magnesium Hardness Reagent C | L56 CTK | B151 | colorimetric | HI719 | / | Magnesium and Total Hardness Reagents | CaCO3 | mg/L | 0 | 2 | 2 | 0.2 | & | 5% | 100% | MR28 | 0 | 0 | 0.41 | 0.8 | 0.18 |  | 1.6 | -0.2 | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI719D-0 | Magnesium Hardness Reagent D | L56 CTK | B152 | colorimetric | HI719 | / | Magnesium and Total Hardness Reagents | CaCO3 | mg/L | 0 | 2 | 2 | 0.2 | & | 5% | 100% | MR28 | 0 | 0 | 0.41 | 0.8 | 0.18 |  | 1.6 | -0.2 | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI720A-0 | Calcium Hardness Reagent A | L56 CTK | B149 | colorimetric | HI720 | / | Calcium Hardness Reagents | CaCO3 | mg/L | 0 | 2.7 | 2 | 0.2 | & | 5% | 100% | MR20 | 0 | 0.75 | 1.25 | 1 | 1.7 |  | 2 | -0.2 | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI720B-0 | Calcium Hardness Reagent B | L56 CTK | B150 | colorimetric | HI720 | / | Calcium Hardness Reagents | CaCO3 | mg/L | 0 | 2.7 | 2 | 0.2 | & | 5% | 100% | MR20 | 0 | 0.75 | 1.25 | 1 | 1.7 |  | 2 | -0.2 | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI720C-0 | Calcium Hardness Reagent C | L56 CTK | B152 | colorimetric | HI720 | / | Calcium Hardness Reagents | CaCO3 | mg/L | 0 | 2.7 | 2 | 0.2 | & | 5% | 100% | MR20 | 0 | 0.75 | 1.25 | 1 | 1.7 |  | 2 | -0.2 | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI721-0 | Iron High Range Reagent | L57 Powder | R041 | colorimetric | HI721 | / | Iron High Range (Fe) | Iron | mg/L | 0 | 5 | 2 | 0.04 | & | 2% | 100% | MR33 | 0 | 0 | 0.04 | 1 | 0.94 | 1.05 | 4 | 3.88 | 4.12 | / | / | / | / | / | / | / | / | / | 5.2 | 4 | 6.5 | / | / | / | / | / | / | 175 | 160 | 190
WK47/2021 | HI7214-0 | Iron High Range Reagent for Swimming Pools | L57 Powder | R041 | colorimetric | HI7214 | / | Iron High Range (Fe) | Iron | mg/L | 0 | 5 | 2 | 0.04 | & | 2% | 100% | MR33 | 0 | 0 | 0.04 | 1 | 0.94 | 1.05 | 4 | 3.88 | 4.12 | / | / | / | / | / | / | / | / | / | 5.2 | 4 | 6.5 | / | / | / | / | / | / | 175 | 160 | 190
WK47/2021 | HI723-0 | Chromium VI High Range Reagent | L57 Powder | R043 | colorimetric | HI723 | / | Chromium VI High Range (Cr) | Chromium VI | µg/L | 0 | 999 | 0 | 5 | & | 4% | 100% | MR43 | 0 | 0 | 5 | 200 | 187 | 212 | 800 | 763 | 837 | / | / | / | / | / | / | / | / | / | 1.3 | 1.1 | 1.6 | / | / | / | / | / | / | 150 | 140 | 160
WK47/2021 | HI726-0 | Nickel High Range Reagent | L57 Powder | R044 | colorimetric | HI726 | / | Nickel High Range(Ni) | Nickel | g/L | 0 | 7 | 2 | 0.1 | & | 5% | 100% | MR26 | 0 | 0 | 0.1 | 1.4 | 1.23 | 1.57 | 5.6 | 5.22 | 5.98 | / | / | / | / | / | / | / | / | / | 10.9 | 10.7 | 11.1 | / | / | / | / | / | / | 750 | 720 | 780
WK47/2021 | HI729S | Fluoride Low Range Reagent | L56 CTK | B153 | colorimetric | HI729 | / | Fluoride Low Range | Fluoride | mg/L | 0 | 2 | 2 | 0.1 | & | 5% | 100% | MR18 | 0 | 0 | 0.1 | 0.4 | 0.29 | 0.53 | 1 | 0.87 | 0.52 | 1.6 |  |  | / | / | / | / | / | / | 0 | 0 | 0.2 | / | / | / | / | / | / | / | / | /
WK47/2021 | HI733A-0 | Ammonia High Range Reagent A | L56 CTK | B135 | colorimetric | HI733 | / | Ammonia High Range | N-Ammonia | mg/L | 0 | 99.9 | 1 | 1 | & | 5% | 100% | MR09 | 0 | 0 | 1 | 20 | 18 | 22 | 80 | 75 | 85 | / | / | / | / | / | / | / | / | / | 12.3 | 12 | 12.6 | 13 | 12.8 | 13.3 | / | / | / | / | / | /
WK47/2021 | HI733B-0 | Ammonia High Range Reagent B | L56 CTK | B154 | colorimetric | HI733 | / | Ammonia High Range | N-Ammonia | mg/L | 0 | 99.9 | 1 | 1 | & | 5% | 100% | MR09 | 0 | 0 | 1 | 20 | 18 | 22 | 80 | 75 | 85 | / | / | / | / | / | / | / | / | / | 12.3 | 12 | 12.6 | 13 | 12.8 | 13.3 | / | / | / | / | / | /
WK47/2021 | HI735A-0 | Total Hardness Low Range Reagent A | L56 CTK | B206 | colorimetric | HI735 | / | Total Hardness ISM as CaCO3 | Ca-Carbonate | mg/L | 0 | 350 | 0 | 3 | & | 3% | 100% | MR20 | 0 | 0 | 3 | 15 | 11 | 19 | 100 | 94 | 106 | 270 | 330 | / | 330 | / | / | / | / | / | / | / | / | 11.3 | 11.1 | 11.5 | 11.3 | 11.1 | 11.5 | / | / | /
WK47/2021 | HI735B-0 | Total Hardness Low Range Reagent B | L56 CTK | B159 | colorimetric | HI735 | / | Total Hardness ISM as CaCO3 | Ca-Carbonate | mg/L | 0 | 350 | 0 | 3 | & | 3% | 100% | MR20 | 0 | 0 | 3 | 15 | 12 | 18 | 100 | 94 | 106 | 270 | 330 | / | 330 | / | / | / | / | / | / | / | / | 11.3 | 11.1 | 11.5 | 11.3 | 11.1 | 11.5 | / | / | /
WK47/2021 | HI735C-0 | Total Hardness Low Range Reagent C | L57 Powder | R050 | colorimetric | HI735 | / | Hardness ISM MR as CaCO3 | Ca-Carbonate | mg/L | 0 | 350 | 0 | 3 | & | 3% | 100% | MR20 | 0 | 0 | 3 | 15 | 12 | 18 | 100 | 94 | 106 | 270 | 330 | / | 330 | / | / | / | / | / | / | / | / | 11.3 | 11.1 | 11.5 | 11.3 | 11.1 | 11.5 | 150 | 140 | 160
WK47/2021 | HI735IND-0 | Total Hardness Low Range Indicator Reagent | L56 CTK | B197 | colorimetric | HI735 | / | Total Hardness ISM as CaCO3 | Ca-Carbonate | mg/L | 0 | 350 | 0 | 3 | & | 3% | 100% | MR20 | 0 | 0 | 3 | 15 | 12 | 18 | 100 | 94 | 106 | 270 | 330 | / | 330 | / | / | / | / | / | / | / | / | 11.3 | 11.1 | 11.5 | 11.3 | 11.1 | 11.5 | / | / | /
WK47/2021 | HI736-0 | Phosphorus Ultra Low Range Reagent | L57 Powder | R071 | colorimetric | HI736 | / | Phosphorus(P) | Phosphorus | µg/L | 0 | 200 | 0 | 5 | & | 5% | 100% | MR24 | 0 | 0 | 5 | 40 | 33 | 47 | 160 | 147 | 173 | / | / | / | / | / | / | / | / | / | 1.5 | 1.2 | 1.6 | / | / | / | / | / | / | 180 | 165 | 195
WK47/2021 | HI739AS | Fluoride HR Reagent A | L56 CTK | B167 | colorimetric | HI739 | / | Fluoride High Range | Fluoride | mg/L | 0 | 20 | 1 | 0.5 | & | 5% | 100% | MR18/ | 0 | 0 | 0.5 | 4 | 3.3 | 4.7 | 16 | 14.8 | 17.4 | / | / | / | / | / | / | / | / | / | 0 | 0 | 0 | 6.6 | 6.4 | 6.8 | / | / | / | / | / | /
WK47/2021 | HI739BS | Fluoride HR Reagent B | L56 CTK | B168 | colorimetric | HI739 | / | Fluoride High Range | Fluoride | mg/L | 0 | 20 | 1 | 0.5 | & | 5% | 100% | MR18/ | 0 | 0 | 0.5 | 4 | 3.3 | 4.7 | 16 | 14.8 | 17.4 | / | / | / | / | / | / | / | / | / | 0 | 0 | 0 | 6.6 | 6.4 | 6.8 | / | / | / | / | / | /
WK47/2021 | HI746-0 | Iron Low Range Reagent | L57 Powder | R054 | colorimetric | HI746 | / | Iron Low Range (Fe) | Iron | µg/L | 0 | 999 | 0 | 20 | & | 5% | 100% | MR15 | 0 | 0 | 20 | 200 | 170 | 230 | 800 | 740 | 860 | / | / | / | / | / | / | / | / | / | 2.9 | 2.7 | 3.1 | / | / | / | / | / | / | 260 | 250 | 270
WK47/2021 | HI747-0 | Copper Low Range Reagent | L57 Powder | R060 | colorimetric | HI747 | / | Copper Low Range | Copper | µg/L | 0 | 999 | 0 | 10 | & | 5% | 100% | MR44 | 0 | 0 | 10 | 200 | 180 | 220 | 800 | 750 | 850 | / | / | / | / | / | / | / | / | / | 6.5 | 6.2 | 6.8 | / | / | / | / | / | / | 270 | 260 | 280
WK47/2021 | HI749-0 | Chromium VI Low Range Reagent | L57 Powder | R043 | colorimetric | HI749 | / | Chromium VI Low Range (Cr) | Chromium VI | µg/L | 0 | 300 | 0 | 3 | & | 5% | 100% | MR43 | 0 | 0 | 3 | 60 | 54 | 65 | 240 | 225 | 255 | / | / | / | / | / | / | / | / | / | 1.3 | 1.1 | 1.6 | / | / | / | / | / | / | 120 | 110 | 130
WK47/2021 | HI753A-0 | Chloride Reagent A | L56 CTK | B175 | colorimetric | HI753 | / | Chloride | Cl | mg/L | 0 | 20 | 1 | 0.5 | & | 6% | 100% | MR39 | 0 | 0 | 0.5 | 4 | 3.3 | 4.7 | 16 | 14.5 | 17.5 | / | / | / | / | / | / | / | / | / |  |  |  | 1.5 | 1 | 2 | / | / | /
WK47/2021 | HI753B-0 | Chloride Reagent A | L56 CTK | B176 | colorimetric | HI753 | / | Chloride | Cl | mg/L | 0 | 20 | 1 | 0.5 | & | 6% | 100% | MR39 | 0 | 0 | 0.5 | 4 | 3.3 | 4.7 | 16 | 14.5 | 17.5 | / | / | / | / | / | / | / | / | / |  |  |  | 1.5 | 1 | 2 | / | / | /
WK47/2021 | HI755S | Alkalinity Reagent | L56 CTK | B104 | colorimetric | HI755 | / | Alkalinity | Ca-Carbonate | mg/L | 0 | 300 | 0 | 5 | & | 5% | 100% | MR05 | 0 | 0 | 5 | 50 | 52 | 67 | 150 | 224 | 258 | 200 | / | / | / | / | / | / | / |  | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI7581 | Calcium Reagent A | L56 CTK | B105 | colorimetric | HI758 | / | Calcium(Ca2+) | Calcium | mg/L | 200 | 600 | 0 | 6 | & | 5% | 100% | MR28+MR60 | 300 | 279 | 321 | 400 | 374 | 426 | 500 | 469 | 531 | / | / | / | / | / | / | / | / | / | 9.9 | 9.8 | 10 | 9.9 | 9.8 | 10 | / | / | / | / | / | /
WK47/2021 | HI7582 | Calcium Reagent B | L57 Powder | R013 | colorimetric | HI758 | / | Calcium(Ca2+) | Calcium | mg/L | 200 | 600 | 0 | 6 | & | 5% | 100% | MR28+MR60 | 300 | 279 | 321 | 400 | 374 | 426 | 500 | 469 | 531 | / | / | / | / | / | / | / | / | / | 9.9 | 9.8 | 10 | 9.9 | 9.8 | 10 | / | / | / | 250 | 235 | 265
WK47/2021 | HI761-0 | Total Chlorine Ultra Low Range Reagent | L57 Powder | R061 | colorimetric | HI761 | / | Total Chlorine Ultra Low Range | Chlorine | µg/L | 0 | 500 | 0 | 5 | & | 5% | 100% | MR67 | 0 | 0 | 5 | 100 | 90 | 110 | 400 | 375 | 425 | / | / | / | / | / | / | / | / | / | 6 | 5.7 | 6.3 | / | / | / | / | / | / | 150 | 135 | 165
WK47/2021 | HI762-0 | Free Chlorine ULR Reagent | L57 Powder | R062 | colorimetric | HI762 | / | Free Chlorine Ultra Low Range | Chlorine | µg/L | 0 | 500 | 0 | 20 | & | 4% | 100% | MR67 | 0 | 0 | 20 | 100 | 76 | 124 | 400 | 364 | 436 | / | / | / | / | / | / | / | / | / | 6 | 5.7 | 6.3 | / | / | / | / | / | / | 120 | 110 | 130
WK02/2022 | HI764-0 | Nitrite Ultra Low Range Reagent | L57 Powder | R031 | colorimetric | HI801 | HI764 | Nitrite Low Range(NO2) | N-Nitrite | µg/L | 0 | 200 | 0 | 8 | & | 4% | 100% | MR13 | 0 | 0 | 10 | 40 | 28 | 51 | 160 | 144 | 176 | / | / | / | / | / | / | / | / | / | 3 | 2.9 | 3.1 | / | / | / | / | / | / | 285 | 275 | 295
WK47/2021 | HI767-0 | Marine Nitrite Low Range Reagent | L57 Powder | R031 | colorimetric | HI767 | / | Marine Nitrite Low Range | N-Nitrite | µg/L | 0 | 999 | 0 | 10 | & | 4% | 100% | MR13 | 0 | 0 | 10 | 200 | 182 | 218 | 800 | 758 | 842 | / | / | / | / | / | / | / | / | / | 3 | 2.9 | 3.1 | / | / | / | / | / | / | 285 | 275 | 295
WK47/2021 | HI770A-0 | Silica High Range Reagent A | L57 Powder | R064 | colorimetric | HI770 | / | Silica High Range (SiO2) | Silica Oxide | mg/L | 0 | 200 | 0 | 2 | & | 5% | 100% | MR82 | 0 | 0 | 2 | 40 | 36 | 44 | 160 | 150 | 170 | / | / | / | / | / | / | / | / | / | 8.2 | 8.1 | 8.4 | 1 | 0.8 | 1.2 | 2.2 | 2 | 2.5 | 120 | 110 | 130
WK47/2021 | HI770B-0 | Silica High Range Reagent B | L57 Powder | R065 | colorimetric | HI770 | / | Silica High Range (SiO2) | Silica Oxide | mg/L | 0 | 200 | 0 | 2 | & | 5% | 100% | MR82 | 0 | 0 | 2 | 40 | 36 | 44 | 159 | 149 | 169 | / | / | / | / | / | / | / | / | / | 8.2 | 8.1 | 8.4 | 1 | 0.8 | 1.2 | 2.2 | 2 | 2.5 | 330 | 315 | 345
WK47/2021 | HI770C-0 | Silica High Range Reagent C | L57 Powder | R014 | colorimetric | HI770 | / | Silica High Range (SiO2) | Silica Oxide | mg/L | 0 | 200 | 0 | 2 | & | 5% | 100% | MR82 | 0 | 0 | 2 | 40 | 36 | 44 | 159 | 149 | 169 | / | / | / | / | / | / | / | / | / | 8.2 | 8.1 | 8.4 | 1 | 0.8 | 1.2 | 2.2 | 2 | 2.5 | 100 | 90 | 110
WK47/2021 | HI771A-0 | Chlorine Ultra High Range Reagent A | L57 Powder | R021 | colorimetric | HI771 | / | Chlorine Ultra High Range | Chlorine | mg/L | 0 | 500 | 0 | 3 | & | 5% | 100% | MR67 | 0 | 0 | 3 | 100 | 92 | 107 | 400 | 377 | 423 | / | / | / | / | / | / | / | / | / | 7 | 6.8 | 7.2 | 1.5 | 1.4 | 1.7 | / | / | / | 130 | 115 | 145
WK47/2021 | HI771B-0 | Chlorine Ultra High Range Reagent B | L57 Powder | R063 | colorimetric | HI771 | / | Chlorine Ultra High Range | Chlorine | mg/L | 0 | 500 | 0 | 3 | & | 5% | 100% | MR67 | 0 | 0 | 3 | 100 | 92 | 107 | 400 | 377 | 423 | / | / | / | / | / | / | / | / | / | 7 | 6.8 | 7.2 | 1.5 | 1.4 | 1.7 | / | / | / | 100 | 90 | 110
WK47/2021 | HI772S | Alkalinity Reagent | L56 CTK | B104 | colorimetric | HI772 | / | Alkalinity | dKH | dKH | 0 | 20 | 1 | 0.3 | & | 5% | 100% | MR05 | 0 | 3.8 | 4.8 | 4 | 9.2 | 10.8 | 10 | 14.7 | 16.9 | 16 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI774-0 | Phosphate Ultra Low Range Reagent | L57 Powder | R071 | colorimetric | HI774 | / | Phosphate Ultra Low Range | Phosphate | mg/L | 0 | 0.9 | 2 | 0.02 | & | 5% | 100% | MR24 | 0 | 0 | 0.02 | 0.18 | 0.15 | 0.21 | 0.72 | 0.66 | 0.78 | / | / | / | / | / | / | / | / | / | 1.5 | 1.2 | 1.6 | / | / | / | / | / | / | 180 | 165 | 195
WK47/2021 | HI7754S | Alkalinity Reagent | L56 CTK | B106 | colorimetric | HI775 | / | Alkalinity | Alkalinity | mg/L | 0 | 500 | 0 | 5 | & | 5% | 100% | MR05 | 100 | 90 | 110 | 200 | 185 | 215 | 400 | 375 | 425 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI775S | Alkalinity Reagent | L56 CTK | B106 | colorimetric | HI775 | / | Alkalinity | Alkalinity | mg/L | 0 | 500 | 0 | 5 | & | 5% | 100% | MR05 | 100 | 90 | 110 | 200 | 185 | 215 | 400 | 375 | 425 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | /
WK02/2022 | HI779-0 | pH Swimming Pool Reagent | L56 CTK | B147 | colorimetric | HI779 |  | pH | pH | pH | 6.5 | 8.5 | 1 | 0.2 | & | 0% | 100% | CM309/CM029 | 6.8 | 6.6 | 7 | 8 | 7.8 | 8.2 |  |  |  | / | / | / | / | / | / | / | / | /
WK02/2022 | HI780-0 | pH Marine Reagent | L56 CTK | B147 | colorimetric | HI780 |  | pH | pH | pH | 6.5 | 8.5 | 1 | 0.2 | & | 0% | 100% | CM023/CM022 | 6.8 | 6.6 | 7 | 8 | 7.8 | 8.2 |  |  |  | / | / | / | / | / | / | / | / | /
WK12/2022 | HI784A-0 | Marine Ammonia Reagent A | L56 CTK | B209 | colorimetric | HI83399 | HI784 | Ammonia | NH3 | mg/L | 0.00 | 2.50 | 2 | 0.05 | & | 5% | 50% | MR09 | 0 | 0 | 0.02 | 0.4 | 0.36 | 0.44 | 1.25 | 1.19 | 1.31 | 2.1 | 2.02 | 2.18 | / | / | / | / | / | / | / | / | / | 10 | 9.9 | 10.1 |  |  |  | / | / | /
WK12/2022 | HI784B-0 | Marine Ammonia Reagent B | L57 Powder | R81 | colorimetric | HI83399 | HI784 | Ammonia | NH3 | mg/L | 0.00 | 2.50 | 2 | 0.05 | & | 5% | 50% | MR09 | 0 | 0 | 0.02 | 0.4 | 0.36 | 0.44 | 1.25 | 1.19 | 1.31 | 2.1 | 2.02 | 2.18 | / | / | / | / | / | / | / | / | / | 10 | 9.9 | 10.1 |  |  |  | / | / | /
WK12/2022 | HI784C-0 | Marine Ammonia Reagent C | L56 CTK | B210 | colorimetric | HI83399 | HI784 | Ammonia | NH3 | mg/L | 0.00 | 2.50 | 2 | 0.05 | & | 5% | 50% | MR09 | 0 | 0 | 0.02 | 0.4 | 0.36 | 0.44 | 1.25 | 1.19 | 1.31 | 2.1 | 2.02 | 2.18 | / | / | / | / | / | / | / | / | / | 10 | 9.9 | 10.1 |  |  |  | / | / | /
WK12/2022 | HI783A-0 | Marine Magnesium Reagent A | L56 CTK | B207 | colorimetric | HI83399 | HI783 | Magnesium | Mg | mg/L | 1000 | 1800 | 0 | 5 | & | 5% | 100% | MR100 | 1100 | 1060 | 1140 | 1400 | 1350 | 1450 | 1700 | 1640 | 1760 | / | / | / | / | / | / | / | / | / | / | / | / | 10 | 9.9 | 10.1 |  |  |  | / | / | /
WK12/2022 | HI783B-0 | Marine Magnesium Reagent B | L56 CTK | B208 | colorimetric | HI83399 | HI783 | Magnesium | Mg | mg/L | 1000 | 1800 | 0 | 5 | & | 5% | 100% | MR100 | 1100 | 1060 | 1140 | 1400 | 1350 | 1450 | 1700 | 1640 | 1760 | / | / | / | / | / | / | / | / | / | / | / | / | 10 | 9.9 | 10.1 |  |  |  | / | / | /
WK47/2021 | HI781A-0 | Marine Nitrate LR Reagent A | L56 CTK | B205 | colorimetric | HI781 | / | Nitrate | Nitrate (NO3) | mg/L | 0 | 5 | 2 | 0.25 | & | 2% | 100% | MR40 | 0 | 0 | 0.25 | 1.02 | 0.75 | 1.2 | 2.52 | 2.22 | 2.82 | 4.03 | 3.72 | 4.34 | / | / | / | / | / | / | 11 | 11 | 11.1 |  |  |  | 2 | 1.9 | 2.1 | / | / | /
WK47/2021 | HI781B-0 | Marine Nitrate LR Reagent B | L57 Powder | R78 | colorimetric | HI781 | / | Nitrate | Nitrate (NO3) | mg/L | 0 | 5 | 2 | 0.25 | & | 2% | 100% | MR40 | 0 | 0 | 0.25 | 1.02 | 0.75 | 1.2 | 2.52 | 2.22 | 2.82 | 4.03 | 3.72 | 4.34 | / | / | / | / | / | / | / | / | / | / | / | / | 2 | 1.9 | 2.1 | / | / | /
WK47/2021 | HI781C-0 | Marine Nitrate LR Reagent C | L57 Powder | R79 | colorimetric | HI781 | / | Nitrate | Nitrate (NO3) | mg/L | 0 | 5 | 2 | 0.25 | & | 2% | 100% | MR40 | 0 | 0 | 0.25 | 1.02 | 0.75 | 1.2 | 2.52 | 2.22 | 2.82 | 4.03 | 3.72 | 4.34 | / | / | / | / | / | / | / | / | / | / | / | / | 2 | 1.9 | 2.1 | / | / | /
WK47/2021 | HI782-0 | Marine Nitrate High Range Reagent | L57 Powder | R080 | colorimetric | HI782 | / | Marine Nitrate High Range | Nitrate (NO3) | mg/L | 0 | 75 | 1 | 2 | & | 5% | 100% | MR40 | 0 | 0 | 2 | 15 | 12.3 | 18.6 | 35 | 31.3 | 38.8 | 60 | 54.6 | 65.4 | [12 empty] | 1.8 | 1.7 | 1.9 | 206 | 196 | 216
WK03/2022 | HI83730A-0 | Peroxide Reagent A (test vial) | L59 COD | D013 | colorimetric | HI83730 | / | ISM peroxide value in edible oils | Oxygen/kg | meq | 0 | 25 | 1 | 0.5 | or | 0% | 100% | MR75 | 0 | 0 | 0.5 | 11 | 10.5 | 20 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | 5 | / | /
WK03/2022 | HI83730B-0 | Peroxide Reagent B | L57 Powder | R021 | colorimetric | HI83730 | / | Peroxide | Oxygen/kg | meq O2 /kg | 0 | 25 | 1 | 0.5 | or | 0% | 100% | MR75 | 0 | 0 | 0.5 | 11 | 10.5 | 11.5 | / | / | / | / | / | / | / | / | / | / | / | / | 7 | 6.8 | 7.2 | 1.5 | 1.4 | 1.7 | / | / | / | 100 | 90 | 110
WK47/2021 | HI83741A-0 | Iron Reagent A | L57 Powder | R023 | colorimetric | HI83741 | / | Iron | Iron | mg/L | 0 | 15 | 1 | 0.2 | & | 5% | 100% | / | 0 | 0 | 0.2 | 0.4 | 0.2 | 0.6 | 1.3 | 1 | 1.6 | / | / | / | / | / | / | / | / | / | 2.7 | 2.6 | 2.9 | / | / | / | / | / | / | 150 | 135 | 165
WK47/2021 | HI83741B-0 | Iron Reagent B | L57 Powder | R022 | colorimetric | HI83741 | / | Iron | Iron | mg/L | 0 | 15 | 1 | 0.2 | & | 5% | 100% | / | 0 | 0 | 0.2 | 0.4 | 0.2 | 0.6 | 1.3 | 1 | 1.6 | / | / | / | / | / | / | / | / | / | 2.7 | 2.6 | 2.9 | / | / | / | / | / | / | 152 | 137 | 167
WK47/2021 | HI83742-0 | Wine Solvent | L56 CTK | B111 | CTK | / | / |  | / |  |  |  | 1 | 0 | / | 0% | 60% |  | / | / | / |  | 0 | / |  | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI83746A-0 | Fehling Solution A | L59 COD | D014 | colorimetric | HI83746 | / | Reducing sugar in wine | Reducing sugar in wine | g/L | 0 | 50 | 2 | 0.5 | & | 5% | 100% | MR58+MR59 | 0 | 0 | 0.5 | 10 | 9 | 10.9 | 25 | 23.25 | 26.75 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI84100-54-0 | Sulphur Dioxide Stabilizer | L57 Powder | R024 | CTK | / | / | / | / |  |  |  | / | 0 | / | 0% | / |  | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | 7 | 6.8 | 7.2 | / | / | / | / | / | / | 270 | 260 | 280
WK47/2021 | HI84500-62-0 | Sulfur Dioxide Stabilizer Reagent | L57 Powder | R024 | CTK | / | / | / | / |  |  |  | / | 0 | / | 0% | / |  | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | 7 | 6.8 | 7.2 | / | / | / | / | / | / | 270 | 260 | 280
WK47/2021 | HI93700A-0 | Ammonia Low Range Reagent A | L56 CTK | B134 | colorimetric | HI801 | HI83399 | Ammonia Low Range | N-Ammonia | mg/L | 0 | 3 | 2 | 0.04 | & | 4% | 100% | MR09 | 0 | 0 | 0.04 | 0.6 | 0.54 | 0.66 | 2.4 | 2.26 | 2.54 | / | / | / | / | / | / | / | / | / | 12.3 | 12 | 12.6 | 13 | 12.8 | 13.3 | / | / | / | / | / | /
WK47/2021 | HI93700B-0 | Ammonia Low Range Reagent B | L56 CTK | B135 | colorimetric | HI801 | HI83399 | Ammonia Low Range | N-Ammonia | mg/L | 0 | 3 | 2 | 0.04 | & | 4% | 100% | MR09 | 0 | 0 | 0.04 | 0.6 | 0.54 | 0.66 | 2.4 | 2.26 | 2.54 | / | / | / | / | / | / | / | / | / | 12.3 | 12 | 12.6 | 13 | 12.8 | 13.3 | / | / | / | / | / | /
WK47/2021 | HI93701-0 | Free Chlorine Reagent | L57 Powder | R025 | colorimetric | HI801 | HI83399 | Free Chlorine | Chlorine | mg/L | 0 | 5 | 2 | 0.03 | & | 3% | 100% | MR67 | 0 | 0 | 0.03 | 0.5 | 0,46 (0.47) | 0,54 - 0.53 | 2 | 1,91 - 1,94 | 2,09 - 2,06 | 8 | / | / | / | / | / | / | / | / | 6 | 5.7 | 6.3 | / | / | / | / | / | / | 120 | 110 | 130 | X
WK47/2021 | HI93701A-F | Free Chlorine Reagent A | L56 CTK | B136 | colorimetric | HI801 | HI83399 | Free Chlorine (liquid) | Chlorine | mg/L | 0 | 5 | 2 | 0.03 | & | 3% | 100% | MR67 | 0 | 0 | 0.03 | 0.5 | 0.46 | 0.54 | 2 | 1.91 | 2.09 | / | / | / | / | / | / | / | / | / | 0.2 | 0 | 0.4 | 6 | 5.7 | 6.3 | / | / | / | / | / | /
WK47/2021 | HI93701A-T | Total Chlorine Reagent A | L56 CTK | B136 | colorimetric | HI801 | HI83399 | Total Chlorine (liquid) | Chlorine | mg/L | 0 | 5 | 2 | 0.03 | & | 3% | 100% | MR67 | 0 | 0 | 0.03 | 0.5 | 0.46 | 0.54 | 2.8 |  |  |  | / | / | / | / | / | / | / | / | 0.2 | 0 | 0.4 | 6 | 5.7 | 6.3 | / | / | / | / | / | /
WK47/2021 | HI93701B-F | Free Chlorine Reagent B | L56 CTK | B137 | colorimetric | HI801 | HI83399 | Free Chlorine (liquid) | Chlorine | mg/L | 0 | 5 | 2 | 0.03 | & | 3% | 100% | MR67 | 0 | 0 | 0.03 | 0.5 | 0.46 | 0.54 | 2 | 1.91 | 2.09 | / | / | / | / | / | / | / | / | / | 0.2 | 0 | 0.4 | 6 | 5.7 | 6.3 | / | / | / | / | / | /
WK47/2021 | HI93701B-T | Total Chlorine Reagent B | L56 CTK | B137 | colorimetric | HI801 | HI83399 | Total Chlorine (liquid) | Chlorine | mg/L | 0 | 5 | 2 | 0.03 | & | 3% | 100% | MR67 | 0 | 0 | 0.03 | 0.5 | 0.46 | 0.54 | 2.8 |  |  |  | / | / | / | / | / | / | / | / | 0.2 | 0 | 0.4 | 6 | 5.7 | 6.3 | / | / | / | / | / | /
WK47/2021 | HI93701C-T | Total Chlorine Reagent C | L56 CTK | B138 | colorimetric | HI801 | HI83399 | Total Chlorine (liquid) | Chlorine | mg/L | 0 | 5 | 2 | 0.03 | & | 3% | 100% | MR67 | 0 | 0 | 0.03 | 0.5 | 0.46 | 0.54 | 2.8 |  |  | 8 | / | / | / | / | / | / | / | / | 0.2 | 0 | 0.4 | 6 | 5.7 | 6.3 | / | / | / | / | / | /
WK47/2021 | HI93702-0 | Copper High Range Reagent | L57 Powder | R026 | colorimetric | HI801 | HI83399 | Copper High Range (Cu) | Copper | mg/L | 0 | 5 | 2 | 0.02 | & | 4% | 100% | MR44 | 0 | 0 | 0.02 | 1 | 0.94 | 1.04 | 4 | 3.82 | 4.18 | / | / | / | / | / | / | / | / | / | 6.4 | 6.2 | 6.7 | / | / | / | / | / | / | 270 | 250 | 290
WK47/2021 | HI93702T-0 | Total Copper Reagent | L57 Powder | R027 | colorimetric | HI801 | HI83399 | Copper High Range (Cu) | Copper | mg/L | 0 | 5 | 2 | 0.02 | & | 4% | 100% | / | 0 | 0 | 0.02 | 1 | 0.94 | 1.06 | 4 | 3.82 | 4.18 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | 105 | 95 | 115
WK47/2021 | HI93703-52-0 | Glycine Powder | L57 Powder | R028 | colorimetric | HI801 | HI83399 | Ozone | Ozone | mg/L | 0 | 2 | 2 | 0.02 | & | 3% | 100% | MR67 | 0 | 0 | 0.02 | 2 | 1.92 | 2.08 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | 125 | 115 | 135
WK47/2021 | HI93704-0 | Hydrazine Reagent | L56 CTK | B144 | colorimetric | HI801 | HI83399 | Hydrazine | Hydrazine | µg/L | 0 | 400 | 0 | 3 | & | 3% | 100% | MR56 | 0 | 0 | 3 | 80 | 94 | 106 | 320 | 307 | 333 | / | / | / | / | / | / | / | / | / | [6 empty] | / | / | / | / | / | /
WK47/2021 | HI93705A-0 | Silica Reagent A | L56 CTK | B145 | colorimetric | HI801 | HI83399 | Silica (SiO2) | Silica Oxide | mg/L | 0 | 2 | 2 | 0.03 | & | 3% | 100% | MR45 | 0 | 0 | 0.03 | 0.4 | 0.36 | 0.44 | 1.6 | 1.52 | 1.68 | / | / | / | / | / | / | / | / | / | 0.6 | 0.4 | 0.8 | 2.4 | 2.3 | 2.6 | 1.4 | 1.2 | 1.6 | / | / | /
WK47/2021 | HI93705B-0 | Silica Reagent B | L57 Powder | R014 | colorimetric | HI801 | HI83399 | Silica (SiO2) | Silica Oxide | mg/L | 0 | 2 | 2 | 0.03 | & | 3% | 100% | MR45 | 0 | 0 | 0.03 | 0.4 | 0.36 | 0.44 | 1.6 | 1.52 | 1.68 | / | / | / | / | / | / | / | / | / | 0.6 | 0.4 | 0.8 | 2.4 | 2.3 | 2.6 | 1.4 | 1.2 | 1.6 | 60 | 50 | 70
WK47/2021 | HI93705C-0 | Silica Reagent C | L57 Powder | R030 | colorimetric | HI801 | HI83399 | Silica (SiO2) | Silica Oxide | mg/L | 0 | 2 | 2 | 0.03 | & | 3% | 100% | MR45 | 0 | 0 | 0.03 | 0.4 | 0.36 | 0.44 | 1.6 | 1.52 | 1.68 | / | / | / | / | / | / | / | / | / | 0.6 | 0.4 | 0.8 | 2.4 | 2.3 | 2.6 | 1.4 | 1.2 | 1.6 | 85 | 75 | 95
WK47/2021 | HI93706A-0 | Phosphorus Reagent A | L56 CTK | B146 | colorimetric | HI801 | HI83399 | Phosphate Hight Range (PO4) | Phosphorus | mg/L | 0 | 15 | 1 | 0.3 | & | 4% | 100% | MR24 | 0 | 0 | 0.3 | 3.1 | 2.7 | 3.5 | 12 | 11.2 | 12.8 | / | / | / | / | / | / | / | / | / | 0.6 | 0.4 | 0.8 | 1.4 | 1.2 | 1.6
WK47/2021 | HI93706B-0 | Phosphorus Reagent B | L57 Powder | R030 | colorimetric | HI801 | HI83399 | Phosphate High Range(PO4) | Phosphorus | mg/L | 0 | 15 | 1 | 0.3 | & | 4% | 100% | MR24 | 0 | 0 | 0.3 | 3.1 | 2.7 | 3.5 | 12 | 11.2 | 12.8 | / | / | / | / | / | / | / | / | / | 0.6 | 0.4 | 0.8 | 1.4 | 1.2 | 1.6 | / | / | / | 170 | 160 | 180
WK47/2021 | HI93707-0 | Nitrite Low Range Reagent | L57 Powder | R031 | colorimetric | HI801 | HI83399 | Nitrite Low Range(NO2) | N-Nitrite | µg/L | 0 | 600 | 0 | 20 | & | 4% | 100% | MR13 | 0 | 0 | 20 | 120 | 95 | 145 | 480 | 441 | 519 | / | / | / | / | / | / | / | / | / | 3 | 2.9 | 3.1 | / | / | / | / | / | / | 285 | 275 | 295
WK47/2021 | HI93708-0 | Nitrite High Range Reagent | L57 Powder | R032 | colorimetric | HI801 | HI83399 | Nitrite High Range(NO2) | N-Nitrite | mg/L | 0 | 150 | 0 | 4 | & | 4% | 100% | MR13 | 0 | 0 | 4 | 30 | 25 | 35 | 120 | 111 | 129 | / | / | / | / | / | / | / | / | / | 1.5 | 1.3 | 1.7 | / | / | / | / | / | / | 325 | 300 | 350
WK47/2021 | HI93709A-0 | Manganese HR Reagent A | L57 Powder | R033 | colorimetric | HI801 | HI83399 | Manganese High Range (Mn) | Manganese | mg/L | 0 | 20 | 1 | 0.2 | & | 3% | 100% | MR37 | 0 | 0 | 0.2 | 4 | 3.7 | 4.3 | 16 | 15.3 | 16.7 | / | / | / | / | / | / | / | / | / | 5.7 | 5 | 6.6 | 5.7 | 5 | 6.6 | / | / | / | 150 | 140 | 160
WK47/2021 | HI93709B-0 | Manganese HR Reagent B | L57 Powder | R034 | colorimetric | HI801 | HI83399 | Manganese High Range (Mn) | Manganese | mg/L | 0 | 20 | 1 | 0.2 | & | 3% | 100% | MR37 | 0 | 0 | 0.2 | 4 | 3.7 | 4.3 | 16 | 15.3 | 16.7 | / | / | / | / | / | / | / | / | / | 5.7 | 5 | 6.6 | 5.7 | 5 | 6.6 | / | / | / | 105 | 95 | 115
WK02/2022 | HI93710-0 | pH Reagent | L56 CTK | B147 | colorimetric | HI83399 | HI801 | pH | pH | pH | 6.5 | 8.5 | 1 | 0.1 | & | 0% | 100% | HI70456(0.1 N NAOH/HI70463(0.1N HCL) | 6.6 | 6.6 | 7 | 7.5 | 7.3 | 7.7 | 8.3 | 8.1 | 8.5 | / | / | / | / | / | / | / | / | /
WK47/2021 | HI93711-0 | Total Chlorine Reagent | L57 Powder | R035 | colorimetric | HI801 | HI83399 | Total Chlorine | Chlorine | mg/L | 0 | 5 | 2 | 0.03 | & | 3% | 100% | MR67 | 0 | 0 | 0.03 | 0.5 | 0.46 | 0.54 | 2.8 | 2.69 | 2.91 | / | / | / | / | / | / | / | / | / | 6.1 | 5.7 | 6.3 | / | / | / | / | / | / | 150 | 135 | 165
WK47/2021 | HI93712A-0 | Aluminum Reagent A | L57 Powder | R023 | colorimetric | HI801 | HI83399 | Aluminum(Al) | Aluminum | mg/L | 0 | 1 | 2 | 0.04 | & | 4% | 100% | MR41 | 0 | 0 | 0.04 | 0.2 | 0.15 | 0.24 | 0.8 | 0.73 | 0.87 | / | / | / | / | / | / | / | / | / | 2.6 | 2.4 | 2.8 | 3.3 | 3.2 | 3.5 | 3.8 | 3.7 | 4 | 200 | 190 | 210
WK47/2021 | HI93712B-0 | Aluminum Reagent B | L57 Powder | R037 | colorimetric | HI801 | HI83399 | Aluminum(Al) | Aluminum | mg/L | 0 | 1 | 2 | 0.04 | & | 4% | 100% | MR41 | 0 | 0 | 0.04 | 0.2 | 0.15 | 0.24 | 0.8 | 0.73 | 0.87 | / | / | / | / | / | / | / | / | / | 2.6 | 2.4 | 2.8 | 3.3 | 3.2 | 3.5 | 3.8 | 3.7 | 4 | 500 | 470 | 530
WK47/2021 | HI93712C-0 | Aluminum Reagent C | L57 Powder | R038 | colorimetric | HI801 | HI83399 | Aluminum(Al) | Aluminum | mg/L | 0 | 1 | 2 | 0.04 | & | 4% | 100% | MR41 | 0 | 0 | 0.04 | 0.2 | 0.15 | 0.24 | 0.8 | 0.73 | 0.87 | / | / | / | / | / | / | / | / | / | 2.6 | 2.4 | 2.8 | 3.3 | 3.2 | 3.5 | 3.8 | 3.7 | 4 | 90 | 80 | 100
WK47/2021 | HI93713-0 | Phosphate Low Range Reagent | L57 Powder | R029 | colorimetric | HI801 | HI83399 | Phosphate Low Range(PO4) | Phosphate | mg/L | 0 | 2.5 | 2 | 0.04 | & | 4% | 100% | MR24 | 0 | 0 | 0.04 | 0.5 | 0.44 | 0.56 | 2 | 1.88 | 2.12 | / | / | / | / | / | / | / | / | / | 1.5 | 1.2 | 1.6 | / | / | / | / | / | / | 180 | 165 | 195
WK47/2021 | HI93714A-0 | Cyanide Reagent A | L57 Powder | R036 | colorimetric | HI801 | HI83399 | Cyanide(CN) | Cyanide | mg/L | 0 | 0.2 | 3 | 0.005 | & | 3% | 100% | CM327 | 0 | 0 | 0.005 | 0.04 | 0.034 | 0.046 | 0.16 | 0.15 | 0.17 | / | / | / | / | / | / | / | / | / | 6.8 | 6.7 | 7 | / | / | / | 6.4 | 6.2 | 6.8 | 17005 | 16805 | 17205
WK47/2021 | HI93714B-0 | Cyanide Reagent B | L57 Powder | R039 | colorimetric | HI801 | HI83399 | Cyanide(CN) | Cyanide | mg/L | 0 | 0.2 | 3 | 0.005 | & | 3% | 100% | CM327 | 0 | 0 | 0.005 | 0.04 | 0.034 | 0.046 | 0.16 | 0.15 | 0.17 | / | / | / | / | / | / | / | / | / | 6.8 | 6.7 | 7 | / | / | / | 6.4 | 6.2 | 6.8 | 200 | 190 | 210
WK47/2021 | HI93714C-0 | Cyanide Reagent C | L57 Powder | R040 | colorimetric | HI801 | HI83399 | Cyanide(CN) | Cyanide | mg/L | 0 | 0.2 | 3 | 0.005 | & | 3% | 100% | CM327 | 0 | 0 | 0.005 | 0.04 | 0.034 | 0.046 | 0.16 | 0.15 | 0.17 | / | / | / | / | / | / | / | / | / | 6.8 | 6.7 | 7 | / | / | / | 6.4 | 6.2 | 6.8 | 95 | 85 | 105
WK47/2021 | HI93715A-0 | Ammonia Medium Range Reagent A | L56 CTK | B134 | colorimetric | HI801 | HI83399 | Ammonia Medium Range | N-Ammonia | mg/L | 0 | 10 | 2 | 0.05 | & | 5% | 100% | MR09 | 0 | 0 | 0.05 | 2.4 | 2.23 | 2.52 | 8 | 7.55 | 8.45 | / | / | / | / | / | / | / | / | / | 12.3 | 12 | 12.6 | 13 | 12.8 | 13.3 | / | / | / | / | / | /
WK47/2021 | HI93715B-0 | Ammonia Medium Range Reagent B | L56 CTK | B135 | colorimetric | HI801 | HI83399 | Ammonia Medium Range | N-Ammonia | mg/L | 0 | 10 | 2 | 0.05 | & | 5% | 100% | MR09 | 0 | 0 | 0.05 | 2.4 | 2.23 | 2.52 | 8 | 7.55 | 8.45 | / | / | / | / | / | / | / | / | / | 12.3 | 12 | 12.6 | 13 | 12.8 | 13.3 | / | / | / | / | / | /
WK47/2021 | HI93716-0 | Bromine Reagent | L57 Powder | R035 | colorimetric | HI801 | HI83399 | Bromine | Bromine | mg/L | 0 | 8 | 2 | 0.08 | & | 3% | 100% | MR67 | 0 | 0 | 0.08 | 1.6 | 1.47 | 1.72 | 6.4 | 6.13 | 6.67 | / | / | / | / | / | / | / | / | / | 6.1 | 5.7 | 6.3 | / | / | / | / | / | / | 150 | 135 | 165 | X
WK47/2021 | HI93717A-0 | Phosphate High Range Reagent A | L56 CTK | B146 | colorimetric | HI801 | HI83399 | Phosphate High Range(PO43-) | Phosphate | mg/L | 0 | 30 | 1 | 1 | & | 4% | 100% | MR24 | 0 | 0 | 1 | 6 | 4.8 | 7.1 | 24 | 22 | 26 | / | / | / | / | / | / | / | / | / | 0.6 | 0.4 | 0.8 | 1.4 | 1.2 | 1.6 | / | / | / | / | / | /
WK47/2021 | HI93717B-0 | Phosphate High Range Reagent B | L57 Powder | R030 | colorimetric | HI801 | HI83399 | Phosphate High Range(PO43-) | Phosphate | mg/L | 0 | 30 | 1 | 1 | & | 4% | 100% | MR24 | 0 | 0 | 1 | 6 | 4.8 | 7.2 | 24 | 22 | 26 | / | / | / | / | / | / | / | / | / | 0.6 | 0.4 | 0.8 | 1.4 | 1.2 | 1.6 | / | / | / | 170 | 160 | 180
WK47/2021 | HI93718-0 | Iodine Reagent | L57 Powder | R035 | colorimetric | HI801 | HI83399 | Iodine | Iodine | mg/L | 0 | 12.5 | 1 | 0.1 | & | 5% | 100% | MR67 | 0 | 0 | 0.1 | 2.5 | 2.3 | 2.7 | 10 | 9.4 | 10.6 | / | / | / | / | / | / | / | / | / | 6.1 | 5.7 | 6.3 | / | / | / | / | / | / | 150 | 135 | 165
WK47/2021 | HI93719A-0 | Magnesium Hardness Reagent A | L56 CTK | B149 | colorimetric | HI83399 | HI801 | Magnesium and Total Hardness Reagents | CaCO3 | mg/L | 0 | 2 | 2 | 0.11 | & | 5% | 100% | MR28 | 0 | 0.65 | 0.95 | 0.8 | 1.41 | 1.79 | 1.6 |  | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI93719B-0 | Magnesium Hardness Reagent B | L56 CTK | B150 | colorimetric | HI83399 | HI801 | Magnesium and Total Hardness Reagents | CaCO3 | mg/L | 0 | 2 | 2 | 0.11 | & | 5% | 100% | MR28 | 0 | 0.65 | 0.95 | 0.8 | 1.41 | 1.79 | 1.6 |  | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI93719C-0 | Magnesium Hardness Reagent C | L56 CTK | B151 | colorimetric | HI83399 | HI801 | Magnesium and Total Hardness Reagents | CaCO3 | mg/L | 0 | 2 | 2 | 0.11 | & | 5% | 100% | MR28 | 0 | 0.65 | 0.95 | 0.8 | 1.41 | 1.79 | 1.6 |  | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI93719D-0 | Magnesium Hardness Reagent D | L56 CTK | B152 | colorimetric | HI83399 | HI801 | Magnesium and Total Hardness Reagents | CaCO3 | mg/L | 0 | 2 | 2 | 0.11 | & | 5% | 100% | MR28 | 0 | 0.65 | 0.95 | 0.8 | 1.41 | 1.79 | 1.6 |  | / | / | / | / | / | / | / | / | / | /
WK02/2022 | HI93720A-0 | Calcium Hardness Reagent A | L56 CTK | B149 | colorimetric | HI83399 | HI801 | Calcium Hardness Reagents | CaCO3 | mg/L | 0 | 2.7 | 2 | 0.08 | & | 4% | 100% | MR20 | 0 | 0.84 | 1.16 | 1 | 1.79 |  | 2 | -0.11 | / | / | / | / | / | / | / | / | / | /
WK02/2022 | HI93720B-0 | Calcium Hardness Reagent B | L56 CTK | B150 | colorimetric | HI83399 | HI801 | Calcium Hardness Reagents | CaCO3 | mg/L | 0 | 2.7 | 2 | 0.08 | & | 4% | 100% | MR20 | 0 | 0.84 | 1.16 | 1 | 1.79 |  | 2 | -0.11 | / | / | / | / | / | / | / | / | / | /
WK02/2022 | HI93720C-0 | Calcium Hardness Reagent C | L56 CTK | B152 | colorimetric | HI83399 | HI801 | Calcium Hardness Reagents | CaCO3 | mg/L | 0 | 2.7 | 2 | 0.08 | & | 4% | 100% | MR20 | 0 | 0.84 | 1.16 | 1 | 1.79 |  | 2 | -0.11 | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI93721-0 | Iron High Range Reagent | L57 Powder | R041 | colorimetric | HI801 | HI83399 | Iron High Range (Fe) | Iron | mg/L | 0 | 5 | 2 | 0.04 | & | 2% | 100% | MR33 | 0 | 0 | 0.04 | 1 | 0.94 | 1.05 | 4 | 3.88 | 4.12 | / | / | / | / | / | / | / | / | / | 5.2 | 4 | 6.5 | / | / | / | / | / | / | 175 | 160 | 190
WK47/2021 | HI93722-0 | Cyanuric Acid Reagent | L57 Powder | R042 | colorimetric | HI801 | HI83399 | Cyanuric Acid (CYA) | CYA | mg/L | 0 | 80 | 0 | 1 | & | 15% | 100% | MR34 | 0 | 0 | 1 | 16 | 13 | 19 | 64 | 53 | 75 | / | / | / | / | / | / | / | / | / | 6 | 5 | 7 | / | / | / | / | / | / | 215 | 200 | 230 | X
WK47/2021 | HI93723-0 | Chromium VI High Range Reagent | L57 Powder | R043 | colorimetric | HI801 | HI83399 | Chromium VI High Range (Cr) | Chromium VI | µg/L | 0 | 1000 | 0 | 5 | & | 4% | 100% | MR43 | 0 | 0 | 5 | 200 | 187 | 212 | 800 | 763 | 837 | / | / | / | / | / | / | / | / | / | 1.3 | 1.1 | 1.6 | / | / | / | / | / | / | 150 | 140 | 160
WK47/2021 | HI93726-0 | Nickel High Range Reagent | L57 Powder | R044 | colorimetric | HI801 | HI83399 | Nickel High Range(Ni) | Nickel | g/L | 0 | 7 | 2 | 0.07 | & | 4% | 100% | MR26 | 0 | 0 | 0.07 | 1.4 | 1.27 | 1.53 | 5.6 | 5.31 | 5.89 | / | / | / | / | / | / | / | / | / | 10.9 | 10.7 | 11.1 | / | / | / | / | / | / | 750 | 720 | 780
WK47/2021 | HI93728-0 | Nitrate Reagent | L57 Powder | R045 | colorimetric | HI801 | HI83399 | Nitrate (NO3-N) | N-Nitrate | mg/L | 0 | 30 | 1 | 0.5 | & | 10% | 100% | MR40 | 0 | 0 | 0.5 | 6 | 4.9 | 7.1 | 24 | 21.1 | 26.9 | / | / | / | / | / | / | / | / | / | 2.7 | 2.5 | 3 | / | / | / | / | / | / | 220 | 205 | 235
WK47/2021 | HI93729-0 | Fluoride Low Range Reagent | L56 CTK | B153 | colorimetric | HI801 | HI83399 | Fluoride Low Range | Fluoride | mg/L | 0 | 2 | 2 | 0.03 | & | 3% | 100% | MR18 | 0 | 0 | 0.03 | 0.4 | 0.37 | 0.45 | 1 | 0.96 | 0.44 | 1.6 | / | / | / | / | / | / | / | / | 0 | 0 | 0.2 | / | / | / | / | / | / | / | / | /
WK47/2021 | HI93730A-0 | Molybdenum Reagent A | L57 Powder | R046 | colorimetric | HI801 | HI83399 | Molybdenum (Mo) | Molybdenum | mg/L | 0 | 40 | 1 | 0.3 | & | 5% | 100% | MR19 | 0 | 0 | 0.3 | 8 | 7.3 | 8.7 | 32 | 30.1 | 33.9 | / | / | / | / | / | / | / | / | / | 4.1 | 3.9 | 4.3 | 4.1 | 3.9 | 4.3 | 4.2 | 4 | 4.4 | 265 | 250 | 280
WK47/2021 | HI93730B-0 | Molybdenum Reagent B | L57 Powder | R047 | colorimetric | HI801 | HI83399 | Molybdenum (Mo) | Molybdenum | mg/L | 0 | 40 | 1 | 0.3 | & | 5% | 100% | MR19 | 0 | 0 | 0.3 | 8 | 7.3 | 8.7 | 32 | 30.1 | 33.9 | / | / | / | / | / | / | / | / | / | 4.1 | 3.9 | 4.3 | 4.1 | 3.9 | 4.3 | 4.2 | 4 | 4.4 | 100 | 90 | 110
WK47/2021 | HI93730C-0 | Molybdenum Reagent C | L57 Powder | R048 | colorimetric | HI801 | HI83399 | Molybdenum (Mo) | Molybdenum | mg/L | 0 | 40 | 1 | 0.3 | & | 5% | 100% | MR19 | 0 | 0 | 0.3 | 8 | 7.3 | 8.7 | 32 | 30.1 | 33.9 | / | / | / | / | / | / | / | / | / | 4.1 | 3.9 | 4.3 | 4.1 | 3.9 | 4.3 | 4.2 | 4 | 4.4 | 140 | 130 | 150
WK47/2021 | HI93731A-0 | Zinc Reagent A | L57 Powder | R049 | colorimetric | HI801 | HI83399 | Zinc (Zn) | Zinc | mg/L | 0 | 3 | 2 | 0.03 | & | 3% | 100% | MR62 | 0 | 0 | 0.03 | 0.6 | 0.55 | 0.65 | 1.8 | 1.72 | 1.88 | 2.9 | 2.78 | 3.02 | 3.5 | 3 | 3 | / | / | / | 8.8 | 8.7 | 9 | / | / | / | / | / | / | 400 | 370 | 430
WK47/2021 | HI93731B-0 | Zinc Reagent B | L56 CTK | B155 | colorimetric | HI801 | HI83399 | Zinc (Zn) | Zinc | mg/L | 0 | 3 | 2 | 0.03 | & | 3% | 100% | MR62 | 0 | 0 | 0.03 | 0.6 | 0.55 | 0.65 | 2.4 | 1.72 | 1.88 | 2.9 | 3.5 | / | 3.5 | / | / | / | / | / | 8.8 | 8.7 | 9 | / | / | / | / | / | / | / | / | /
WK47/2021 | HI93733A-0 | Ammonia High Range Reagent A | L56 CTK | B135 | colorimetric | HI801 | HI83399 | Ammonia High Range | N-Ammonia | mg/L | 0 | 100 | 1 | 0.5 | & | 5% | 100% | MR09 | 0 | 0 | 0.5 | 20 | 18.5 | 21.5 | 80 | 75.5 | 84.5 | / | / | / | / | / | / | / | / | / | 12.3 | 12 | 12.6 | 13 | 12.8 | 13.3 | / | / | / | / | / | /
WK47/2021 | HI93735A-HR | Hardness High Range Reagent A | L56 CTK | B156 | colorimetric | HI801 | HI83399 | Hardness ISM HR as CaCO3 | Ca-Carbonate | mg/L | 0 | 750 | 0 | 10 | & | 2% | 100% | MR20 | 0 | 431 | 469 | 500 | 480 | 520 | 600 | 578 | 622 | 750 | / | / |  | / | / | / | / | / | / | / | / | 11.3 | 11.1 | 11.5 | 11.3 | 11.1 | 11.5 | / | / | /
WK47/2021 | HI93735A-LR | Hardness Low Range Reagent A | L56 CTK | B206 | colorimetric | HI801 | HI83399 | Hardness ISM LR as CaCO3 | Ca-Carbonate | mg/L | 0 | 250 | 0 | 5 | & | 4% | 100% | MR20 | 0 | 43 | 57 | 100 | 91 | 109 | 200 | 187 | 213 | 250 | 250 | / |  | / | / | / | / | / | / | / | / | 11.3 | 11.1 | 11.5 | 11.3 | 11.1 | 11.5 | / | / | /
WK47/2021 | HI93735A-MR | Hardness Medium Range Reagent A | L56 CTK | B158 | colorimetric | HI801 | HI83399 | Hardness ISM MR as CaCO3 | Ca-Carbonate | mg/L | 0 | 500 | 0 | 7 | & | 3% | 100% | MR20 | 0 | 236 | 265 | 250 | 333 | 368 | 350 | 430 | 471 | 450 | / | / | 500 | [5 empty] | / | / | / | 11.3 | 11.1 | 11.5 | 11.3 | 11.1 | 11.5 | / | / | /
WK47/2021 | HI93735B-0 | Hardness Buffer Reagent B | L56 CTK | B159 | colorimetric | HI801 | HI83399 | Hardness ISM HR-LR-MR as CaCO3 | Ca-Carbonate | mg/L | 0 | 750 | 0 | 7 | & | 3% | 100% | MR20 | 0 | 90 | 110 | 100 | 381 | 419 | 400 | 624 | 677 | 650 | / | / | / | / | / | / | / | / | / | / | / | 11.3 | 11.1 | 11.5 | 11.3 | 11.1 | 11.5 | / | / | /
WK47/2021 | HI93735C-0 | Fixing Reagent | L57 Powder | R050 | colorimetric | HI801 | HI83399 | Hardness ISM MR as CaCO3 | Ca-Carbonate | mg/L | 0 | 750 | 0 | 7 | & | 3% | 100% | MR20 | 0 | 90 | 110 | 200 | 381 | 419 | 450 | 624 | 677 | 600 | / | / | / | / | / | / | / | / | / | / | / | 11.3 | 11.1 | 11.5 | 11.3 | 11.1 | 11.5 | 150 | 140 | 160
WK47/2021 | HI93735IND-HR | Hardness Indicator Reagent HR | L56 CTK | B197 | colorimetric | HI801 | HI83399 | Hardness ISM HR as CaCO3 | Ca-Carbonate | mg/L | 0 | 750 | 0 | 10 | & | 2% | 100% | MR20 | 0 | 480 | 520 | 500 | 578 | 622 | 600 | 725 | 775 | 750 | / | / | / | / | / | / | / | / | / | / | / | 11.3 | 11.1 | 11.5 | 11.3 | 11.1 | 11.5 | / | / | /
WK47/2021 | HI93735IND-LR | Hardness Indicator Reagent LR | L56 CTK | B197 | colorimetric | HI801 | HI83399 | Hardness ISM MR as CaCO3 | Ca-Carbonate | mg/L | 0 | 250 | 0 | 5 | & | 4% | 100% | MR20 | 0 | 0 | 5 | 100 | 91 | 109 | 200 | 187 | 213 | 250 | [8 empty] | / | / | / | 11.3 | 11.1 | 11.5 | 11.3 | 11.1 | 11.5 | / | / | /
WK47/2021 | HI93735IND-MR | Hardness Indicator Reagent MR | L56 CTK | B197 | colorimetric | HI801 | HI83399 | Hardness ISM LR as CaCO3 | Ca-Carbonate | mg/L | 0 | 500 | 0 | 7 | & | 3% | 100% | MR20 | 0 | 236 | 265 | 250 | 333 | 368 | 350 | 430 | 471 | 450 | / | / | 500 | [5 empty] | / | / | / | 11.3 | 11.1 | 11.5 | 11.3 | 11.1 | 11.5 | / | / | /
WK47/2021 | HI93737A-0 | Silver Reagent A | L56 CTK | B160 | colorimetric | HI801 | HI83399 | Silver | Ag | µg/L | 0 | 1000 | 0 | 20 | & | 5% | 100% | CM102 | 0 | 0 | 20 | 432 | 390 | 474 | 863 | 800 | 926 | / | / | / | / | / | / | / | / | / | 9.4 | 9.3 | 9.5 | 9.5 | 9.3 | 9.7 | 9.5 | 9.3 | 9.7
WK47/2021 | HI93737B-0 | Silver Reagent B | L56 CTK | B161 | colorimetric | HI801 | HI83399 | Silver | Ag | µg/L | 0 | 1000 | 0 | 20 | & | 5% | 100% | CM102 | 0 | 0 | 20 | 432 | 390 | 474 | 863 | 800 | 926 | / | / | / | / | / | / | / | / | / | 9.4 | 9.3 | 9.5 | 9.5 | 9.3 | 9.7 | 9.5 | 9.3 | 9.7
WK47/2021 | HI93737C-0 | Silver Reagent C | L56 CTK | B162 | colorimetric | HI801 | HI83399 | Silver | Ag | µg/L | 0 | 1000 | 0 | 20 | & | 5% | 100% | CM102 | 0 | 0 | 20 | 432 | 390 | 474 | 863 | 800 | 926 | / | / | / | / | / | / | / | / | / | 9.4 | 9.3 | 9.5 | 9.5 | 9.3 | 9.7 | 9.5 | 9.3 | 9.7
WK47/2021 | HI93737D-0 | Silver Reagent D | L56 CTK | B163 | colorimetric | HI801 | HI83399 | Silver | Ag | µg/L | 0 | 1000 | 0 | 20 | & | 5% | 100% | CM102 | 0 | 0 | 20 | 432 | 390 | 474 | 863 | 800 | 926 | / | / | / | / | / | / | / | / | / | 9.4 | 9.3 | 9.5 | 9.5 | 9.3 | 9.7 | 9.5 | 9.3 | 9.7
WK47/2021 | HI93738A-0 | Chlorine Dioxide Reagent A | L56 CTK | B164 | colorimetric | HI801 | HI83399 | Chlorine Dioxide(ClO2) | Chlorine dioxide | mg/L | 0 | 2 | 2 | 0.1 | & | 5% | 100% | CM357 | 0 | 0 | 0.1 | 0.4 | 0.28 | 0.52 | 1.6 | 1.42 | 1.78 | / | / | / | / | / | / | / | / | / | / | / | / | 10 | 9.9 | 10 | / | / | / | / | / | /
WK47/2021 | HI93738B-0 | Chlorine Dioxide Reagent B | L57 Powder | R051 | colorimetric | HI801 | HI83399 | Chlorine Dioxide(ClO2) | Chlorine dioxide | mg/L | 0 | 2 | 2 | 0.1 | & | 5% | 100% | CM357 | 0 | 0 | 0.1 | 0.4 | 0.28 | 0.52 | 1.6 | 1.42 | 1.78 | / | / | / | / | / | / | / | / | / | / | / | / | 10 | 9.9 | 10 | / | / | / | 155 | 145 | 165
WK47/2021 | HI93738C-0 | Chlorine Dioxide Reagent C | L56 CTK | B165 | colorimetric | HI801 | HI83399 | Chlorine Dioxide(ClO2) | Chlorine dioxide | mg/L | 0 | 2 | 2 | 0.1 | & | 5% | 100% | CM357 | 0 | 0 | 0.1 | 0.4 | 0.28 | 0.52 | 1.6 | 1.42 | 1.78 | / | / | / | / | / | / | / | / | / | / | / | / | 10 | 9.9 | 10 | / | / | / | / | / | /
WK47/2021 | HI93738D-0 | Chlorine Dioxide Reagent D | L56 CTK | B166 | colorimetric | HI801 | HI83399 | Chlorine Dioxide(ClO2) | Chlorine dioxide | mg/L | 0 | 2 | 2 | 0.1 | & | 5% | 100% | CM357 | 0 | 0 | 0.1 | 0.4 | 0.28 | 0.52 | 1.6 | 1.42 | 1.78 | / | / | / | / | / | / | / | / | / | / | / | / | 10 | 9.9 | 10 | / | / | / | / | / | /
WK47/2021 | HI93739A-0 | Fluoride High Range Reagent A | L56 CTK | B167 | colorimetric | HI801 | HI83399 | Fluoride High Range | Fluoride | mg/L | 0 | 20 | 1 | 0.5 | & | 3% | 100% | MR18 | 0 | 0 | 0.5 | 4 | 3.4 | 4.6 | 16 | 15.1 | 17.1 | / | / | / | / | / | / | / | / | / | 0 | 0 | 0 | 6.6 | 6.4 | 6.8 | / | / | / | / | / | /
WK47/2021 | HI93739B-0 | Fluoride High Range Reagent B | L56 CTK | B168 | colorimetric | HI801 | HI83399 | Fluoride High Range | Fluoride | mg/L | 0 | 20 | 1 | 0.5 | & | 3% | 100% | MR18 | 0 | 0 | 0.5 | 4 | 3.4 | 4.6 | 16 | 15.1 | 17.1 | / | / | / | / | / | / | / | / | / | 0 | 0 | 0 | 6.6 | 6.4 | 6.8 | / | / | / | / | / | /
WK02/2022 | HI93740A-0 | Nickel Low Range Reagent A | L57 Powder | R052 | colorimetric | HI801 | HI83399 | Nickel Low Range (Ni) | Nickel | µg/L | 0 | 1000 | 0 | 10 | & | 7% | 100% | MR21 | 0 | 0 | 7 | 200 | 187 | 213 | 800 | 769 | 831 | / | / | / | / | / | / | / | / | / | 4.9 | 4.7 | 5.1 | / | / | / | 5.3 | 5.1 | 5.5 | 200 | 185 | 215
WK02/2022 | HI93740B-0 | Nickel Low Range Reagent B | L56 CTK | B169 | colorimetric | HI801 | HI83399 | Nickel Low Range (Ni) | Nickel | µg/L | 0 | 1000 | 0 | 10 | & | 7% | 100% | MR21 | 0 | 0 | 7 | 200 | 187 | 213 | 800 | 769 | 831 | / | / | / | / | / | / | / | / | / | 4.9 | 4.7 | 5.1 | / | / | / | 5.3 | 5.1 | 5.5
WK02/2022 | HI93740C-0 | Nickel Low Range Reagent C | L57 Powder | R053 | colorimetric | HI801 | HI83399 | Nickel Low Range (Ni) | Nickel | µg/L | 0 | 1000 | 0 | 10 | & | 7% | 100% | MR21 | 0 | 0 | 7 | 200 | 187 | 213 | 800 | 769 | 831 | / | / | / | / | / | / | / | / | / | 4.9 | 4.7 | 5.1 | / | / | / | 5.3 | 5.1 | 5.5 | 200 | 185 | 215
WK47/2021 | HI93746-0 | Iron Low Range Reagent | L57 Powder | R054 | colorimetric | HI801 | HI83399 | Iron Low Range (Fe) | Iron | mg/L | 0 | 1.6 | 3 | 0.01 | & | 8% | 100% | MR15 | 0 | 0 | 0.01 | 0.080 |  |  | 0.320 |  |  | / | / | / | / | / | / | / | / | / | 2.9 | 2.7 | 3.1 | / | / | / | / | / | / | 260 | 250 | 270 | X
WK02/2022 | HI93748A-0 | Manganese Low Range Reagent A | L57 Powder | R023 | colorimetric | HI801 | HI83399 | Manganese Low Range(Mn) | Manganese | µg/L | 0 | 300 | 0 | 7 | & | 3% | 100% | MR29 | 0 | 0 | 10 | 60 | 48 | 72 | 240 | 223 | 257 | / | / | / | / | / | / | / | / | / | 3 | 2.9 | 3.2 | / | / | / | / | / | / | 50 | 40 | 60
WK02/2022 | HI93748B-0 | Manganese Low Range Reagent B | L56 CTK | B170 | colorimetric | HI801 | HI83399 | Manganese Low Range(Mn) | Manganese | µg/L | 0 | 300 | 0 | 7 | & | 3% | 100% | MR29 | 0 | 0 | 10 | 60 | 48 | 72 | 240 | 223 | 257 | / | / | / | / | / | / | / | / | / | 3 | 2.9 | 3.2 | / | / | / | / | / | / | 50 | 40 | 60
WK02/2022 | HI93748C-0 | Manganese Low Range Reagent C | L56 CTK | CP-B171 | colorimetric | HI801 | HI83399 | Manganese Low Range(Mn) | Manganese | µg/L | 0 | 300 | 0 | 7 | & | 3% | 100% | MR29 | 0 | 0 | 10 | 60 | 48 | 72 | 240 | 223 | 257 | / | / | / | / | / | / | / | / | / | 3 | 2.9 | 3.2 | / | / | / | / | / | / | 50 | 40 | 60
WK47/2021 | HI93749-0 | Chromium VI LR Reagent | L57 Powder | R043 | colorimetric | HI801 | HI83399 | Chromium VI Low Range (Cr) | Chromium VI | µg/L | 0 | 300 | 0 | 10 | & | 4% | 100% | MR43 | 0 | 0 | 10 | 60 | 48 | 71 | 240 | 220 | 260 | / | / | / | / | / | / | / | / | / | 1.3 | 1.1 | 1.6 | / | / | / | / | / | / | 120 | 110 | 130
WK02/2022 | HI93750A-0 | Potassium Reagent A | L56 CTK | B021 | colorimetric | HI801 | HI83399 | Potassium MR (K) | Potassium | mg/L | 0 | 200 | 1 | 3 | & | 7% | 100% | MR30 | 0 | 0 | 3 | 4 | 0.7 | 7.2 | 16 | 11.9 | 20.1 | 80 | / | / | 180 | [5 empty] | 4 | 3.8 | 4.2 | 6.2 | 5.9 | 6.5 | / | / | / | / | / | /
WK02/2022 | HI93750B-0 | Potassium Reagent B | L57 Powder | R055 | colorimetric | HI801 | HI83399 | Potassium MR (K) | Potassium | mg/L | 0 | 200 | 1 | 3 | & | 7% | 100% | MR30 | 0 | 0 | 3 | 4 | 0.7 | 7.3 | 16 | 11.9 | 20.1 | 80 | / | / | 180 | / | / | / | / | / | 4 | 3.8 | 4.2 | 6.2 | 5.9 | 6.5 | / | / | / | 110 | 100 | 120
WK47/2021 | HI93751-0 | Sulfate Reagent | L57 Powder | R056 | colorimetric | HI801 | HI83399 | Sulfate (SO4) | Sulphate | mg/L | 0 | 150 | 0 | 5 | & | 3% | 100% | MR32 | 0 | 0 | 5 | 20 | 14 | 26 | 80 | 73 | 87 | / | / | / | / | / | / | / | / | / | 2 | 1.8 | 2.2 | / | / | / | / | / | / | 140 | 130 | 150
WK47/2021 | HI93752A-CA | Calcium Reagent A | L56 CTK | B172 | colorimetric | HI801 | HI83399 | Calcium | Calcium | mg/L | 0 | 400 | 0 | 10 | & | 5% | 100% | MR20 | 0 | 0 | 10 | 80 | 66 | 94 | 320 | 294 | 346 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI93752A-MG | Magnesium Reagent A | L56 CTK | B173 | colorimetric | HI801 | HI83399 | Magnesium | Magnesium | mg/L | 0 | 150 | 0 | 5 | & | 3% | 100% | MR28 | 0 | 0 | 5 | 30 | 24 | 36 | 120 | 111 | 129 | / | / | / | / | / | / | / | / | / | / | / | / | 8.4 | 8.3 | 8.5 | / | / | / | / | / | /
WK47/2021 | HI93752B-CA | Calcium Reagent B | L56 CTK | B174 | colorimetric | HI801 | HI83399 | Calcium | Calcium | mg/L | 0 | 400 | 0 | 10 | & | 5% | 100% | MR20 | 0 | 0 | 10 | 80 | 66 | 94 | 320 | 294 | 346 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI93752B-MG | Magnesium Reagent B | L56 CTK | B157 | colorimetric | HI801 | HI83399 | Magnesium | Magnesium | mg/L | 0 | 150 | 0 | 5 | & | 3% | 100% | MR28 | 0 | 0 | 5 | 30 | 24 | 36 | 120 | 111 | 129 | / | / | / | / | / | / | / | / | / | / | / | / | 8.4 | 8.3 | 8.5 | / | / | / | / | / | /
WK47/2021 | HI93753A-0 | Chloride Reagent A | L56 CTK | B175 | colorimetric | HI801 | HI83399 | Chloride | Cl | mg/L | 0 | 20 | 1 | 0.5 | & | 6% | 100% | MR39 | 0 | 0 | 0.5 | 4 | 3.3 | 4.7 | 16 | 14.5 | 17.5 | / | / | / | / | / | / | / | / | / |  |  |  | 1.5 | 1 | 2 | / | / | /
WK47/2021 | HI93753B-0 | Chloride Reagent A | L56 CTK | B176 | colorimetric | HI801 | HI83399 | Chloride | Cl | mg/L | 0 | 20 | 1 | 0.5 | & | 6% | 100% | MR39 | 0 | 0 | 0.5 | 4 | 3.3 | 4.7 | 16 | 14.5 | 17.5 | / | / | / | / | / | / | / | / | / |  |  |  | 1.5 | 1 | 2 | / | / | /
WK47/2021 | HI93754A-0 | COD Low Range Reagent | L59 COD | D001 | colorimetric | HI801 | HI83399 | COD LR | Oxygen | mg/L | 0 | 150 | 0 | 5 | or | 4% | 100% | MR50 | 0 | 0 | 5 | 20 | 15 | 25 | 50 | 45 | 55 | 75 | 70 | 80 | 100 | 95 | 105 | 140 | 134 | 151 | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI93754B-0 | COD Medium Range Reagent | L59 COD | D002 | colorimetric | HI801 | HI83399 | CODMR | Oxygen | mg/L | 0 | 1500 | 0 | 15 | or | 4% | 100% | MR50 | 0 | 0 | 15 | 200 | 185 | 215 | 500 | 485 | 515 | 750 | 720 | 780 | 1000 | 960 | 1040 | 1400 | 1344 | 1471 | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI93754C-0 | COD High Range EPA (13 mm vial) | [8 empty] | 0 | 15000 | 0 | 150 | or | 2%
WK47/2021 | HI93754D-0 | COD Low Range Reagent | L59 COD | D006 | colorimetric | HI801 | HI83399 | COD-LR.W/HG | Oxygen | mg/L | 0 | 150 | 0 | 5 | or | 4% | 100% | MR50 | 0 | 0 | 5 | 20 | 15 | 25 | 50 | 45 | 55 | 75 | 70 | 80 | 100 | 95 | 105 | 140 | 134 | 151 | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI93754E-0 | COD Medium Range Reagent | L59 COD | D008 | colorimetric | HI801 | HI83399 | COD-MR.W/HG | Oxygen | mg/L | 0 | 1500 | 0 | 15 | or | 3% | 100% | MR50 | 0 | 0 | 15 | 200 | 185 | 215 | 500 | 485 | 515 | 750 | 728 | 773 | 1000 | 970 | 1030 | 1400 | 1358 | 1457 | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI93754F-0 | COD Low Range Reagent | L59 COD | D004 | colorimetric | HI801 | HI83399 | COD-LR.ISO | Oxygen | mg/L | 0 | 150 | 0 | 5 | or | 4% | 100% | MR50 | 0 | 0 | 5 | 20 | 15 | 25 | 50 | 45 | 55 | 75 | 70 | 80 | 100 | 95 | 105 | 140 | 134 | 151 | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI93754G-0 | COD Medium Range Reagent | L59 COD | D007 | colorimetric | HI801 | HI83399 | COD-MR.ISO | Oxygen | mg/L | 0 | 1500 | 0 | 15 | or | 3% | 100% | MR50 | 0 | 0 | 15 | 200 | 185 | 215 | 400 | 385 | 415 | 550 | 534 | 567 | 750 | 728 | 773 | 950 | 922 | 994 | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI93754J-0 | COD ULTRA HIGHT RANGE Reagent | L59 COD | D29 | Colorimetric | HI801 | HI83399 | COD-UHR | Oxygen | g/L | 0 | 60 | 1 | 0.5 | & | 3% | 100% | MR50 | 0 | 0 | 0.5 | 6 | 5.3 | 6.7 | 18 | 17 | 19 | 30 | 28.6 | 31.4 | 36 | 34.4 | 37.6 | 54 | 51.9 | 56.1 | [9 empty] | 7.427 | 7.39 | 7.464
WK47/2021 | HI93757-0 | Ozone Reagent | L57 Powder | R035 | colorimetric | HI801 | HI83399 | Ozone | Ozone | mg/L | 0 | 2 | 2 | 0.02 | & | 3% | 100% | MR67 | 0 | 0 | 0.02 | 0.4 | 0.37 | 0.43 | 1.6 | 1.53 | 1.67 | / | / | / | / | / | / | / | / | / | 6.1 | 5.7 | 6.3 | / | / | / | / | / | / | 150 | 135 | 165
WK47/2021 | HI93758-0 | Phosphorus Reagent | L57 Powder | R029 | colorimetric | HI801 | HI83399 | Phosphorus Total (P) | Phosphorus | mg/L | 0 | 1.6 | 2 | 0.05 | or | 5% | 60% | MR24 | 0 | 0 | 0.05 | 0.25 | 0.2 | 0.3 | 1 | 0.95 | 1.05 | / | / | / | / | / | / | / | / | / | 1.5 | 1.2 | 1.6 | / | / | / | / | / | / | 180 | 165 | 195
WK47/2021 | HI93758A-0 | Reactive Phosphorus Reagent Vial | L59 COD | D012 | colorimetric | HI801 | HI83399 | Phosphorus reactive | Phosphorus | mg/L | 0 | 1.6 | 2 | 0.05 | or | 4% | 100% | MR24 | 0 | 0 | 0.05 | 0.3 | 0.25 | 0.35 | 1.2 | 1.15 | 1.25 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI93758V-0 | Phosphorus Reagent Vial | L59 COD | D015 | colorimetric | HI801 | HI83399 | Phosphorus Total LR | Phosphorus | mg/L | 0 | 1.15 | 2 | 0.05 | or | 6% | 100% | MR52 | 0 | 0 | 0.05 | 0.25 | 0.2 | 0.3 | 1 | 0.94 | 1.06 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI93758V-0AH | Acid Hydrolyzable Phosphorus Reagent Vial | L59 COD | D015 | colorimetric | HI801 | HI83399 | Phosphorus Acid hydrolizabil | Phosphorus | mg/L | 0 | 1.6 | 2 | 0.05 | or | 5% | 100% | MR51 | 0 | 0 | 0.05 | 0.3 | 0.25 | 0.35 | 1.2 | 1.14 | 1.26 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI93758V-0HR | Total Phosphorus High Range Reagent Vial | L59 COD | D015 | colorimetric | HI801 | HI83399 | Phosphorus Total HR | Phosphorus | mg/L | 0 | 32.6 | 1 | 0.5 | or | 5% | 100% | MR52 | 0 | 0 | 0.5 | 6.5 | 6 | 7 | 26 | 24.7 | 27.3 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI93763A-0 | Reactive Phosphorous High Range Reagent Vial | L59 COD | D017 | colorimetric | HI801 | HI83399 | Phosphorus reactive HR | Phosphorus | mg/L | 0 | 32.6 | 1 | 0.5 | or | 4% | 100% | MR24 | 0 | 0 | 0.5 | 6.5 | 6 | 7 | 26 | 25 | 27 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI93763B-0 | Reactive Phosphorous High Range Reagent B | L56 CTK | B148 | colorimetric | HI801 | HI83399 | Total Phosphorus HR | Phosphorus | mg/L | 0 | 32.6 | 1 | 0.5 | or | 5% | 100% | MR52 | 0 | 0 | 0.5 | 6.5 | 6 | 7 | 26 | 24.7 | 27.3 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI93764-0 | Nessler Reagent | L56 CTK | B135 | colorimetric | HI801 | HI83399 | Ammonia HR | N-Ammonia | mg/L | 0 | 100 | 1 | 1 | or | 5% | 100% | MR09 | 0 | 0 | 1 | 20 | 19 | 21 | 80 | 76 | 84 | / | / | / | / | / | / | / | / | / | 12.3 | 12 | 12.6 | 12.6 | 12.4 | 12.8 | / | / | / | / | / | /
WK47/2021 | HI93764A-0 | Ammonia Low Range Reagent Vial | L59 COD | D016 | colorimetric | HI801 | HI83399 | Ammonia LR | N-Ammonia | mg/L | 0 | 3 | 2 | 0.1 | or | 5% | 100% | MR09 | 0 | 0 | 0.1 | 2.5 | 2.4 | 2.63 | / | / | / | / | / | / | / | / | / | / | / | / | 12.3 | 12 | 12.6 | 13 | 12.8 | 13.3 | / | / | / | / | / | /
WK47/2021 | HI93764B-0 | Ammonia High Range Reagent Vial | L59 COD | D018 | colorimetric | HI801 | HI83399 | Ammonia HR | N-Ammonia | mg/L | 0 | 100 | 1 | 1 | or | 5% | 100% | MR09 | 0 | 0 | 1 | 20 | 19 | 21 | 80 | 76 | 84 | / | / | / | / | / | / | / | / | / | 12.3 | 12 | 12.6 | 12.6 | 12.4 | 12.8 | / | / | / | / | / | /
WK47/2021 | HI93766-0 | Nitrate Reagent | L57 Powder | R059 | colorimetric | HI801 | HI83399 | Nitrate (NO3-N) | N-Nitrate | mg/L | 0 | 30 | 1 | 1 | or | 3% | 100% | MR83 | 0 | 0 | 1 | 20 | 19 | 21 | / | / | / | / | / | / | / | / | / | / | / | / | 4.3 | 4.1 | 4.5 | / | / | / | / | / | / | 100 | 90 | 110
WK47/2021 | HI93766V-0 | Nitrate Reagent Vial | L59 COD | D019 | colorimetric | HI801 | HI83399 | Nitrate | N-Nitrate | mg/L | 0 | 30 | 1 | 1 | or | 3% | 100% | MR83 | 0 | 0 | 1 | 20 | 19 | 21 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI93766V-0HR | Total Nitrogen High Range Reagent Vial | L59 COD | D019 | colorimetric | HI801 | HI83399 | Nitrogen Total HR | Nitrogen | mg/L | 0 | 150 | 0 | 3 | or | 4% | 100% | MR09 | 0 | 0 | 3 | 75 | 72 | 78 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI93766V-0LR | Total Nitrogen Low Range Reagent Vial | L59 COD | D019 | colorimetric | HI801 | HI83399 | Nitrogen Total LR | Nitrogen | mg/L | 0 | 25 | 1 | 1 | or | 5% | 100% | MR09 | 0 | 0 | 1 | 20 | 19 | 21 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI93767-0 | Total Nitrogen Reagent | L57 Powder | R059 | colorimetric | HI801 | HI83399 | Nitrogen total HR | Nitrogen | mg/L | 0 | 30 | 1 | 1 | or | 3% | 100% | MR83 | 0 | 0 | 1 | 20 | 19 | 21 | / | / | / | / | / | / | / | / | / | / | / | / | 4.3 | 4.1 | 4.5 | / | / | / | / | / | / | 100 | 90 | 110 | X
WK47/2021 | HI93767A-B | Total Nitrogen Digestion Vial | L59 COD | D020 | colorimetric | HI801 | HI83399 | Nitrogen Total LR | Nitrogen | mg/L | 0 | 25 | 1 | 1 | or | 5% | 100% | MR09 | 0 | 0 | 1 | 20 | 19 | 21 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI93767B-B | Total Nitrogen High Range Digestion Vial | L59 COD | D020 | colorimetric | HI801 | HI83399 | Nitrogen Total HR | Nitrogen | mg/L | 0 | 15 | 0 | 3 | or | 4% | 100% | MR09 | 0 | 0 | 10 | 50 | 47 | 53 | 75 | 72 | 78 | 120 | 115 | 125 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI94754A-0 | COD Low Range Reagent with Barcode | L59 COD | D001 | colorimetric | HI801 | HI83399 | COD94LR | Oxygen | mg/L | 0 | 150 | 0 | 5 | or | 5% | 100% | MR50 | 0 | 0 | 5 | 20 | 15 | 25 | 50 | 45 | 55 | 75 | 70 | 80 | 100 | 95 | 105 | 140 | 133 | 152 | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI94754B-0 | COD Medium Range Reagent with Barcode | L59 COD | D002 | colorimetric | HI801 | HI83399 | COD94MR | Oxygen | mg/L | 0 | 1500 | 0 | 15 | or | 4% | 100% | MR50 | 0 | 0 | 15 | 200 | 185 | 215 | 500 | 485 | 515 | 750 | 720 | 780 | 1000 | 960 | 1040 | 1400 | 1344 | 1471 | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI94754D-0 | COD Low Range Reagent with Barcode | L59 COD | D006 | colorimetric | HI801 | HI83399 | COD94LR.W/HG | Oxygen | mg/L | 0 | 150 | 0 | 5 | or | 5% | 100% | MR50 | 0 | 0 | 5 | 20 | 15 | 25 | 50 | 45 | 55 | 75 | 70 | 80 | 100 | 95 | 105 | 140 | 133 | 152 | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI94754E-0 | COD Medium Range Reagent with Barcode | L59 COD | D008 | colorimetric | HI801 | HI83399 | COD94MR.W/HG | Oxygen | mg/L | 0 | 1500 | 0 | 15 | or | 4% | 100% | MR50 | 0 | 0 | 15 | 200 | 185 | 215 | 500 | 485 | 515 | 750 | 720 | 780 | 1000 | 960 | 1040 | 1400 | 1344 | 1471 | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI94754F-0 | COD Low Range Reagent with Barcode | L59 COD | D004 | colorimetric | HI801 | HI83399 | COD94LR.ISO | Oxygen | mg/L | 0 | 150 | 0 | 5 | or | 5% | 100% | MR50 | 0 | 0 | 5 | 20 | 15 | 25 | 50 | 45 | 55 | 75 | 70 | 80 | 100 | 95 | 105 | 140 | 133 | 152 | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI94754G-0 | COD Medium Range Reagent with Barcode | L59 COD | D007 | colorimetric | HI801 | HI83399 | COD94MR.ISO | Oxygen | mg/L | 0 | 1500 | 0 | 15 | or | 4% | 100% | MR50 | 0 | 0 | 15 | 200 | 185 | 215 | 400 | 385 | 415 | 550 | 528 | 572 | 750 | 720 | 780 | 950 | 912 | 1003 | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI94758A-0 | Reactive Phosphorus Reagent Vial with Barcode | L59 COD | D012 | colorimetric | HI801 | HI83399 | Phosphorus reactive | Phosphorus | mg/L | 0 | 1.6 | 2 | 0.05 | or | 5% | 100% | MR24 | 0 | 0 | 0.05 | 0.3 | 0.25 | 0.35 | 1.2 | 1.14 | 1.26 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI94758V-0 | Phosphorus Reagent Vial | L59 COD | D015 | colorimetric | HI801 | HI83399 | Phosphorus Total LR | Phosphorus | mg/L | 0 | 1.15 | 2 | 0.05 | or | 6% | 100% | MR52 | 0 | 0 | 0.05 | 0.25 | 0.2 | 0.3 | 1 | 0.94 | 1.06 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI94758V-0AH | Acid Hydrolyzable Phosphorus Reagent Vial with Barcode | L59 COD | D015 | colorimetric | HI801 | HI83399 | Phosphorus Acid hydrolizabil | Phosphorus | mg/L | 0 | 1.6 | 2 | 0.05 | or | 5% | 100% | MR51 | 0 | 0 | 0.05 | 0.3 | 0.25 | 0.35 | 1.2 | 1.14 | 1.26 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI94758V-0HR | Total Phosphorus High Range Reagent Vial with Barcode | L59 COD | D015 | colorimetric | HI801 | HI83399 | Phosphorus Total HR | Phosphorus | mg/L | 0 | 32.6 | 1 | 0.5 | or | 5% | 100% | MR52 | 0 | 0 | 0.5 | 6.5 | 6 | 7 | 26 | 24.7 | 27.3 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI94763A-0 | Reactive Phosphorous High Range Reagent Vial with Barcode | L59 COD | D017 | colorimetric | HI801 | HI83399 | Phosphorus reactive HR | Phosphorus | mg/L | 0 | 32.6 | 1 | 0.5 | or | 5% | 100% | MR24 | 0 | 0 | 0.5 | 6.5 | 6 | 7 | 26 | 24.7 | 27.3 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI94764A-0 | Ammonia Low Range Reagent Vial with Barcode | L59 COD | D016 | colorimetric | HI801 | HI83399 | Ammonia LR | N-Ammonia | mg/L | 0 | 3 | 2 | 0.1 | or | 5% | 100% | MR09 | 0 | 0 | 0.1 | 2.5 | 2.4 | 2.63 | / | / | / | / | / | / | / | / | / | / | / | / | 12.3 | 12 | 12.6 | 13 | 12.8 | 13.3 | / | / | / | / | / | /
WK47/2021 | HI94764B-0 | Ammonia High Range Reagent Vial with Barcode | L59 COD | D018 | colorimetric | HI801 | HI83399 | Ammonia HR | N-Ammonia | mg/L | 0 | 100 | 0 | 1 | or | 5% | 100% | MR09 | 0 | 0 | 1 | 20 | 19 | 21 | 80 | 76 | 84 | / | / | / | / | / | / | / | / | / | 12.3 | 12 | 12.6 | 13 | 12.8 | 13.3 | / | / | / | / | / | /
WK47/2021 | HI94766V-0 | Nitrate and Total Nitrogen Reagent Vial with Barcode | L59 COD | D019 | colorimetric | HI801 | HI83399 | Nitrate | N-Nitrate | mg/L | 0 | 30 | 1 | 1 | or | 5% | 100% | MR83 | 0 | 0 | 1 | 20 | 19 | 21 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI94766V-0HR | Total Nitrogen High Range Reagent Vial with Barcode | L59 COD | D019 | colorimetric | HI801 | HI83399 | Nitrogen Total HR | Nitrogen | mg/L | 10 | 150 | 0 | 3 | or | 4% | 100% | MR09 | 0 | 0 | 3 | 75 | 72 | 78 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI94766V-0LR | Total Nitrogen Low Range Reagent Vial with Barcode | L59 COD | D019 | colorimetric | HI801 | HI83399 | Nitrogen Total LR | N-Nitrate | mg/L | 0 | 25 | 1 | 1 | or | 5% | 100% | MR09 | 0 | 0 | 1 | 20 | 19 | 21 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI94767A-B | Total Nitrogen Low Range Digestion Vial with Barcode | L59 COD | D020 | colorimetric | HI801 | HI83399 | Nitrogen Total LR | N-Nitrate | mg/L | 0 | 25 | 1 | 1 | or | 5% | 100% | MR09 | 0 | 0 | 1 | 20 | 19 | 21 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI94767B-B | Total Nitrogen High Range Digestion Vial with Barcode | L59 COD | D020 | colorimetric | HI801 | HI83399 | Nitrogen Total HR | Nitrogen | mg/L | 10 | 150 | 0 | 3 | or | 4% | 100% | MR09 | 0 | 0 | 10 | 50 | 47 | 53 | 75 | 72 | 78 | 120 | 115 | 125 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI95747-0 | Copper Low Range Reagent | L57 Powder | R060 | colorimetric | HI801 | HI83399 | Copper Low Range (Cu) | Copper | mg/L | 0 | 1.5 | 3 | 0.01 | & | 5% | 100% | MR44 | 0 | 0 | 0.01 | 0.3 | 0.275 | 0.325 | 1.2 | 1.13 | 1.27 | / | / | / | / | / | / | / | / | / | 6.5 | 6.2 | 6.8 | / | / | / | / | / | / | 270 | 260 | 280
WK47/2021 | HI95761-0 | Total Chlorine Reagent | L57 Powder | R061 | colorimetric | HI801 | HI83399 | Total Chlorine Ultra Low Range | Chlorine | mg/L | 0 | 0.5 | 3 | 0.02 | & | 3% | 100% | MR67 | 0 | 0 | 0.02 | 0.1 | 0.077 | 0.123 | 0.4 | 0.368 | 0.432 | / | / | / | / | / | / | / | / | / | 6 | 5.7 | 6.3 | / | / | / | / | / | / | 150 | 135 | 165
WK47/2021 | HI95762-0 | Free Chlorine Reagent | L57 Powder | R062 | colorimetric | HI801 | HI83399 | Free Chlorine Ultra Low Range | Chlorine | mg/L | 0 | 0.5 | 3 | 0.02 | & | 3% | 100% | MR67 | 0 | 0 | 0.02 | 0.1 | 0.077 | 0.122 | 0.4 | 0.368 | 0.432 | / | / | / | / | / | / | / | / | / | 6 | 5.7 | 6.3 | / | / | / | / | / | / | 120 | 110 | 130
WK47/2021 | HI95769A-0 | Anionic Surfactants Reagent A | L56 CTK | B071 | colorimetric | HI801 | HI83399 | Surfactant, anionic | SDBS | mg/L | 0 | 3.5 | 2 | 0.04 | & | 3% | 100% | MR92 | 0.7 | 0.64 | 0.76 | 2.8 | 2.68 | 2.92 | [9 empty] | / | / | / | 10 | 9.8 | 10.2 | 10 | 9.8 | 10.2 | / | / | / | / | / | /
WK47/2021 | HI95769B-0 | Anionic Surfactants Reagent A | L56 CTK | B072 | colorimetric | HI801 | HI83399 | Surfactant, anionic | SDBS | mg/L | 0 | 3.5 | 2 | 0.04 | & | 3% | 100% | MR92 | 0.7 | 0.64 | 0.76 | 2.8 | 2.68 | 2.92 | [9 empty] | / | / | / | 10 | 9.8 | 10.2 | 10 | 9.8 | 10.2 | / | / | / | / | / | /
WK47/2021 | HI95771A-0 | Chlorine Ultra High Range Reagent A | L57 Powder | R021 | colorimetric | HI801 | HI83399 | Chlorine Ultra High Range | Chlorine | mg/L | 0 | 500 | 0 | 3 | & | 3% | 100% | MR67 | 0 | 0 | 3 | 100 | 94 | 106 | 400 | 385 | 415 | / | / | / | / | / | / | / | / | / | 7 | 6.8 | 7.2 | 1.5 | 1.4 | 1.7 | / | / | / | 130 | 115 | 145
WK47/2021 | HI95771B-0 | Chlorine Ultra High Range Reagent B | L57 Powder | R063 | colorimetric | HI801 | HI83399 | Chlorine Ultra High Range | Chlorine | mg/L | 0 | 500 | 0 | 3 | & | 3% | 100% | MR67 | 0 | 0 | 3 | 100 | 94 | 106 | 400 | 385 | 415 | / | / | / | / | / | / | / | / | / | 7 | 6.8 | 7.2 | 1.5 | 1.4 | 1.7 | / | / | / | 100 | 90 | 110
WK47/2021 | HI96770A-0 | Silica High Range Reagent A | L57 Powder | R064 | colorimetric | HI801 | HI83399 | Silica High Range (SiO2) | Silica Oxide | mg/L | 0 | 200 | 0 | 1 | & | 5% | 100% | MR82 | 0 | 0 | 1 | 40 | 37 | 43 | 160 | 151 | 169 | / | / | / | / | / | / | / | / | / | 8.2 | 8.1 | 8.4 | 1 | 0.8 | 1.2 | 2.2 | 2 | 2.5 | 120 | 110 | 130
WK47/2021 | HI96770B-0 | Silica High Range Reagent B | L57 Powder | R065 | colorimetric | HI801 | HI83399 | Silica High Range (SiO2) | Silica Oxide | mg/L | 0 | 200 | 0 | 1 | & | 5% | 100% | MR82 | 0 | 0 | 1 | 40 | 37 | 43 | 160 | 151 | 169 | / | / | / | / | / | / | / | / | / | 8.2 | 8.1 | 8.4 | 1 | 0.8 | 1.2 | 2.2 | 2 | 2.5 | 330 | 315 | 345
WK47/2021 | HI96770C-0 | Silica High Range Reagent C | L57 Powder | R014 | colorimetric | HI801 | HI83399 | Silica High Range (SiO2) | Silica Oxide | mg/L | 0 | 200 | 0 | 1 | & | 5% | 100% | MR82 | 0 | 0 | 1 | 40 | 37 | 43 | 160 | 151 | 169 | / | / | / | / | / | / | / | / | / | 8.2 | 8.1 | 8.4 | 1 | 0.8 | 1.2 | 2.2 | 2 | 2.5 | 100 | 90 | 110
WK47/2021 | HI96773A-0 | Oxygen Scavenger Reagent A | L57 Powder | R072 | colorimetric | HI83399 | HI801 | Oxygen Scavengers Reagents | Carbohydrazide | mg/L | 0 | 150 | 2 | 0.02 | & | 3% | 100% |  | / | / | / | / | / |  | / | -0.02 | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI96773B-0 | Oxygen Scavenger Reagent B | L56 CTK | B198 | colorimetric | HI83399 | HI801 | Oxygen Scavengers Reagents | Carbohydrazide | mg/L | 0 | 150 | 2 | 0.02 | & | 3% | 100% | MR93 | 0 | / | / | 0.35 | / |  | 0.7 | -0.02 | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI96776-0 | Iron(II) Reagent | L57 Powder | R073 | colorimetric | HI801 | HI83399 | Iron (II) | Iron | mg/L | 0 | 6 | 2 | 0.1 | & | 2% | 100% |  | 0 | 0 | 0.1 | 2 | 1.86 | 2.14 | 5 | 4.8 | 5.2 | / | / | / | / | / | / | / | / | / | 5.2 | 4 | 6.5 | / | / | / | / | / | / | 140 | 125 | 155
WK47/2021 | HI96777A-0 | Iron(II)/(III) Reagent A | L57 Powder | R073 | colorimetric | HI801 | HI83399 | Iron (II)/(III) | Iron | mg/L | 0 | 6 | 2 | 0.1 | & | 2% | 100% |  | 0 | 0 | 0.1 | 2 | 1.86 | 2.14 | 5 | 4.8 | 5.2 | / | / | / | / | / | / | / | / | / | 5.2 | 4 | 6.5 | / | / | / | / | / | / | 140 | 125 | 155
WK47/2021 | HI96777B-0 | Iron(II)/(III) Reagent B | L57 Powder | R074 | colorimetric | HI801 | HI83399 | Iron (II)/(III) | Iron | mg/L | 0 | 6 | 2 | 0.1 | & | 2% | 100% |  | 0 | 0 | 0.1 | 2 | 1.86 | 2.14 | 5 | 4.8 | 5.2 | / | / | / | / | / | / | / | / | / | 5.2 | 4 | 6.5 | 5.2 | 4 | 6.5 | / | / | / | 105 | 95 | 115
WK47/2021 | HI96778A-0 | Total Iron Reagent A | L56 CTK | B199 | colorimetric | HI801 | HI83399 | Iron Total | Iron | mg/L | 0 | 7 | 2 | 0.2 | or | 3% | 20% | CM357 | 0 | 0 | 0.2 | 2 | 1.8 | 2.2 | 5 | 4.8 | 5.15 | / | / | / | / | / | / | / | / | / | 5.2 | 4 | 6.5 | 5.2 | 4 | 6.5 | / | / | / | / | / | /
WK47/2021 | HI96778B-0 | Total Iron Reagent B | L57 Powder | R041 | colorimetric | HI801 | HI83399 | Iron Total | Iron | mg/L | 0 | 7 | 2 | 0.2 | or | 3% | 100% | CM357 | 0 | 0 | 0.2 | 2 | 1.8 | 2.2 | 5 | 4.8 | 5.15 | / | / | / | / | / | / | / | / | / | 5.2 | 4 | 6.5 | 5.2 | 4 | 6.5 | / | / | / | 175 | 160 | 190
WK47/2021 | HI96778V-0 | Total Iron Digestion Vial | L59 COD | D022 | colorimetric | HI801 | HI83399 | Iron Total | Iron | mg/L | 0 | 7 | 2 | 0.2 | or | 3% | 100% | CM357 | 0 | 0 | 0.2 | 2 | 1.8 | 2.2 | 5 | 4.8 | 5.15 | / | / | / | / | / | / | / | / | / | 5.2 | 4 | 6.5 | 5.2 | 4 | 6.5 | / | / | / | / | / | /
WK47/2021 | HI96779A-0 | Chlorine Dioxide Reagent A | L56 CTK | B201 | colorimetric | HI801 | HI83399 | Chlorine Dioxide, Rapid method | Chlorine dioxide | mg/L | 0 | 2 | 2 | 0.2 | & | 3% | 100% | CM357 | 0 | 0 | 0.2 | 0.55 | 0.33 | 0.76 | 1.65 | 1.4 | 1.9 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI96779B-0 | Chlorine Dioxide Reagent B | L57 Powder | R025 | colorimetric | HI801 | HI83399 | Chlorine Dioxide, Rapid method | Chlorine dioxide | mg/L | 0 | 2 | 2 | 0.2 | & | 3% | 100% | CM357 | 0 | 0 | 0.2 | 0.55 | 0.33 | 0.76 | 1.65 | 1.4 | 1.9 | / | / | / | / | / | / | / | / | / | / | / | / | 6 | 5.7 | 6.3 | / | / | / | 120 | 110 | 130
WK47/2021 | HI96780V-0 | Nonionic Surfactants Reagent Vial | L59 COD | CP-96780V | colorimetric | HI801 | HI83399 | Surfactant, nonionic | TRITONX-100 | mg/L | 0 | 6 | 2 | 0.1 | & | 5% | 100% | MR90 | 0 | 0 | 0.1 | 0.5 | 0.38 | 0.63 | 2 | 1.8 | 2.2 | 3.5 | 3.23 | 3.78 | 5.5 | 5.13 | 5.88 |  |  |  | 8 | 7.8 | 8.2 | / | / | / | / | / | / | / | / | /
WK47/2021 | HI96781A-0 | Chromium Reagent A | L57 Powder | R075 | colorimetric | HI801 | HI83399 | Chromium (VI) Total | Chromium | µg/L | 0 | 1000 | 0 | 10 | or | 3% | 100% | MR91 | 0 | 0 | 10 | 200 | 190 | 210 | 400 | 390 | 410 | 800 | 776 | 824 | 900 | 873 | 927 |  |  |  | 1.5 | 1 | 2 | 1.5 | 1 | 2 | / | / | / | 60 | 50 | 70
WK47/2021 | HI96781B-0 | Chromium Reagent B | L57 Powder | R076 | colorimetric | HI801 | HI83399 | Chromium (VI) Total | Chromium | µg/L | 0 | 1000 | 0 | 10 | or | 3% | 100% | MR91 | 0 | 0 | 10 | 200 | 190 | 210 | 400 | 390 | 410 | 800 | 776 | 824 | 900 | 873 | 927 |  |  |  | 1.5 | 1 | 2 | 1.5 | 1 | 2 | / | / | / | 125 | 115 | 135
WK47/2021 | HI96781V-0 | Chromium Digestion Vial | L59 COD | D025 | colorimetric | HI801 | HI83399 | Chromium (VI) Total | Chromium | µg/L | 0 | 1000 | 0 | 10 | or | 3% | 100% | MR91 | 0 | 0 | 10 | 200 | 190 | 210 | 400 | 390 | 410 | 800 | 776 | 824 | 900 | 873 | 927 | / | / | / | 1.5 | 1 | 2 | 1.5 | 1 | 2 | / | / | / | / | / | /
WK47/2021 | HI96782A-0 | Anionic Surfactants Buffer Reagent A | L56 CTK | B202 | colorimetric | HI801 | HI83399 | Surfactant, anionic | SDBS | mg/L | 0 | 3.5 | 2 | 0.1 | & | 5% | 100% | MR92 | 0 | 0 | 0.1 | 0.5 | 0.38 | 0.62 | 1 | 0.85 | 1.15 | 2 | 3 | / | 3 | / | / |  |  |  | 10 | 9.8 | 10.2 | 10 | 9.8 | 10.2 | / | / | / | / | / | /
WK47/2021 | HI96782B-0 | Anionic Surfactants Indicator Reagent B | L56 CTK | B203 | colorimetric | HI801 | HI83399 | Surfactant, anionic | SDBS | mg/L | 0 | 3.5 | 2 | 0.1 | & | 5% | 100% | MR92 | 0 | 0 | 0.1 | 0.5 | 0.38 | 0.62 | 1 | 0.85 | 1.15 | 2 | 1.84 | 2.16 | 3 | 2.81 | 3.19 |  |  |  | 10 | 9.8 | 10.2 | 10 | 9.8 | 10.2 | / | / | / | / | / | /
WK47/2021 | HI96782V-0 | Anionic Surfactants Reagent Vial | L59 COD | D026 | colorimetric | HI801 | HI83399 | Surfactant, anionic | SDBS | mg/L | 0 | 3.5 | 2 | 0.1 | & | 5% | 100% | MR92 | 0 | 0 | 0.1 | 0.5 | 0.38 | 0.63 | 1 | 0.85 | 1.15 | 2 | 1.8 | 2.2 | 3 | 2.75 | 3.25 | / | / | / | 10 | 9.8 | 10.2 | 10 | 9.8 | 10.2 | / | / | / | / | / | /
WK47/2021 | HI96783-0 | Nitrite Low Range Reagent | L57 Powder | R077 | colorimetric | HI801 | HI83399 | Nitrite LR | Nitrite | µg/L | 0 | 600 | 0 | 10 | & | 3% | 100% | MR13 | 0 | 0 | 10 | 60 | 48 | 72 | 300 | 281 | 319 | 480 | 456 | 504 | / | / | / |  |  |  | 2.7 | 2.5 | 3 | / | / | / | / | / | / | 73 | 63 | 83
WK47/2021 | HI96783V-0 | Nitrite Low Range Reagent Vial | L59 COD | D027 | colorimetric | HI801 | HI83399 | Nitrite LR | Nitrite | µg/L | 0 | 600 | 0 | 10 | & | 3% | 100% | MR13 | 0 | 0 | 10 | 60 | 48 | 72 | 300 | 281 | 319 | 480 | 456 | 504 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | /
WK47/2021 | HI96784-0 | Nitrite Medium Range Reagent | L57 Powder | R077 | colorimetric | HI801 | HI83399 | Nitrite MR | Nitrite | mg/L | 0 | 6 | 2 | 0.1 | & | 3% | 100% | MR13 | 0 | 0 | 0.1 | 0.6 | 0.48 | 0.72 | 3 | 2.81 | 3.19 | 4.8 | 4.56 | 5.04 |  | / | / |  |  |  | 2.7 | 2.5 | 3 | / | / | / | / | / | / | 73 | 63 | 83
WK47/2021 | HI96784V-0 | Nitrite Medium Range Reagent Vial | L59 COD | D028 | colorimetric | HI801 | HI83399 | Nitrite MR | Nitrite | mg/L | 0 | 6 | 2 | 0.1 | & | 3% | 100% | MR13 | 0 | 0 | 0.1 | 0.6 | 0.48 | 0.72 | 3 | 2.81 | 3.19 | 4.8 | 4.56 | 5.04 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | /
WK02/2022 | HI96786V-0 | Iron Reagent Vial | L59 COD | D031 | colorimetric | HI801 | HI83399 | Iron | Fe | mg/L | 0 | 6 | 2 | 0.1 | & | 3% | 100% | MR33 | 0 | 0 | 0.1 | 1 | 0.87 | 1.13 | 3 | 2.81 | 3.19 | 5 | 4.75 | 5.25 | / | / | / | / | / | / | 2 | 1 | 2.2 | / | / | / | / | / | / | / | / | / | YES
WK02/2022 | HI96786-0 | Iron Vial | L57 Powder | R041 | colorimetric | HI801 | HI83399 | Iron | Fe | mg/L | 0 | 6 | 2 | 0.1 | & | 3% | 100% | MR33 | 0 | 0 | 0.1 | 1 | 0.87 | 1.13 | 3 | 2.81 | 3.19 | 5 | 4.75 | 5.25 | / | / | / | / | / | / | 2 | 1 | 2.2 | / | / | / | / | / | / | / | / | / | YES
WK02/2022 | HI96785V-0 | Cationic Surfactants Reagent Vial | L59 COD | D030 | colorimetric | HI801 | HI83399 | Cationic Surfactants | CTAB | mg/L | 0 | 2.5 | 2 | 0.15 | & | 3% | 100% | MR94 | 0 | 0 | 0.15 | 0.5 | 0.34 | 0.67 | 1 | 0.82 | 1.18 | 2 | 1.79 | 2.21 | / | / | / | / | / | / | 2 | 1 | 2.2 | / | / | / | / | / | / | / | / | / | YES
WK02/2022 | HI96785-0 | Cationic Reagent Vial | L57 Powder | R082 | colorimetric | HI801 | HI83399 | Cationic Surfactants | CTAB | mg/L | 0 | 2.5 | 2 | 0.15 | & | 3% | 100% | MR94 | 0 | 0 | 0.15 | 0.5 | 0.34 | 0.67 | 1 | 0.82 | 1.18 | 2 | 1.79 | 2.21 | / | / | / | / | / | / | 2 | 1 | 2.2 | / | / | / | / | / | / | / | / | / | YES
WK47/2021 | MANNITOL | Mannitol Reagent | L57 Powder | R066 | CTK | / | / | Boron Test Kit | Boron | mg/L | 0 | 5 | 1 | 0.2 | or | 0% | 100% |  | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | 6.2 | 6 | 6.4 | / | / | / | / | / | / | 300 | 270 | 330
WK47/2021 | MI512-0 | Phosphate Reagent | L57 Powder | R029 | colorimetric | / | / | Phosphate Low Range(PO4) | Phosphate | mg/L | 0 | 2.5 | 2 | 0.04 | & | 4% | 100% | MR24 | 0 | 0 | 0.04 | 0.5 | 0.44 | 0.56 | 2 | 1.88 | 2.12 | 3 | 2.84 | 3.16 | / | / | / | / | / | / | 1.5 | 1.2 | 1.6 | / | / | / | / | / | / | 180 | 165 | 195
WK47/2021 | MI524-0 | Total Chlorine Reagent | L57 Powder | R035 | colorimetric | / | / | Total Chlorine | Chlorine | mg/L | 0 | 5 | 2 | 0.03 | & | 3% | 100% | MR67 | 0 | 0 | 0.03 | 0.5 | 0.46 | 0.55 | 2.8 | 2.69 | 2.91 | / | / | / | / | / | / | / | / | / | 6 | 5.7 | 6.3 | / | / | / | / | / | / | 150 | 135 | 165
WK47/2021 | MI526-0 | Free Chlorine Reagent | L57 Powder | R025 | colorimetric | MI526 | / | Free Chlorine | Chlorine | mg/L | 0 | 5 | 2 | 0.03 | & | 3% | 100% | MR67 | 0 | 0 | 0.03 | 0.5 | 0.46 | 0.55 | 2 | 1.91 | 2.09 | / | / | / | / | / | / | / | / | / | 6 | 5.7 | 6.3 | / | / | / | / | / | / | 120 | 110 | 130
WK47/2021 | MI527-0 | Iodine Reagent | L57 Powder | R035 | colorimetric | MI527 | / | Iodine | Iodine | mg/L | 0 | 12.5 | 1 | 0.1 | & | 5% | 100% | MR67 | 0 | 0 | 0.1 | 2.5 | 2.3 | 2.7 | 10 | 9.4 | 10.6 | / | / | / | / | / | / | / | / | / | 6.1 | 5.7 | 6.3 | / | / | / | / | / | / | 150 | 135 | 165
WK47/2021 | MI528-0 | Iron Reagent | L57 Powder | R041 | colorimetric | MI528 | / | Iron High Range (Fe) | Iron | mg/L | 0 | 5 | 2 | 0.04 | & | 2% | 100% | MR33 | 0 | 0 | 0.04 | 1 | 0.94 | 1.05 | 4 | 3.88 | 4.12 | / | / | / | / | / | / | / | / | / | 5.2 | 4 | 6.5 | / | / | / | / | / | / | 175 | 160 | 190
WK47/2021 | MI552A-0 | Iron Reagent A | L57 Powder | R023 | colorimetric | / | / | Iron | Iron | mg/L | 0 | 15 | 1 | 0.2 | & | 5% | 100% | MR33 | 0 | 0 | 0.2 | 0.4 | 0.2 | 0.6 | 1.3 | 1 | 1.6 | / | / | / | / | / | / | / | / | / | 2.7 | 2.6 | 2.9 | / | / | / | / | / | / | 150 | 135 | 165
WK47/2021 | MI552B-0 | Iron Reagent B | L57 Powder | R022 | colorimetric | / | / | Iron | Iron | mg/L | 0 | 15 | 1 | 0.2 | & | 5% | 100% | MR33 | 0 | 0 | 0.2 | 0.4 | 0.2 | 0.6 | 1.3 | 1 | 1.6 | / | / | / | / | / | / | / | / | / | 2.7 | 2.6 | 2.9 | / | / | / | / | / | / | 152 | 137 | 167
WK47/2021 | MI555-006-0 | Stabilizer SO2 | L57 Powder | R024 | colorimetric | / | / | / | / |  |  |  | / | 0 | / | 0% | / |  | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | 7 | 6.8 | 7.2 | / | / | / | / | / | / | 270 | 260 | 280
WK03/2022 | MI590A-0 | Peroxide Reagent A (test vial) | L59 COD | D013 | Colorimetric | HI83730 | MI490 | ISM peroxide value in edible oils | Oxygen / kg | meq | 0 | 25 | 1 | 0.5 | or | 0% | 100% | MR75 | 0 | 0 | 0.5 | 11 | 10.5 | 20 |  | 0 | 0 | / | / | / | / | / | / | / | / | / | 7 | 6.8 | 7.2 | 1.5 | 1.3 | 1.7 |  |  |  | 5
WK03/2022 | MI590B-0 | Peroxide Reagent B | L57 Powder | R021 | colorimetric | HI83730 | MI490 | Peroxide | Peroxide | meq O2 /kg | 0 | 25 | 1 | 0.5 | or | 0% | 100% | MR75 | 0 | / | 0.5 | 11 | 10.5 | 11.5 | / | / | / | / | / | / | / | / | / | / | / | / | 7 | 6.8 | 7.2 | 1.5 | 1.4 | 1.7 | / | / | / | 100 | 90 | 110
WK47/2021 | MT5002-0 | Potassium Reagent | L57 Powder | R008 | colorimetric | / | / | Potassium Test | Potassium | mg/L | 0 | 72 | 0 | 0 | / | 0% | 100% | MR30 | 16 | 16 | 16 | 36 | 36 | 36 | 72 | 72 | 72 | / | / | / | / | / | / | / | / | / | 11.7 | 11.5 | 11.8 | / | / | / | / | / | / | 105 | 95 | 115
WK47/2021 | MT5009-0 | Nitrate Reagent | L57 Powder | R009 | colorimetric | / | / | Nitrogen Test | Nitrogen | mg/L | 0 | 7 | 0 | 0 | / | 0% | 100% | MR09 | 2 | 2 | 2 | 3 | 3 | 3 | 7 | 7 | 7 | / | / | / | / | / | / | / | / | / | 1.9 | 1.8 | 2.1 | / | / | / | / | / | / | 290 | 270 | 310
WK47/2021 | MT5010-0 | Phosphorus Reagent | L57 Powder | R029 | colorimetric | / | / | Phosphate Low Range(PO4) | Phosphate | mg/L | 0 | 2.5 | 2 | 0.04 | & | 4% | 100% | MR24 | 0 | 0 | 0.04 | 0.5 | 0.44 | 0.56 | 2 | 1.88 | 2.12 | / | / | / | / | / | / | / | / | / | 1.5 | 1.2 | 1.6 | / | / | / | / | / | / | 180 | 165 | 195
WK47/2021 | OXALATE | Oxalate Reagent | L57 Powder | R067 | CTK | / | / | Calcium Test Kit | Calcium | mg/L | 0 | 250 | 1 | 2 | & | 0% | 100% | MR20 | / | / | / | 15 | 13 | / | 100 | 98 | 102 | / | / | / | / | / | / | / | / | / | 6.2 | 6 | 6.4 | / | / | / | / | / | / | 225 | 210 | 240
WK47/2021 | PEPSIN | Pepsin | L57 Powder | R068 | CTK | / | / | / | / |  |  |  | / | 0 | / | 0% | / |  | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | 3.6 | 3.4 | 3.8 | / | / | / | / | / | / | 800 | 750 | 850
WK47/2021 | PERSULFATE/I | Potassium Persulfate Reagent | L57 Powder | R069 | colorimetric | HI801 | HI83399 | Iron Total | Iron | mg/L | 0 | 7 | 2 | 0.2 | or | 3% | 100% | CM357 | 0 | 0 | 0.2 | 2 | 1.8 | 2.2 | 5 | 4.8 | 5.15 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | 100 | 90 | 110
WK47/2021 | PERSULFATE/N | Potassium Persulfate Reagent | L57 Powder | R069 | colorimetric | HI801 | HI83399 | Nitrogen Total LR (N) | Nitrogen | mg/L | 0 | 25 | 1 | 1 | or | 5% | 60% | MR09 | 0 | 0 | 1 | 20 | 19 | 20.6 | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | 200 | 180 | 220 | X
WK02/2022 | pH-0 | pH Reagent | L56 CTK | B147 | colorimetric | HI83399 | MI411 | pH | pH | pH | 6.5 | 8.5 | 1 | 0.2 | & | 0% | 100% | HI70456(0.1 N NAOH/HI70463(0.1N HCL) | 6.6 | 6.6 | 7 | 7.5 | 7.8 | 8.2 | 7.9 |  |  | / | / | / | / | / | / | / | / | /
WK47/2021 | SULFAMIC | Sulfamic Acid | L57 Powder | R003 | CTK | / | / | / | / |  |  |  | / | 0 | / | 0% | / |  | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | / | 1 | 0.8 | 1.2 | / | / | / | / | / | / | 310 | 285 | 335

## STD prep MR L
MR Bottle | MR Description | MR Parameter | FW MR Parameter | MR Location | MR Stock q.ty | MR Stock Unit | MR EXP | MR physical state | MR Purity (%) | MR Value | MR Unit | STD Exp. | Storage STD
l'operatore sceglie dal MR data base la bottiglia aperta - dalla scelta compariranno i dati delle colonne F - G - H - I - J - K - L - M - N - O | Ammonium 1000 mg/l NH₄ traceable to SRM from NIST NH₄Cl in H₂O | NH4 | 18 |  | 600 | dal data base in mL se L, in g se S |  | L | 100% | 1000 | mg/L | 1 Day
Hanna Code | MR CODE | Parameter Formula | FW Parameter Formula | Conc Hanna Parameter (Calcolo) | Measurement Unit | STD MATRIX | STD Volume (mL) | STD Exp (days)
HI784A-0 | MR09 | NH3 | 17.0067 | 944.8167 | mg/L | SW4 | 500 | 10
|  |  |  | questo valore viene moltiplicato per mille se Measurement Unit è in µg/L | questa unità di misura è riferita anche agli STD (può essere mg/L o µg/L)
STD1Value | q.ty for STD1 |  |  |  | STD2Value | q.ty for STD2 |  |  |  | STD3Value | q.ty for STD3
0 | 0 | mg/L |  |  | 0.4 | 0.21168 | mg/L |  |  | 1.25 | 0.6615
STD4Value | q.ty for STD4 |  |  |  | STD5Value | q.ty for STD5 |  |  |  | STD6Value | q.ty for STD6
2.1 | 1.1113 |  |  |  | 3 | 1.5876 |  |  |  | / | #VALUE!
La somma dei valori in blu deve essere inferiore a MR STOCK SOPRA

## STD prep MS2
MR Bottle | MR Description | MR Parameter | FW MR Parameter | MR Location | MR Stock q.ty | MR Stock Unit | MR EXP | MR physical state | MR Purity (%) | MR Value | MR Unit | STD Exp. | Storage STD
l'operatore sceglie dal MR data base la bottiglia aperta - dalla scelta compariranno i dati delle colonne successive | Silicon STD AAS TraceCERT®, 1000 mg/L Si in NaOH | Si | 54.94 |  | dal data base in mL se L, in g se S | mL |  | L | 100% | 1000 | mg/L | 1 Day
MS2 dil | MS1 Volume (mL) | q.ty MR to add (calcolo) | MS2 Value (calcolo) | MS2 Exp (days)
10 | 100 | 10 | 100 | 10
| sceglie l'operatore | questo valore deve essere inferiore a Mr stock in verde sopra
Hanna Code | MR CODE | Hanna Parameter | FW Parameter Formula | Conc Hanna Parameter (Calcolo) | Measurement Unit | STD MATRIX | STD Volume (mL) | STD Exp (days)
HI705A-0 | MR45 | SiO2 | 54.94 | 100000 | mg/L | DWA | 250 | 10
|  |  |  | quasto valore viene moltiplicato per 1000 se Measurement Unit è in µg/L | questa unità di misura è riferita anche agli STD (può essere mg/L o µg/L)
STD1Value | q.ty for STD1 |  |  |  | STD2Value | q.ty for STD2 |  |  |  | STD3Value | q.ty for STD3
0 | 0 |  |  |  | 60 | 0.15 |  |  |  | 240 | 0.6
STD4Value | q.ty for STD4 |  |  |  | STD5Value | q.ty for STD5 |  |  |  | STD6Value | q.ty for STD6
2.1 | 0.00525 |  |  |  | 3 | 0.0075 |  |  |  | / | #VALUE!
La somma dei valori in blu deve essere inferiore a MS2 volume in girigio sopra

## STD prep MS1
MR Bottle | MR Description | MR Parameter | FW MR Parameter | MR Location | MR Stock q.ty | MR Stock Unit | MR EXP | MR physical state | MR Purity (%) | MR Value | MR Unit | STD Exp. | Storage STD
l'operatore sceglie dal MR data base la bottiglia aperta - dalla scelta compariranno i dati delle colonne F - G - H - I - J - K - L - M - N - O | Calcium carbonate BioUltra ≥99.0% (KT) - 250 gr | CaCO3 | 100.0869 |  | dal data base in mL se L, in g se S | g |  | S | 99% |  |  | 1 Day
MS1 Conc | MS1 Volume (mL) | q.ty MR (calcolo) | MS1 Exp (days)
5000 | 100 | 500 | 10
| sceglie l'operatore | questo valore deve essere inferiore a Mr stock in verde sopra
Hanna Code | MR CODE | Parameter Formula | FW Parameter Formula | Conc Hanna Parameter (Calcolo) | Measurement Unit | STD MATRIX | STD Volume (mL) | STD Exp (days)
HI755 | MR05 | CaCO3 | 100.0869 | 4950 | mg/L | SW4 | 500 | 10
|  |  |  | quasto valore viene moltiplicato per 1000 se Measurement Unit è in µg/L | questa unità di misura è riferita anche agli STD (può essere mg/L o µg/L)
STD1Value | q.ty for STD1 |  |  |  | STD2Value | q.tyfor STD2 |  |  |  | STD3Value | q.ty for STD3
0 | 0 |  |  |  | 50 | 5.0505 |  |  |  | 150 | 15.1515
STD4Value | q.ty for STD4 |  |  |  | STD5Value | q.ty for STD5 |  |  |  | STD6Value | q.ty for STD6
2.1 | 0.21212 |  |  |  | 3 | 0.30303 |  |  |  | / | #VALUE!
La somma dei valori in blu deve essere inferiore a MS1 volume in girigio sopra